MSP - 013

PROCEDURE FOR MANAGING INTERNAL QUALITY CONTROL PROGRAM

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Procedure	7

4.1	General Principles	7

4.2	Pre-Analytical Quality Control	8

4.3	Quality Control Test Performance	10

4.4	Qualitative Methods	11

4.5	Quantitative Methods	12

4.6	Post-Analytical Quality Control	14

4.7	Calibration Verification	15

4.8	Analytical Measurement Range Validation	15

4.9	Tools for data collection and analysis	16

4.10	Corrective Action	16

4.11	Reviews	17

5.	Quality Control	18

6.	Related Documents	18

7.	Reference Regulations & Guidelines	19

8.	Related Records	19

- APPROVAL AUTHORITY

| Prepared by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Issued by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Approved by: | Chairman & Managing Director | Mr. Venkata Cherukuri Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BME | Biomedical Engineering |
| 3 | CLSI | Clinical & Laboratory Standards Institute |
| 4 | CV | Coefficient of Variation |
| 5 | EP | Evaluation Protocols |
| 6 | FOM | Form |
| 7 | GE | General |
| 8 | ISO | International Organization for Standardization |
| 9 | MSDS | Material Safety Data Sheets |
| 10 | MSP | Management System Procedure |
| 11 | MAN | Manual |
| 12 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 13 | QA | Quality Assurance |
| 14 | QC | Quality Control |
| 15 | SD | Standard Deviation |
| 16 | TDPL | TRUSTlab Diagnostics Private Limited |

Objective

- The objective of this MSP is to outline a step-by-step procedure for managing the Internal Quality Control at TRUSTlab.

Scope

- This procedure applies to all testing activities within TRUSTlab.

Responsibilities

  - Accession Department

  - Lab Manager

  - Lab Head

Definitions

Sample: A biological specimen of a patient/test requestor that is received by TRUSTlab for specific diagnostic testing.

Procedure

General Principles

- Control specimens are tested in the same manner and by the same personnel as patient samples.

- Tolerance limits are defined for control procedures.

- Controls are used within the date of expiration.

- Reagents used to calibrate are not used as controls within the same run or analytical interval.

- The results of controls are verified for acceptability before patient results are reported.

- All control results are recorded.

- In case calibration and control materials are not commercially available, a method is defined to verify the accuracy of patient test results.

- If more than one method for testing an analyte is used, the methods are checked against each other twice a year for correlation of results.

- Where appropriate, manufacturer’s certification of lot acceptability will be retained as documentation of acceptable quality control performance. (This usually applies to media.)

- High complexity testing at TRUSTlab is performed by technical staff who meet the requirements for Lab Supervisor.

- Technical staff performing high complexity testing may perform testing, result testing, and review testing results in the absence of a Lab Supervisor.

- Technical staff are trained and competent to perform specific tasks as assigned.

- The identification of technical staff performing patient and quality control testing is documented.

Pre-Analytical Quality Control

- Specimen Acceptability should be defined for each department or method as applicable.

  - Patient identification and Specimen labeling is defined.

  - Specimen type, container, preservative that may be used, any special collection instructions, and conditions for transport.

  - Specimen processing or pre-analytical treatment instructions.

  - Specimen storage temperature requirements and length of storage limits.

  - As applicable for method, procedures define method of identification and sample integrity for capillary specimens, aliquots, and dilutions.

  - Criteria for rejection of unacceptable specimens is defined.

  - Instructions for special handling of sub-optimal specimens are outlined for methods as applicable.

- Test kit, shelf life and storage should be defined for each method as applicable.

  - All reagents, standards, calibrators and quality control (QC) materials purchased by the laboratory must have a date of receipt indicated on the container/box.

    - Reagent kit Inserts should be available for each item and maintained in the appropriate Laboratory Department.

    - Package inserts from each new lot must be reviewed to identify any changes. All changes must be acted upon and subsequently maintained in the Laboratory Department.

    - New or revised MSDS sheets are sent to BME department to be included in the laboratory's on-line QT QMs and Material Safety Data Sheet (MSDS) Folder (TDPL/GE/FOL-003).

  - The Name and Lot number of the reagent and/or test kit should be printed on the container or box.

  - The expiration date should be printed on the reagent and/or test kit container. Reagents and/or test kits will not be used past their expiration without the approval of the Pathologist.

  - Reagents will not be exchanged or used with other lots unless allowed by the manufacturer.

  - The “opened on” date should be mentioned on the container; the new expiration date/time as applicable should be on the container; the initials of the tech should be on the container.

  - The reagents and/or test kit must be stored at the temperature required by the manufacturer.

  - Any reagents prepared by the laboratory must be stored in a labeled container with the following information:

    - Reagent name

    - Date prepared

    - Expiration date

    - Storage temperature

    - Tech initials

- Materials used in testing should meet or exceed manufacturer's recommendations.

  - The type of water necessary for each method is defined.

  - Specifications for Laboratory water are defined.

  - It is ensured that adequate supply is available.

  - Frequency of testing the water quality is defined.

  - Glassware is cleaned in a manner that removes detergent residue that may interfere with testing.

  - Pipettes and pipette tips meet or exceed the accuracy required for the method.

  - The frequency of Pipette calibration is defined.

- Control reagents must be handled as required by the manufacturer.

  - Control reagents are reconstituted as described by the manufacturer’s instructions, when applicable.

  - Control reagent labels contain the name, lot, level (if applicable), expiration date/time, date opened, and initials of the Tech.

  - Control reagents are stored as per the manufacturer’s specifications.

  - Control reagents are not used beyond their expiration date/time.

- Instrument maintenance is defined as per manufacturer’s recommendations and compliance agencies’ requirements.

  - All analyzers are tested before being placed into service.

  - Instruments are put through defined validation protocols before being placed into service.

  - Required and/or scheduled, instrument checks and maintenance should be completed prior to testing.

  - Instrument checks and maintenance details are documented.

Quality Control Test Performance

- All new lots of reagents, kits, and control materials must be verified against current lots in use, prior to being placed into service.

  - This can be accomplished by the use of known QC materials. If known matrix effect problems exist with control material, then patient samples can be used to determine acceptability as well. A minimum of 3-5 patient samples should be used and the samples should include known normal and abnormal values.

  - "Lot-to-lot" testing should include known patient samples. In methods in which specimen acceptance criteria do not allow for extended storage, using patient samples for lot-to-lot verification may not be possible.

- An analytical run is defined as the interval within which the accuracy and precision of an instrument or method is expected to be stable, based upon manufacturer's recommendations and shall not exceed 24 hours.

  - Depending on the instrument or test, controls may be required every eight hours or once every 24 hours. See individual methods for QC instructions.

  - The manufacturer's instructions should be followed for the frequency of testing of external controls on test kits that utilize Internal controls.

- The material used for quality control may come with the test kit or be a commercially available preserved liquid product or a lyophilized product.

  - Control reagents in qualitative test kits are labeled with the expected result.

  - The commercial product may be assayed by the manufacturer or may be non-assayed.

    - Package insert ranges for assayed controls may be used only until there is sufficient data to calculate the mean and standard deviation for each analyte used.

    - Non-assayed controls may not be placed in service until the mean and standard deviation for each analyte has been determined.

- A number of factors will affect the stability of the instrument/reagent analytical system. These include:

  - Environmental factors such as temperature and humidity

  - Quality of water used by the system, or for reconstitution of controls and reagents

  - Recalibration of an existing lot of reagent

  - Change of reagent lot

  - Service to the analyzer

  - Deterioration of reagent or of control

  - Instrument mechanical wear or lamp problems

- Control reagents should be tested whenever there is a change during a run:

  - New reagents added

  - Instrument adjusted

  - Power interruption

  - Patient results are suspect

  - Internal control does not react as expected

- If reagents are not used according to instrument and/or kit manufacturer's instructions, one of the following applies:

  - Approval of the change must be sent by the manufacturer on the manufacturer’s letterhead and signed by a company representative authorized to approve the change.

  - There must be validation data to support that modifications produce reliable results consistent with the original assay claims of the manufacturer.

  - The Pathologist may approve a temporary variance. This must be documented in the Variance Approval Form (TDPL/GE/FOM-032).

Qualitative Methods

- Use of controls

  - Controls are run on reagents or test kits each day of use or with each run depending upon method, manufacturer’s instructions and compliance requirements.

  - The manufacturer's instructions should be followed for the frequency of testing of external controls on test kits that utilize Internal controls.

  - For external controls, at minimum, a positive and a negative control should be run.

  - Reactive, weakly reactive, and nonreactive controls are used where results are reported in this fashion.

  - Semi-quantitative (titer) procedures include a positive control, run at two or more levels, and a negative control.

  - Whenever possible, patient samples are included in lot-to-lot testing.

- Acceptable Control results

  - A "positive" control must be positive; a "negative" control must be negative.

  - Internal controls must react as described in the manufacturer's instructions.

  - Semi-quantitative (titer) procedure control results should agree ± 1 level.

  - Graded control results may be reported as +1, +2, etc. or Small, Moderate, Large, or as defined for the specific method; and must come within the required range.

  - Percent CV is within range set by manufacturer.

- For qualitative tests that use a cut-off value to distinguish positive from negative, the cut-off value should be verified initially, and every 6 months thereafter.

Quantitative Methods

- Every analyte tested will have at least 2 levels or values of controls run per analytical run.

- Acceptable limits for each new lot of controls are established by using laboratory specific values collected over a period of time, defined in the specific Department Plan, prior to being placed into service.

- Record daily control values into the LIMS or Daily IQC Data Monitoring Form (TDPL/GE/FOM-012) and determine standard deviation and range.

  - Compare data with the suggested acceptable limits or assay values from the manufacturer (or inter-laboratory peer group data, if available).

  - For unassayed controls, if there appears to be significant variation, run assayed control if available, or calibrators as unknown samples.

  - Based on these findings, determine mean and upper and lower limits of acceptability for control results.

  - Control limits or mean should be adjusted subsequently only after careful review of all data and instrument maintenance and lot numbers in use.

- Where appropriate, the control limits for moving averages based on patient results must be established including the following:

  - The frequency of calculations.

  - The basis for the selection of the upper and lower limits.

- Ensure: the process is appropriately sensitive to detect alterations in analyzer calibration status.

- Note: Recalibration is not required for minor calibration variations of no clinical consequence.

- Should there be an unusual change in coefficient of variation (CV) from one month to the next, or our CV is markedly smaller or larger than our interlaboratory group for a specific analyte/control, then an investigation must occur.

  - This can include running calibrators or assayed controls as unknowns, re-calibration, linearity check or receiving assistance from the instrument or QC manufacturer.

  - The investigation must be reviewed by the Lab Head or designee as soon as possible.

- Controls should be within ± 2 SD of the assigned mean

  - If two controls are used and both are within ±2 SD the run is acceptable

  - If three controls are used and all three are within ± 2 SD the run is acceptable.

  - If two controls are used and one is within ±2 SD and one is between 2-3 SD, the run is acceptable.

  - If three controls are used and two are within ± 2 SD and one is between 2-3 SD, the run is acceptable.

  - If one control is between 2-3 SD for two consecutive runs, the second run is not acceptable

  - The following would indicate a shift is occurring; Tech must investigate and notify manager or designee if QC cannot be resolved or if patient results are questionable.

    - Two or more results ± 3 SD

    - Two consecutive results more than 2 SD from the mean on the same side.

    - Of three consecutive results, two are greater than ± 2 SD from the mean on the same side.

- All QC failures and variances must have corrective action documented.

- Dilution protocols and diluents (or concentration protocols) are specified for all methods for which the Clinical Reportable Range exceeds the Analytical Measurement Range.

Post-Analytical Quality Control

- The Laboratory Head or Pathologist designee reviews and approves the content and format of patient reports annually.

- Laboratory Reports include the following elements:

  - Name and address of testing laboratory (see note below)

  - Patient name and identification number, or unique patient identifier

  - Name of the physician of record, or let all authorized person ordering test, as appropriate

  - Date and time of specimen collection, when appropriate

  - Date of release of report (if not on the report, this information should be readily accessible)

  - Time of release of report, if applicable (if not on the report, this information should be readily accessible)

  - Specimen source, when applicable

  - Test result(s) (and units of measurement, when applicable)

  - Reference intervals, as applicable (see Note below)

  - Conditions of specimen that may limit adequacy of testing.

- Each Laboratory Department has defined a system to detect and correct significant clerical and analytical errors, and unusual results.

  - The system provides for timely correction of results.

  - The system is documented.

- When results are revised, the revised and original data are clearly identified as such and readily accessible. Multiple sequential corrections of a single test result are referenced in sequential order on subsequent reports.

- Critical ranges are defined for test results when appropriate

  - Notification of critical values is defined and documented.

  - There is a policy for verification "read-back" of critical values that are communicated verbally or by phone.

Calibration Verification

- Calibration Verification is the process of confirming that the current calibration settings remain valid for a method.

  - Some calibrations require a series of known standards; at least three calibrators that span the reportable range should be run as "unknowns".

  - Some calibrations may require only one-point calibration due to the instrument methodology.

  - Controls should be included in the run.

  - If calibration verification confirms the settings are valid, it is not necessary to perform a full calibration or recalibration of the method.

- Indications for Calibration

  - As specified by equipment/reagent manufacturer.

  - New lot of reagents, complete change of reagents, or new reagent manufacturer being placed into use; depending on requirements for specific method.

  - Unacceptable QC results from 2 different bottles of the same control when apparent instrument problem exists.

  - Failure of calibration/ verification procedure.

  - At least every 6 months.

  - Following major equipment maintenance or major component replacement.

- Acceptable Results

  - A calibration is successful when QC is acceptable and when calibrators run as unknowns are within acceptable limits.

  - The recovered values should be within ± 10% of the stated values.

  - All documentation of calibration must be kept in appropriate logs.

  - A failure of this process will require investigation and recalibration before any patient specimens can be analyzed.

Analytical Measurement Range Validation

- The Analytical Measurement Range (AMR) is the range of analyte values that a method can directly measure without any dilution, concentration or other pretreatment, not part of the normal assay process.

  - AMR validation is the process of confirming that the assay system will correctly recover the concentration or activity of the analyte over the established AMR.

  - Materials used for validation must be known to have the appropriate matrix characteristics for the specific method. In many cases the manufacturer will recommend suitable materials.

  - Test specimens must have values that, at a minimum, are near the low, mid, and high values of the AMR.

- AMR validation must be performed at least once every 6 months.

- If the materials used for Calibration or calibration verification include low, midpoint and high values that are near the stated AMR, and if the data are within the acceptance criteria, the AMR has been verified and no additional procedures are required.

Tools for data collection and analysis

  - Westgard Rules

  - Levey-Jennings Chart

    - Used when predicting the expected range of outcomes from a process.

    - Used when determining whether or not a process is stable or in statistical control

    - shows a measurement against time and is used to look for trends or changes in the average

  - Inter-Laboratory Cumulative Data

  - Evaluation Protocols (EP Series), CLSI

Corrective Action

- Any control value outside of acceptable limits must be investigated and resolved. Patient results must be held back, and no further patients tested, during the investigation.

- Immediate action

  - Rerun the control with fresh aliquot

  - If another vial of the control is available, use that. Otherwise prepare a fresh vial.

  - No more than 2 control reruns should be made.

- Do not continue to rerun controls until an in-control result is obtained.

- Troubleshooting

  - Check the analyzer (or kit) for problems.

  - Has the control been in the analyzer for too long? Is the control beyond the open vial expiration? In either case, pour a fresh cup or reconstitute a new vial as indicated

  - Has the reagent expired, or been open for longer than the manufacturer's recommendations? Replace and rerun.

  - Has the elapsed time of calibration exceeded manufacturer's recommendations? Recalibrate.

  - Has a lot or reagent been changed, with or without calibration? Calibrate.

  - Has there been any maintenance done on the analyzer or other equipment? May need calibration and/or reassignment of control mean.

  - Has there been a recent calibration? May need to recalibrate.

- Patient results

  - No patient results are to be reported until the out-of-control situation has been resolved.

  - If the investigation indicates that the run was not in control:

    - Recall all patient results back to the last valid control results.

    - Repeat all patient testing once problem is resolved.

- All instances of control problems must be documented, along with the corrective (remedial) action, in the Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021).

- If QC results were unacceptable, not performed, or not documented, and patient results were released, a lab incident report must be generated.

  - Incident reports are managed as per the specifications in Laboratory Incident Form (TDPL/GE/FOM-018).

  - QA reviews the report and works with the Lab Manager/Department Supervisor to determine the need for Process Improvement.

Reviews

  - The Technical staff performing the test verifies that the results of control runs are within acceptable limits before patient-results are reported.

  - Weekly Review: The manager or designate reviews the Quality Control results.

    - Review for correct entry of data.

    - Review for adequate documentation of quality control failures and corrective actions.

    - Review for trends and shifts.

    - The reviewer initials and dates the appropriate form(s), or is identified electronically, to document that the review was performed.

    - Document any additional corrective actions.

  - Monthly Review: The Quality Manager, Department Head, or other appropriate supervisory staff, reviews Quality Control charts, logs, and/or forms.

    - Review results for acceptability

    - Review documented corrective actions for appropriateness and completeness.

    - Review records for shifts, trends and bias.

    - Investigate variances, take corrective action, and document.

    - The reviewer initials and dates the appropriate form(s) to document that the review was performed.

  - Quality Control data, printouts, forms and logs will be retained for two years.

- NOTE: Any exceptions to or variance from the policies in this document must be approved by the pathologist and documented in the Variance Approval form (TDPL/GE/FOM-049).

Quality Control

- The Quality Manager or designee will do random Quality Control checks to ensure that this SOP is implemented as written. If Quality control checks reveal non-compliance with this SOP, the SOP deviations and violations will be investigated, corrective and preventive action initiated, implemented, reviewed and closed out. All corrective and preventive action should be documented in the Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021).

Related Documents

- N/A.

Reference Regulations & Guidelines

Specific Criteria for Accreditation of Medical Laboratories (NABL 112 A; Issue 01)

Guidance document: Medical Laboratories (NABL 112B; Issue 01)

ISO 15189:2022 Medical Laboratories – Requirements for Quality and Competence

Related Records

Material Safety Data Sheet (MSDS) Folder (TDPL/GE/FOL-003)

Daily IQC Data Monitoring Form (TDPL/GE/FOM-012)

Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021)

Laboratory Incident Form (TDPL/GE/FOM-018)

Variance Approval form (TDPL/GE/FOM-032)