SOP – 004

PROCEDURE FOR ALBUMIN TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	8

7.	Type of Sample	9

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	10

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	11

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	13

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | BCG | Bromocresol Green |
| 4 | °C | Degree Centigrade |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EDTA | Ethylenediaminetetraacetic Acid |
| 8 | EQAS | External Quality Assurance Scheme |
| 9 | HOD | Head of Department |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | pH | Potential of Hydrogen |
| 14 | QC | Quality Control |
| 15 | QM | Quality Manager |
| 16 | QNS | Quantity Not Sufficient |
| 17 | SD | Standard Deviation |
| 18 | SOP | Standard Operating Procedure |
| 19 | SST | Serum Separator Tube |
| 20 | TRF | Test Requisition Form |
| 21 | TDPL | TRUSTlab Diagnostics Private Limited |
| 22 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for quantitative estimation of Albumin in serum and plasma samples using Beckman coulter automated biochemistry analyser. This SOP gives information about Albumin estimation, clinical significance etc.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- An albumin serum test measures the amount of albumin, a protein produced by the liver, in your blood. It's a blood test that can help assess liver and kidney function, as well as nutritional status.

- Test Principle

Albumin binds with bromocresol green (BCG) to form a coloured complex, at pH 4.2 causing a shift in absorbance of the yellow BCG dye. The blue–green colour formed is proportional to the concentration of albumin present, which is measured bichromatically (600/800 nm) with a maximum absorbance at 625 nm.

pH 4.2

Albumin + Bromocresol green (BCG) 	Green complex

- Specific Performance Characteristics

Limit of Quantitation:

      - LLOQ: 1.5 g/dL

      - ULOQ: 6.0 g/dl

Analytical measurement range (AMR): 1.5 – 6.0 g/dl

Clinical Reportable Range (CRR):  3.5 – 5.2 g/dl

- Type of Sample

A minimum of 2ml of Serum/Plasma

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Collection in wrong vacutainer

      - No TRF/Grossly incomplete TRF/Errors in labeling

      - Grossly haemolysed specimens.

      - Highly lipemic specimens.

      - Highly icteric specimens.

      - Sample beyond stability period

      - Incomplete TRF details or labelling errors are clarified with PSC/ Doctor/ Patient before sample acceptance / processing. If critical information is still not obtained, the sample is rejected

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

EDTA- Purple top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No other specific preparation required.

- Equipment & Reagents

Equipment

- Beckman DxC 700 AU Chemistry Analyzer

Reagents

      - Final concentration of reactive ingredients:

- Succinate buffer (pH 4.2)                             100 mmol/l

- Bromocresol green                                       0.2 mmol/l

- Preservative

      - The concentrations of the reactive components of the reagents shown on the kit label are the actual concentrations in the individual R1/R2 vials. The reagent composition which is shown in the Instructions for Use is the final concentration of these components in the reaction cuvette after addition of R1, Sample, and R2.

      - All reagents are ready to use.

      - Storage & Stability:

- The reagent is stable, unopened, up to the stated expiry date when stored at 2 to 25°C. Once open, reagent stored on board the instrument is stable for 90 days.

- Safety

For in vitro diagnostic use only. Do not pipette by mouth. Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 10% Sodium hydroxide.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

- Calibration Procedure

Calibrator Required: System Calibrator

- The calibrator albumin value is traceable to IFCC (International Federation of Clinical Chemistry) standard CRM 470.

When to Calibrate

- Recalibrate the assay when any of the following occur:

      - Change in reagent lot number or significant shift in control values.

      - Major preventative maintenance was performed on the analyser or a critical part was replaced.

      - Once every 30 days using the same reagent lot or after 7 days if the reagents are onboard.

- Test Procedure

Refer to the appropriate Beckman Coulter AU analyzer User Guide.

Sample Processing Using Racks:

      - Place Each sample in a rack (if volume is not sufficient transfer sample into a validated low volume container).

      - Open the top and/or the front cover of the Rack Input area.

      - Place the Rack in the input area with the Rack ID barcode facing the back of the analyzer.

      - Do not leave any empty space between the racks.

      - Close the rack Input area cover.

      - Select the Start button.

      - To monitor sample status on the Analyzer, select results.

- Quality Control Procedure

Internal Quality Control

    - Biorad. Two levels of Quality Control (L1 & L2) sera i.e. normal and high are run once daily (as per NABL 112).

    - Preparation of Controls:

      - Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

      - Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

      - Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

      - Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

      - Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided).

      - Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”. Store the aliquots at -20°C.

    - Storage and Stability of Controls:

      - Unopened: Stable until the expiry date when stored at 2 to 8°C.

      - Reconstituted: Stable for 7 days when stored at 2 to 8°C; stable up to 30 days when stored at -20°C. Note: Aliquots are stored at the above temperatures as required.

    - Procedure

      - Run Quality Control according to TRUS lab-defined Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20°C, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8°C, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion

      - Place a portion of the material in a sample cup (approximately 350 µL).

      - Place the cup in the tray positions indicated by the screen Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levy Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

      - Verify the following points and if required, consult with HOD/QM:

          - Review the environmental conditions and recent events and look for any deviation.

          - Check on-board reagents (for stability, deposits, fungal growth) reference fluid and controls.

          - Check the quality of the sterile water for injections used for the preparation of controls and calibrators, contaminants, i.e. bacterial growth may contribute to inaccurate results.

          - Check reaction temperature.

          - Check the expiry date of kit and contents.

          - Check instrument settings and light source.

          - If there are any control failures, take appropriate corrective and preventive actions as per TRUSTlab’s Quality Control MSP.

          - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

          - Take appropriate corrective actions and re-run the QC.

          - Contact application support/service support from vendor if required corrective action fails to yield acceptable results.

    - Repeat and Recheck criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation – CMC EQAS

- Interferences & Cross Reactions

Lysed serum specimens may give falsely elevated values.

Do not use if the re-agents is turbid as it indicates contamination of the reagent.

Haemoglobin in the tested level of 150 mg/dL has no significant interference on the assay.

Lipemia in the tested level of 2000 mg/dl has no significant interference on the assay.

Bilirubin in the tested level of 40 mg/dl has no significant interference on the assay.

Ascorbic Acid in the tested level of 30 mg/dL has no significant interference on the assay.

- Result Calculation

TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

| Age | Male (g/dl) | Female (g/dl) |
| --- | --- | --- |
| 1 day | 2.8-4.4 | 2.8-4.4 |
| 1day - 14 years | 3.8-5.4 | 3.8-5.4 |
| 14 – 60 years | 3.5-5.2 | 3.5-5.2 |
| 60 - 90 years | 3.2-4.6 | 3.2-4.6 |
| > 90 years | 2.9-4.5 | 2.9-4.5 |

- Reportable Interval

Analytical Measurement Range (AMR):  1.5 – 6.0 g/dl

Clinical Reportable Range (CRR):  3.5 – 5.2 g/dl

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

      - In children: < 1.7 g/dl and > 6.8 g/dl.

- Laboratory Clinical Interpretation

Albumin is the most abundant serum protein in blood plasma, representing 55-65% of the total protein. It is synthesised in the liver and has a half-life of 2 to 3 weeks.

The main biological functions of albumin are to maintain the water balance in serum and plasma and to transport and store a wide variety of ligands e.g. fatty acids, calcium, bilirubin and hormones such as thyroxine.  Albumin also provides an endogenous source of amino acids.

Increased absolute serum albumin content is not seen as a natural condition. Relative increase may occur in hemoconcentration. Absolute increase may occur artificially by infusion of hyperoncotic albumin suspensions.

Decreased serum albumin is seen in states of decreased synthesis (malnutrition, malabsorption, liver disease, and other chronic diseases), increased loss (nephrotic syndrome, many GI conditions, thermal burns, etc.), and increased catabolism (thyrotoxicosis, cancer chemotherapy, Cushing's disease, familial hypoproteinemia.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid DO NOT USE as it indicates contamination of the reagent.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are strongly recommended for the test.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of Beckman Coulter AU680 / AU 480 / DXC 700

      - Procedure for Managing Internal QC Program (TDPL/AQ/MAN-001: MSP-013)