SOP – 007

PROCEDURE FOR AFP TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	11

13.	Test Procedure	11

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	13

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	16

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AFP | Alpha-Fetoprotein |
| 2 | AMR | Analytical Measurement Range |
| 3 | BC | Biochemistry |
| 4 | BSA | Bovine Serum Albumin |
| 5 | °C | Degree Centigrade |
| 6 | CA | Carbohydrate Antigen |
| 7 | CRR | Clinical Reportable Range |
| 8 | HOD | Head of Department |
| 9 | LLOQ | Lower Limit of Quantitation |
| 10 | EDTA | Ethylenediaminetetraacetic Acid |
| 11 | EQAS | External Quality Assurance Scheme |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | SST | Serum Separator Tube |
| 15 | TDPL | TRUSTlab Diagnostics Private Limited |
| 16 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of AFP in serum samples using Beckman DxI 800 analyzer.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- An alpha-fetoprotein (AFP) test measures the amount of AFP in your blood. It can be used to screen for birth defects in pregnancy, or to diagnose certain cancers in adults.

- Test Principle

- The Access AFP assay is a two-site immunoenzymatic (“sandwich”) assay. A sample is added to a reaction vessel with mouse monoclonal anti-AFP-alkaline phosphatase conjugate, and paramagnetic particles coated with a second mouse monoclonal anti-AFP antibody. The AFP in the sample binds to the immobilized monoclonal anti-AFP on the solid phase while, at the same time, the monoclonal anti-AFP-alkaline phosphatase conjugate reacts with different antigenic sites on the sample AFP. After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is directly proportional to the concentration of AFP in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.5 ng/ml

      - ULOQ: 3000 ng/ml

Analytical measurement range (AMR): 0.5–3000 ng/ml

Clinical Reportable Range (CRR): >300000 (1:100 dilution)

Sensitivity: 95%

Specificity: NA

- Note: If requested by treating Clinician/Doctor, values greater than CRR may be reported, with approval of HOD.

- Type of Sample

Only non hemolysed serum/plasma

      - Ambient (+15 to 25°C): 24 hrs

      - Refrigerator (+2 to +8°C): 4 days

      - Frozen (-20°C): long term storage.

- Type of Container & Additives

- Red/SST yellow topped vacutainer tubes. Do not add any additives /preservatives.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No other specific preparation required.

- Equipment & Reagents

Instruments

      - DxIi 800

Material

      - R1: Access AFP Reagent Pack

          - R1a: Paramagnetic particles coated with mouse monoclonal anti-AFP antibodies suspended in TRIS buffered saline, with surfactant, bovine serum albumin (BSA) matrix, < 0.1% sodium azide, and 0.1% ProClin 300.

          - R1b: Mouse monoclonal anti-AFP alkaline phosphatase (bovine) conjugate diluted in phosphate buffered saline, with surfactant, BSA matrix, proteins (goat, rabbit, mouse), < 0.1% sodium azide, and 0.25% ProClin 300.

      - S0-S6:	Access AFP Calibrators: 2.5 mL/ vial

      - Access Substrate

- Cat. No. 81906: 4 x 130 mL

      - DxI Wash Buffer II

      - Access AFP Sample Diluent.

Reagent Storage & Stability

      - Reagent Kit must be stored at 2-8°C and may be used immediately after removal from 2-8°C storage.

      - When stored and handled as directed, reagents are stable until the expiration date.

      - Reagent Kit may be stored on board the DxI System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking on board time, refer to the DXI System Operations Manual,

      - Reagents may be stored on or off the DxI System. If reagents are removed from the system, store them at 2-8°C (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright.

      - If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded.

      - After reagents are removed from the system, you must initiate a scan to update the onboard stability timer

- Safety

Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the AFP Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

- Calibration Procedure

- An active calibration curve is required for all tests. For the Access AFP assay, calibration is required every 28 days. The Access AFP Calibrators are provided at seven levels - zero and approximately 2.5, 5.0, 25, 100, 500 and 3000 ng/mL- prepared gravimetrically from human AFP and buffered BSA matrix. Assay calibration data are valid up to 28 days.

- Test Procedure

- Refer to the appropriate system manuals and/or Help system for information on managing samples, configuring tests, requesting tests, and reviewing test results.

Specimen Dilution Procedures:

      - For assaying samples containing <3000 ng/mL AFP, select AFP as the test name.

      - Select Dil-AFP as the test name for assaying samples containing > 3000 ng/mL

      - For detailed information on ordering dilutions, refer to the DXI System Operations Manual, Section 5.

Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure

      - To handle Customer request/complaints upon approval by HOD

Handling of Repeat Test Results:

      - In case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever higher).

      - If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

    - Biorad- Three levels of Quality Control sera i.e. low, normal and high range, any two levels are run once daily

      - Run QC according to TRUSTlab’s Quality Control described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC storage, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of material in a sample cup (approximately 350 µL).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal QC once at the beginning of the day.

      - Check QC values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levy Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible

    - Handling of QC failure:

      - Verify the following points and if required, consult with HOD/QM:

          - Review the environmental conditions and recent events and look for any deviation.

          - Check on-board reagents (for stability, deposits, fungal growth) reference fluid and controls.

          - Check the quality of the sterile water for injections used for the preparation of controls and calibrators, contaminants, i.e. bacterial growth may contribute to inaccurate results.

          - Check reaction temperature.

          - Check the expiry date of kit and contents.

          - Check instrument settings and light source.

          - If there are any control failures, take appropriate corrective and preventive actions as per TRUSTlab’s Quality Process MSP.

          - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032)

          - Take appropriate corrective actions and re-run the QC.

          - Contact application support/service support from vendor if required corrective action fails to yield acceptable results.

    - Storage and Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2 to 8 °C.

      - Reconstituted: Stable for 5 days when stored at 2 to 8º C.

      - Longer storage: -20°C.

External Quality Control

- EQAS participation – Metropolis EQAS.

- Interferences & Cross Reactions

Samples containing up to 25 mg/dL (427 µmol/L) bilirubin, lipemic samples containing the equivalent of 520 mg/dL triglycerides, 560 mg/dL phospholipid, 380 µEq/L free fatty acid, and hemolyzed samples containing up to 1200 mg/dL (12 g/L) hemoglobin do not affect the concentration of AFP assayed. In addition, samples with 0.6 g/dL (6 g/L) of bovine albumin added to the endogenous albumin in the samples do not affect the concentration of AFP assayed.

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

- 0.0 to 9.0 ng/mL

- Reportable Interval

- 0.5 – 3000 ng/mL

- CRR > 3000 ng/mL

- Dilution Procedure:

- For UniCel DxI systems:

- Samples containing > 3000 ng/mL can be processed using the DxI onboard dilution feature. The DxI system onboard dilution feature automates the dilution process, using one volume of sample with 100 volumes of UniCel DxI Access Immunoassay Systems Wash Buffer II, allowing samples to be quantitated up to 303,000 ng/mL. The system reports the results adjusted for the dilution.

- Alternatively, samples containing > 3000 ng/mL for AFP or > 51,000 ng/mL for Dil-AFP can also be processed via off-line pre-dilution following these examples:

  - For AFP values > 3000 ng/mL, dilute as follows:

  - Dilute one volume of sample with 100 volumes of Access Wash Buffer II or Access AFP Sample Diluent for serum samples. The pre-dilution factor is 101.

- For Dil-AFP values > 51,000 ng/mL, dilute as follows:

  - For serum samples, dilute one volume of sample with 10 volumes of Access Wash Buffer II or Access AFP Sample Diluent. The pre-dilution factor is 11.

  - Type in the pre-dilution factor when entering the test request. Order the AFP or Dil-AFP test. The system will automatically multiply the result by the pre-dilution factor and report that value.

  - If the pre-dilution factor is not used when entering the request, multiply the calculated value by the pre-dilution factor after assaying the diluted sample using the Access AFP assay or the Dil-AFP assay.

  - If the system reports a pre-diluted AFP result as < 0.50 ng/mL or a pre-diluted Dil-AFP result < 2700 ng/mL, then re-dilute with a lesser dilution.

- Refer to the appropriate system manuals and/or Help system for additional instructions on processing pre-diluted samples.

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

      - N/A.

- Laboratory Clinical Interpretation

AFP is first detected in the fetal circulation30 days after conception. After reaching a peak concentration at 13 weeks gestation, levels gradually diminish until birth. By 2 years of age, only trace levels of AFP can be detected in normal individuals.

Elevated AFP levels are found in primary hepatic carcinoma and other malignancies most importantly nonseminomatous testicular carcinoma.

Elevated levels of AFP in nonseminomatous testicular carcinoma greatly facilitated the differential diagnosis of germ cell tumors, since pure seminoma is not associated with elevated AFP levels. Changing AFP levels have assisted in the diagnosis, prognosis and management of patients with nonseminomatous testicular carcinoma.

Elevated AFP levels have also been found in association with ataxia telangiectasia, hereditary tyrosinemia, neonatal hyper bilirubinemia, acute and chronic viral hepatitis, cirrhosis, and other malignancies.

Additionally, decreasing levels following therapy generally indicate successful intervention whereas rising levels following therapy usually indicate residual tumor or recurrence.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - Do not use if the reagents are turbid as it indicates contamination of the reagent.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

      - Strong lipemic and haemolytic sera should not be used.

- References

Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

Access DxI AFP Kit Insert

Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).