SOP – 009

PROCEDURE FOR

BETA HCG TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	9

8.	Type of Container & Additives	9

9.	Patient Preparation	10

10.	Equipment & Reagents	10

11.	Safety	11

12.	Calibration Procedure	11

13.	Test Procedure	12

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	15

17.	Biological Reference Interval	16

18.	Reportable Interval	16

19.	Procedure when Results are not in Range	16

20.	Critical Alert Values	18

21.	Laboratory Clinical Interpretation	18

22.	Potential Source of Variability	18

23.	References	19

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | ꞵ-hCG | Beta Human Chorionic Gonadotropin |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EDTA | Ethylenediaminetetraacetic Acid |
| 8 | EQAS | External Quality Assurance Scheme |
| 9 | FSH | Follicle Stimulating Hormone |
| 10 | HOD | Head of Department |
| 11 | IFCC | International Federation for Clinical Chemistry |
| 12 | IQC | Internal Quality Control |
| 13 | LH | Luteinizing Hormone |
| 14 | LLOQ | Lower Limit of Quantitation |
| 15 | MSP | Management System Procedure |
| 16 | N/A | Not Applicable/Available |
| 17 | pH | Potential of Hydrogen |
| 18 | QC | Quality Control |
| 19 | QM | Quality Manager |
| 20 | QNS | Quantity Not Sufficient |
| 21 | SD | Standard Deviation |
| 22 | SOP | Standard Operating Procedure |
| 23 | SST | Serum Separator Tube |
| 24 | TRF | Test Requisition Form |
| 25 | TDPL | TRUSTlab Diagnostics Private Limited |
| 26 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of total Beta-hCG in serum using Beckman Coulter DxI 800. This SOP also provides information about the test principle, clinical significance and other test specific details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- The beta-hCG test is a blood test that measures the level of human chorionic gonadotropin (hCG) in the blood. hCG is a hormone produced by the placenta during pregnancy. This test is primarily used to confirm pregnancy, monitor its progress, and detect potential complications. It can also be used to diagnose ectopic pregnancies, miscarriages, and certain cancers.

- Test Principle

The Access Total β-hCG assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of total β-hCG levels in human serum and plasma using the Access Immunoassay Systems.

Human chorionic gonadotropin (hCG) is a glycoprotein hormone, produced by the placenta, with structural similarity to the pituitary hormones FSH, TSH, and LH. The alpha subunit (MW 15,000-20,000 daltons) is common to all of these hormones, but the beta subunits differ and confer immunological and biological specificity. Beta hCG (MW 25,000-30,000 daltons) shares several peptide sequences with beta LH but has a unique carboxyl terminal region.

Shortly after the implantation of a fertilized ovum into the uterine wall, the trophoblast begins to produce hCG. The hormone maintains steroid secretions of the corpus luteum until the placenta can do so. During a normal pregnancy, hCG is generally approximately 50 mIU/ml (IU/l) in the week after conception and doubles every 1.5 - 3 days for the first six weeks. Levels continue to rise until the end of the first trimester, then gradually fall to a lower level for the remainder of the pregnancy. After delivery, hCG returns to < 5 mIU/ml (IU/l) and is usually undetectable for several days postpartum.

This hormone is an excellent marker for pregnancy. Healthy, non-pregnant individuals have low (< 5 mIU/ml [IU/l]) to undetectable hCG levels. During pregnancy, hCG concentrations increase as noted above and then show a gradual decrease after the first trimester. Unusually low or rapidly declining levels may indicate an abnormal condition such as an ectopic pregnancy or impending spontaneous abortion.

Originally bioassay systems measured hCG by measuring gonadal tissue response in various animals. These methods exhibited insufficient sensitivity, were difficult to perform, and required large volumes of sample. Tests to measure urine hCG traditionally employed latex agglutination or agglutination inhibition methods. With the development of radioimmunoassay techniques for the measurement of hCG by Vaitukaitis et al. in 1972, more sensitive and rapid assays for hCG became available. Subsequent development of two-site immuno-radiometric assays (IRMA) provided assays with increased sensitivity, specificity and precision.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation

      - LLOQ: 0.5 mIU/ml

      - ULOQ: 1000 mIU/ml

Analytical measurement range (AMR): 0.5-1000 mIU/ml

Clinical Reportable Range (CRR): 2.39–2,00,000 mIU/ml

The DxI system onboard dilution feature automates the dilution process, using one volume of sample with 199 volumes of UniCel DxI Access Immunoassay Systems Wash Buffer II, allowing samples to be quantitated up to 2,00,000 mIU/ml.

- Type of Sample

A minimum of 2 ml of Serum/Plasma

Collect blood sample by venipuncture

For serum, allow samples to clot adequately before centrifugation.

Keep tubes stoppered at all times.

Physically separate serum or plasma from contact with cells as soon as possible.

Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - •	Follow blood collection tube manufacturer’s recommendations for centrifugation

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800

Reagents

    - Access Total βhCG Reagent Pack Cat. No. 33500: 100 determinations, 2 packs, 50 tests/pack

    - Provided ready to use.

    - Store upright and refrigerate at 2 to 10°C.

    - Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

    - Stable until the expiration date stated on the label when stored at 2 to 10°C.

    - Stable at 2 to 10°C for 28 days after initial use.

    - Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

    - If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

    - All antisera are polyclonal unless otherwise indicated.

    - R1a: Paramagnetic particles coated with goat anti-mouse IgG: mouse monoclonal anti-βhCG complexes suspended in TRIS buffered saline, with surfactant, bovine serum albumin (BSA),

    - < 0.1% sodium azide, and 0.1% ProClin 300.

    - R1b: Rabbit anti-βhCG alkaline phosphatase (bovine) conjugate diluted in TRIS buffered saline, with surfactant, BSA, protein (goat, rabbit, mouse), < 0.1% sodium azide, and 0.25% ProClin 300.

    - CAUTION:

- Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

    - Do not use reagents beyond the expiration date

- Safety

For in vitro diagnostic use only.

Do not pipette by mouth.

Do not ingest. Harmful if swallowed

Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

Used cuvettes contain human body fluids, it should be handled with appropriate care to avoid skin contact.

Reagent packs: Sealed cartridges are stable until the Expiry date.

- Calibration Procedure

Calibrator Required:

- N/A.

Calibration Schedule:

- An active calibration curve is required for all tests. Calibration is required every 28 days. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When quality control results are found to be consistently outside acceptable range.

- Test Procedure

Refer to the appropriate Beckman Coulter DxI 800 User Guide.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use ten (10) μL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) μL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

The system default unit of measure for sample results is ng/ml. To change sample reporting units to the International System of Units (SI units), µg/l, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply ng/ml by multiplication factor

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Analyzer generated results are reviewed by the operator & only clinically valid results are saved for approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad - Level 1, Level 2 and Level 3; Any two levels of Quality Control sera i.e. low, normal and high as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

Samples containing up to 10 mg/dl (171 μmol/l) bilirubin, lipemic samples containing the equivalent of 1800 mg/dl (20.32 mmol/l) triglycerides, and hemolyzed samples containing up to 500 mg/dl (5 g/l) hemoglobin do not affect the concentration of total βhCG assayed. In addition, samples with 3 g/dl (30 g/l) of human albumin added to the endogenous albumin in the samples do not affect the concentration of total βhCG assayed.

No significant cross reactivity (< 0.5 mIU/ml [IU/l]) was observed when hLH, hFSH, or hTSH were added to the Access Total βhCG Calibrator S0 (zero) at 1000 mIU/ml (IU/l), 1000 mIU/ml (IU/l), 1 mIU/ml (IU/l), respectively.

- Result Calculation

Each laboratory should establish its own reference ranges to assure proper representation of specific populations.

Total βhCG concentrations were measured in human serum samples collected from apparently healthy adult males and non-pregnant females using the Access Total βhCG assay. The observed range of total βhCG concentrations are shown in the following table.

| Reference Group | n | Median (mIU/ml) | 95% Range (mIU/ml) |
| --- | --- | --- | --- |
| Males | 250 | < 0.5 | < 0.5 - 2.67 |
| Non-pregnant females | 250 | < 0.5 | < 0.5 - 2.90 |

Representative hCG ranges during normal pregnancy are summarized below. Because other clinical reference citations may show different values, each laboratory should establish its own reference ranges.

| Approximate Gestational Age (weeks) | Approximate hCG Range (mIU/ml, IU/l) | Approximate Gestational Age (weeks) | Approximate hCG Range (mIU/ml, IU/l) |
| --- | --- | --- | --- |
| 0.2-1 | 5-50 | 4-5 | 1,000-50,000 |
| 1-2 | 50-500 | 5-6 | 10,000-100,000 |
| 2-3 | 100-5,000 | 6-8 | 15,000-200,000 |
| 3-4 | 500-10,000 | 8-12 | 10,000-100,000 |

- Biological Reference Interval

| Females - Age | Range |
| --- | --- |
| 18 - 40 Yrs | 0.0 - 0.6 mIU/ml |
| > 40 Yrs | 0.0 – 3.1 mIU/ml |
| Post-menopausal | 0.1 -11.6 mIU/ml |

- Reportable Interval

Analytical Measurement Range (AMR): 0.5 - 1000 mIU/ml

Clinical Reportable Range (CRR): 0.5 - 200000 mIU/ml

- Procedure when Results are not in Range

This assay is intended for early detection of pregnancy.

Samples can be accurately measured within the analytic range of the lower limit of detection and the highest calibrator value (approximately 0.5-1000 mIU/ml [IU/l]).

To accurately measure samples containing 1000-200,000 mIU/ml, select the Dil-hCG2 test. This test uses the TβhCG2 pack. When Dil-hCG2 is requested, the system autodilutes the sample and reads the resulting dose off the TβhCG2 calibration curve. The system multiplies by the dilution factor defined in the software (200) to calculate final test results.

For UniCel DxI Systems:

      - Samples containing > 1000 mIU/ml, can be processed using the DxI onboard dilution feature.

      - The DxI system onboard dilution feature automates the dilution process, using one volume of sample with 199 volumes of UniCel DxI Access Immunoassay Systems Wash Buffer II, allowing samples to be quantitated up to 200,000 mIU/ml.

      - The system reports the results adjusted for the dilution.

      - If the sample contains less than 800 mIU/ml, the system will report results as < 800 mIU/ml.

      - Samples containing > 1000 mIU/ml can also be processed via off-line pre-dilution following these steps:

      - Dilute one volume of sample with 199 volumes of Wash Buffer (1/200) or dilute per laboratory dilution protocol. NOTE: Dilution with alternate buffers may cause erroneous results.

      - Type in the pre-dilution factor when entering the test request. Order the TβhCG or Dil-hCG2 test.

      - The system will automatically multiply the result by the pre-dilution factor and report that value. NOTE: If the system reports a pre-diluted Dil-hCG2 result as < 1000 mIU/ml (IU/l), redilute the sample such that it will read between 1,000 and 200,000 mIU/ml (IU/l). A neat sample reading < 1000 mIU/ml (IU/l) in the Dil-hCG2 assay should be retested in the TβhCG2 assay.

      - If the pre-dilution option is not selected, multiply the calculated value by the dilution factor 200 (or by another selected factor) after assaying the diluted sample using the Access Total βhCG assay.

      - If the calculated value of the diluted sample using the Access Total βhCG assay is < 5 mIU/ml (IU/l), re-dilute one volume of the neat sample with 99 volumes of Wash Buffer (1/100) and re-assay, remembering to multiply the calculated value by the dilution factor 100.

      - Refer to the appropriate system manuals and/or Help system for additional instructions on processing pre-diluted samples

DO NOT reuse small sample volumes that have been resident on the analyzer for more than 1 hour.

The Access Total βhCG assay has no discernible “hook effect” at 1,000,000 mIU/ml.

For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce antibodies, e.g. HAMA, that interfere with immunoassays. Additionally, other heterophile antibodies such as human anti-goat antibodies may be present in patient samples.

Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

The role of preanalytical factors in laboratory testing has been described in a variety of published literature. Following blood collection tube manufacturers’ specimen collection and handling recommendations will help to reduce preanalytical errors.

The Access Total βhCG results should be interpreted in light of the total clinical presentation of the patient, including symptoms, clinical history, data from additional tests, and other appropriate information.

If the total βhCG level is inconsistent with clinical presentation, results should be confirmed by an alternate hCG method or a urine-based assay.

Trophoblastic or non-trophoblastic neoplastic conditions should be ruled out before reporting results.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

NEGATIVE:

- A single HCG reading is not enough information for most diagnosis. Whenever there is a question regarding the health of the pregnancy, multiple testing of hCG done a couple of days apart gives a more accurate assessment of the situation.

POSITIVE:

- Repeat after 72 hrs. Values in pregnancy should double every 3 days for the first 6 weeks.

Decreasing HCG concentrations indicate threatened or missed abortion, recent termination of pregnancy, ectopic pregnancy, gestosis or intrauterine death.

Both normal and ectopic pregnancies generally yield positive results of pregnancy tests. The comparison of quantitative HCG measurements with the results of transvaginal ultrasonography (TVUS) may aid in the diagnosis of ectopic pregnancy.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of βhCG, Beckman Coulter DxI 800

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)