SOP – 014

PROCEDURE FOR

CA 125 TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	11

13.	Test Procedure	11

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	15

17.	Biological Reference Interval	16

18.	Reportable Interval	16

19.	Procedure when Results are not in Range	16

20.	Critical Alert Values	17

21.	Laboratory Clinical Interpretation	17

22.	Potential Source of Variability	18

23.	References	18

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CA 125 | Cancer Antigen 125 |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EDTA | Ethylenediaminetetraacetic Acid |
| 8 | EQAS | External Quality Assurance Scheme |
| 9 | HOD | Head of Department |
| 10 | IFCC | International Federation for Clinical Chemistry |
| 11 | IQC | Internal Quality Control |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MSP | Management System Procedure |
| 14 | N/A | Not Applicable/Available |
| 15 | pH | Potential of Hydrogen |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | SD | Standard Deviation |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TRF | Test Requisition Form |
| 23 | TDPL | TRUSTlab Diagnostics Private Limited |
| 24 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Cancer Antigen 125 (CA-125) in serum using Beckman Coulter DxI 800. This SOP also provides information about the test principle, clinical significance and other test specific details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- CA 125 is a blood test that measures the amount of a protein called cancer antigen 125 in the blood. It's a tumor marker, and elevated levels can be a sign of ovarian cancer, but also other conditions.

- Test Principle

The Access OV Monitor assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of CA 125 antigen levels in human serum and plasma using the Access Immunoassay Systems. This device is indicated for use in the measurement of CA 125 antigen to aid in the management of ovarian cancer patients. Serial testing for patient CA 125 antigen concentrations should be used in conjunction with other clinical methods used for monitoring ovarian cancer.

The CA 125 antigen is an epitope on a large mucin-like glycoprotein (MW ~ 1000 kDa) that may be found in elevated concentrations in certain ovarian malignancies. There is no known function for CA 125.

The Access OV Monitor assay is a two-site immunoenzymatic (“sandwich”) assay. A sample is added to a reaction vessel along with mouse monoclonal anti-CA 125 antigen alkaline phosphatase conjugate and paramagnetic particles coated with a second mouse monoclonal anti-CA 125 antigen antibody. The CA 125 antigen in the sample binds to the immobilized monoclonal anti-CA 125 antigen on the solid phase, while the conjugate antibody reacts with a different antigenic site on the CA 125 antigen molecule. After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is directly proportional to the concentration of CA 125 antigen in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: 95%

Limit of Quantitation:

      - LLOQ: 0.5 U/ml

      - ULOQ: 5000 U/ml

Analytical measurement range (AMR): 0.5–5000 U/ml

Clinical Reportable Range (CRR): 0.5–5000 U/ml

- Type of Sample

A minimum of 2 ml of Serum (preferred).

Collect blood sample by venipuncture

For serum, allow samples to clot adequately before centrifugation.

Keep tubes stoppered at all times.

Physically separate serum or plasma from contact with cells as soon as possible.

Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - Follow blood collection tube manufacturer’s recommendations for centrifugation

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800

Reagents

    - Access OV Monitor Reagent Pack Cat. No. 386357: 100 determinations, 2 packs, 50 tests/pack.

    - Provided ready to use.

    - Store upright and refrigerate at 2 to 10°C.

    - Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

    - Stable until the expiration date stated on the label when stored at 2 to 10°C.

    - Stable at 2 to 10°C for 28 days after initial use.

    - Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

    - If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

    - All antisera are polyclonal unless otherwise indicated.

| R1a: | Paramagnetic particles, coated with goat anti-biotin antibodies, biotinylated anti CA 125 antigen mouse monoclonal antibodies, bovine serum albumin, < 0.1% sodium azide, and 0.1% ProClin300. |
| --- | --- |
| R1b: | Mouse monoclonal anti-CA 125 antigen-alkaline phosphatase (bovine) conjugate, bovine serum albumin, < 0.1% sodium azide, and 0.1% ProClin 300. |
| R1c: | Buffered protein solution (bovine, goat, mouse), < 0.1% sodium azide and 0.1% ProClin 300. |

    - CAUTION:

- Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

    - Do not use reagents beyond the expiration date

- Safety

For in vitro diagnostic use only.

Use this reagent only for the DXI 800.

Do not pipette by mouth.

Do not ingest. Harmful if swallowed

Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

Used cuvettes contain human body fluids, it should be handled with appropriate care to avoid skin contact.

Reagent packs: Sealed cartridges are stable until the Expiry date.

- Calibration Procedure

Calibrator Required:

- 6 Point Access OV Monitor Calibrators.

Calibration Schedule:

- An active calibration curve is required for all tests. Calibration is required every 28 days. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When quality control results are found to be consistently outside acceptable range.

- Test Procedure

Refer to the appropriate Beckman Coulter DxI 800 User Guide.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use ten (10) μL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) μL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

The system default unit of measure for sample results is ng/ml. To change sample reporting units to the International System of Units (SI units), µg/l, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply ng/ml by multiplication factor

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Analyzer generated results are reviewed by the operator & only clinically valid results are saved for approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1 and Level 2, Two levels of Quality Control sera i.e. low, normal and high as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- ILC –- Unipath

- Interferences & Cross Reactions

Some individuals have antibodies to mouse protein (HAMA) which can cause interference in immunoassays that employ antibodies derived from mice. In particular, it has been reported that serum samples from patients who have undergone therapeutic or diagnostic procedures that include infusion of mouse monoclonal antibodies may produce erroneous results in such assays. Additionally, other heterophile antibodies may be present in patient samples. Therefore, results for such patients should be used only in conjunction with results from some other diagnostic procedures and with information available from the clinical evaluation of the patient.

The following table describes the cross-reactivity of the assay with common chemotherapeutic agents and other potential interferents

| Substance | Concentration Added | Expected (U/ml) | Observed (U/ml) | Mean % Recovery |
| --- | --- | --- | --- | --- |
| Doxorubicin | 100 μg/ml | 15.0 | 15.4 | 102.7 |
| Amethopterin | 500 μg/ml | 15.0 | 14.9 | 99.3 |
| Carboplatin | 1,000 μg/ml | 15.0 | 15.2 | 101.3 |
| Cyclophosphamide | 1,000 μg/ml | 15.0 | 14.6 | 97.3 |
| 5-uorouracil | 1,000 μg/ml | 15.0 | 14.8 | 98.7 |
| Cisplatin | 2,000 μg/ml | 15.0 | 14.7 | 98.0 |
| Melphalan | 100 μg/ml | 15.0 | 14.6 | 97.3 |
| Acetominophen | 200 μg/ml | 15.0 | 15.2 | 101.3 |
| Aspirin | 500 μg/ml | 15.0 | 15.0 | 100.0 |
| Paclitaxel | 10 ng/ml | 15.0 | 15.1 | 100.7 |
| Biotin | 50 ng/ml | 15.0 | 14.4 | 96.0 |
| Vitamin D2 | 1 U/ml | 15.0 | 14.9 | 99.3 |

- Result Calculation

The distribution of Access OV Monitor results, presented below, were determined from a total of 889 serum samples from apparently healthy females and males and females with non-malignant and malignant conditions.

| Subject Category | Number of Subjects | 0-35 U/ml | 35.1-65 U/ml | 65.1-100 U/ml | > 100 U/ml |
| --- | --- | --- | --- | --- | --- |
| Apparently Healthy Females Males | 170 51 | 166 51 | 3 0 | 1 0 | 0 0 |
| Malignant Conditions (including previously treated patients) |  |  |  |  |  |
| Ovarian | 141 | 94 | 10 | 6 | 31 |
| Breast | 48 | 43 | 3 | 1 | 1 |
| Cervical | 27 | 22 | 1 | 1 | 3 |
| Colon | 50 | 47 | 2 | 1 | 0 |
| Endometrial | 8 | 5 | 2 | 1 | 0 |
| Esophageal | 9 | 7 | 1 | 0 | 1 |
| Fallopian Tube | 5 | 3 | 0 | 1 | 1 |
| Gastric/Stomach | 13 | 7 | 2 | 2 | 2 |
| Lung | 50 | 28 | 9 | 3 | 10 |
| Pancreatic | 25 | 14 | 0 | 2 | 9 |
| Uterine | 26 | 19 | 1 | 1 | 5 |
| Vaginal/Vulva | 12 | 11 | 1 | 0 | 0 |
| Non-Malignant Conditions (including previously treated patients) |  |  |  |  |  |
| Ovarian | 28 | 28 | 0 | 0 | 0 |
| Colon | 43 | 43 | 0 | 0 | 0 |
| Cystitis | 9 | 8 | 1 | 0 | 0 |
| Endometriosis | 22 | 21 | 0 | 1 | 0 |
| Gastric/Stomach | 39 | 37 | 1 | 0 | 1 |
| Pelvic Inflammatory Disease | 8 | 7 | 1 | 0 | 0 |
| Uterine Fibroids | 30 | 25 | 1 | 1 | 3 |
| Pregnancy | 75 | 42 | 22 | 4 | 7 |

- Biological Reference Interval

| Age | Range |
| --- | --- |
| - | 5.5–1000 U/ml |

- Reportable Interval

Analytical Measurement Range (AMR): 0.5–5000 U/ml

Clinical Reportable Range (CRR): 0.5–5000 U/ml

- Procedure when Results are not in Range

Samples can be accurately measured within the analytic range of the lower limit of detection and the highest calibrator value (approximately 0.5 U/ml and 5,000 U/ml).

If a sample contains less than the lower limit of detection for the assay, report the results as less than that value (i.e., < 0.5 U/ml). When the DxI system onboard dilution feature is used, the system will report results as less than 4,250 U/ml.

If a sample contains more than the stated value of the highest Access OV Monitor Calibrator (S5), report the result as greater than that value (i.e., > 5,000 U/ml). Alternatively, dilute one volume of sample with 9 or 19 volumes of Access OV Monitor Calibrator S0 (zero) or Access Sample Diluent A. Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution. The DxI system onboard dilution feature automates the dilution process, using one volume of sample with nineteen volumes of Access Sample Diluent A, allowing samples to be quantitated up to approximately 100,000 U/ml. The system reports the results adjusted for the dilution.

For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce antibodies, e.g. HAMA, that interfere with immunoassays. Additionally, other heterophile antibodies such as human anti-goat antibodies may be present in patient samples. Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

The Access OV Monitor results should be interpreted in light of the total clinical presentation of the patient, including symptoms, clinical history, data from additional tests and other appropriate information. Serum or plasma OV Monitor concentrations should not be interpreted as absolute evidence for the presence or absence of cancer. Elevated concentrations may be observed in the serum or plasma of patients with benign conditions or other non-cancer disorders, as well as in ovarian cancer and other metastatic diseases. The Access OV Monitor assay should not be used as a cancer screening test.

The Access OV Monitor assay does not demonstrate any “hook” effect up to 120,000 U/ml.

- Critical Alert Values

> 200 U/ml

- Laboratory Clinical Interpretation

CA-125 is a large glycoprotein found on the surface of many ovarian cancer cells and in some normal tissues. It is a product of the MUC16 gene.

CA-125 is recommended, together with transvaginal ultrasound, for early detection of ovarian cancer in women with hereditary syndromes and also recommended as an adjunct in distinguishing benign from malignant suspicious pelvic masses, particularly in postmenopausal women.

CA-125 values increased in malignant disease, Fallopian tube tumors, non-mucinous epithelial ovarian carcinoma, cervical adenocarcinoma, endometrial adenocarcinoma and squamous cell carcinoma of vulva or cervix.

CA-125 is used for evaluating patient’s response to ovarian cancer therapy and predicting the recurrent ovarian cancer. This test is not useful for cancer detection screening in the normal population.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of CA-125, Beckman Coulter DxI 800

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)