SOP – 016

PROCEDURE FOR CEA TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	11

13.	Test Procedure	11

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 |  |  |
| 2 | AMR | Analytical Measurement Range |
| 3 | BC | Biochemistry |
| 4 | BSA | Bovine Serum Albumin |
| 5 | °C | Degree Centigrade |
| 6 | CA | Carbohydrate Antigen |
| 7 | CEA | Carcinoembryonic Antigen |
| 8 | CLIA | Chemiluminescence Immunoassay |
| 9 | CV | Coefficient of Variation |
| 10 | CRR | Clinical Reportable Range |
| 11 | EDTA | Ethylenediaminetetraacetic Acid |
| 12 | EQAS | External Quality Assurance Scheme |
| 13 | HOD | Head of Department |
| 14 | IQC | Internal Quality Control |
| 15 | LLOQ | Lower Limit of Quantitation |
| 16 | MAb | Monoclonal Antibodies |
| 17 | QC | Quality Control |
| 18 | QM | Quality Manager |
| 19 | SD | Standard Deviation |
| 20 | SST | Serum Separator Tube |
| 21 | TDPL | TRUSTlab Diagnostics Private Limited |
| 22 | ULOQ | Upper Limit of Quantitation |

- Objective

  - The purpose of this SOP is to describe the procedure for quantitative estimation of Carcinoembryonic Antigen (CEA) in serum samples using Beckman Coulter DxI analyzer.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- The CEA test, or Carcinoembryonic Antigen test, is a blood test that measures the level of a protein called CEA in the blood. It's used primarily to monitor the progress of certain cancers, assess the effectiveness of treatment, and detect potential cancer recurrence. While elevated CEA levels can sometimes indicate cancer, they can also be elevated in non-cancerous conditions.

- Test Principle

- The Access CEA assay is a two-site immunoenzymatic “sandwich” assay using two mouse monoclonal anti-CEA antibodies (MAb) which react with different epitopes of CEA. A sample is added to a reaction vessel, along with the first anti-CEA MAb-alkaline phosphatase conjugate and the second anti-CEA MAb bound to paramagnetic particles. The incubation is followed by magnetic separation and washing. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is proportional to the concentration of CEA in the sample. The amount of analyte in the sample is determined by means of a stored, multi-point calibrator curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.1 ng/ml

      - ULOQ: 1000 ng/ml

Analytical measurement range (AMR): 0.1–1000 ng/ml

Clinical Reportable Range (CRR): > 10000 ng/ml (1:10 dilution)

Sensitivity: 95%

Specificity: N/A

- Note: If requested by treating Clinician/Doctor, values greater than CRR may be reported with further dilution, with approval of HOD.

- Type of Sample

Only non hemolysed serum/plasma

      - Ambient (+15 to 25°C): 24 hours

      - Refrigerated (+2 to +8°C): 4 days

      - Frozen (-20°C): long term storage.

- Type of Container & Additives

- Red/SST yellow topped vacutainer tubes. Do not add any additives /preservatives.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No other specific preparation required.

- Equipment & Reagents

Equipment

- Beckman DxI 800 CLIA analyzer

Reagents

- Access CEA Reagent Pack

- Cat. No. 33200: 100 determinations, 2 packs, 50 tests/pack

- Provided ready to use.

- Store upright and refrigerate packs at 2 to 10°C.

- Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

- Stable until the expiration date stated on the label when stored at 2 to 10°C.

- After initial use, the pack is stable at 2 to 10°C up to 28 days.

- Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

- If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

| R1a | Solid phase: Paramagnetic particles coated with mouse anti-CEA MAb, suspended in TRIS buffered bovine serum albumin (BSA), with < 0.1% sodium azide and 0.1% ProClin* 300. |
| --- | --- |
| R1b | Diluent: Phosphate buffer, protein (bovine, murine) with < 0.1% sodium azide and 0.1% ProClin 300. |
| R1c | Conjugate: Mouse anti-CEA MAb bound to alkaline phosphatase (bovine), diluted in phosphate buffer, protein (bovine), < 0.1% sodium azide and 0.1% ProClin 300. |

- Safety

Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the CEA Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

Handling precautions are described in detail in the UniCel DxI System Operations Manual.

- Calibration Procedure

- An active calibration curve is required for all tests. For the Access CEA assay, calibration is required every 28 days. Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

- Test Procedure

- Refer to the appropriate system manuals and/or Help system for a specific description of installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

Procedure

      - Mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument. Do not invert open (punctured) packs, instead, mix reagents by swirling gently

      - Use ten (10) µL of sample for each determination in addition to the sample container and system dead volume. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

      - The system default unit of measure for sample results is ng/ml.

Handling of Repeat Test Results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever higher).

      - If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

    - Biorad Two levels of Quality Control sera i.e. normal and pathological range are run once daily

    - Preparation of Controls:

      - Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

      - Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

      - Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

      - Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

      - Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided).

      - Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”. Store the aliquots at -20°C.

    - Storage and Stability of Controls:

      - Unopened: Stable until the expiry date when stored at 2 to 8°C.

      - Reconstituted: Stable for 7 days when stored at 2 to 8°C; stable up to 30 days when stored at -20°C. Note: Aliquots are stored at the above temperatures as required.

    - Procedure

      - Run Quality Control according to TRUS lab-defined Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)

      - Remove the QC material stored at -20°C, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8°C, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion

      - Place a portion of the material in a sample cup (approximately 350 µL).

      - Place the cup in the tray positions indicated by the screen Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levy Jennings chart and calculate the cumulative precision i.e. %CV.

    - Handling of QC failure:

      - Verify the following points and if required, consult with HOD/QM:

          - Review the environmental conditions and recent events and look for any deviation.

          - Check on-board reagents (for stability, deposits, fungal growth) reference fluid and controls.

          - Check the quality of the sterile water for injections used for the preparation of controls and calibrators, contaminants, i.e. bacterial growth may contribute to inaccurate results.

          - Check reaction temperature.

          - Check the expiry date of kit and contents.

          - Check instrument settings and light source.

          - If there are any control failures, take appropriate corrective and preventive actions as per TRUSTlab’s Quality Process MSP.

          - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

          - Take appropriate corrective actions and re-run the QC.

          - Contact application support/service support from vendor if required corrective action fails to yield acceptable results.

    - Repeat and Recheck criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation – ILC Unipath

- Interferences & Cross Reactions

Hemoglobin, triglycerides, bilirubin, human serum albumin and rheumatoid factor, tested up to the following concentrations, respectively, 500 mg/dl, 1,800 mg/dl, 30 mg/dl, 5 g/dl and 500 IU/ml, do not interfere with the Access CEA assay.

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

- Non-smokers: 0.0 - 3.0 ng/ml

- Smokers: 0.0 - 5.0 ng/ml

- Reportable Interval

- 0.1 to 1000 ng/ml

- CRR > 10000 ng/ml

- Dilution Procedure:

- If a sample contains more than the stated value of the highest Access CEA Calibrator (S5), report the result as greater than that value (i.e. > 1000 ng/ml). Alternatively, dilute one volume of the sample with nine volumes of Access CEA Calibrator S0 (zero) or Access CEA Diluent. After assaying the diluted sample, multiply the obtained value by the dilution factor, ten (10). If a sample contains less than the lower limit of detection for the assay, report the result as less than the lower limit of detection (< 0.1 ng/ml). Refer to the appropriate system manuals and/or Help system for instructions on how to enter a sample dilution in a test request.

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

      - N/A.

- Laboratory Clinical Interpretation

The measurement of serum CEA has shown substantial benefit in the prognosis and management of patients with malignant diseases, especially colorectal cancer. Serial measurements can be used to monitor patients for progression, regression or recurrence of cancer following treatment. A persistent elevation of CEA following therapeutic or surgical intervention signals residual disease or recurrence, whereas decreasing levels to within the normal range is indicative of successful intervention.

CEA is also elevated in the serum of patients with non-malignant diseases and in heavy smokers, therefore CEA should not be used in the diagnosis of cancer or for screening asymptomatic patients.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - Do not use if the reagents are turbid as it indicates contamination of the reagent.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Access DXI CEA Kit insert

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)