SOP – 026

PROCEDURE FOR

C-REACTIVE PROTEIN (CRP) TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	11

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	13

16.	Result Calculation	13

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	14

22.	Potential Source of Variability	14

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CM | Chylomicrons |
| 5 | CV | Coefficient of Variation |
| 6 | CRP | C-Reactive Protein |
| 7 | CRR | Clinical Reportable Range |
| 8 | EDTA | Ethylenediaminetetraacetic Acid |
| 9 | EQAS | External Quality Assurance Scheme |
| 10 | HOD | Head of Department |
| 11 | IFCC | International Federation for Clinical Chemistry |
| 12 | IQC | Internal Quality Control |
| 13 | LLOQ | Lower Limit of Quantitation |
| 14 | MSP | Management System Procedure |
| 15 | N/A | Not Applicable/Available |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | RPPHS | Reference Preparation for Proteins in Human Serum |
| 20 | SD | Standard Deviation |
| 21 | SOP | Standard Operating Procedure |
| 22 | SST | Serum Separator Tube |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |
| 26 |  |  |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of C-Reactive Protein (CRP) levels in serum samples using Beckman Coulter automated biochemistry analyzer. This SOP also provides information about the test principle, clinical significance etc.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- C-reactive protein (CRP) is a protein produced by the liver in response to inflammation. It's a highly sensitive acute-phase reactant, meaning its levels rise quickly in the blood following injury or inflammation.

- Test Principle

Immune complexes formed in solution scatter light in proportion to their size, shape, and concentration. Turbidimeters measure the reduction of incidence light due to reflection, absorption, or scatter. In this procedure, the measurement of the rate of decrease in light intensity transmitted (increase in absorbance) through particles suspended in solution is the result of complexes formed during the antigen-antibody reaction.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 5 mg/l

      - ULOQ: 300 mg/l

Analytical measurement range (AMR): 5 - 300 mg/l

Clinical Reportable Range (CRR): 5 - 300 mg/l

- N/A.

- Type of Sample

Whole Blood (A minimum of 2 ml)

Collect blood sample by venipuncture and separate serum by centrifugation.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxC 700 AU Chemistry Analyzer

Reagents

      - Concentrations

| - Reactive Ingredient | - Final Concentration |
| --- | --- |
| Tris buffer (pH 7.5) | - 80 mmol/l |
| Sodium Chloride | - 125 mmol/l |
| Polyethylene glycol 6000 | - 1.5% |
| Goat anti-CRP Antibodies | - ~ 0.6 g/l |
| Preservative | - N/A |

      - Reagent Preparation

          - Reagents are ready for use.

          - Can be directly placed on board the instrument.

      - Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-8°C.

          - Once open, reagents stored on board the instrument are stable for 90 days.

- Safety

For in vitro diagnostic use only.

Do not pipette by mouth.

Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

- Calibration Procedure

Calibrator Required:

      - Use System Calibrator: Cat. No.Cat # ODR3021

      - Calibrator traceable to IFCC (International Federation of Clinical Chemistry) standard CRM 470 (RPPHS)

Calibration Schedule:

- Calibration verification is done at least once every 6 months. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When required as per government regulations.

      - If quality control results are consistently outside acceptable range.

Calibrator Preparation:

      - Calibrator material should be at room temperature before reconstitution. Keep the vials at room temperature for approximately 2 Hours, after taking out from the Deep freezer.

      - Slowly invert the calibrator or diluent bottles several times to mix the contents thoroughly. DO NOT SHAKE.

      - Gently tap the lyophilized calibrator vial on the counter several times to dislodge any material adhering to the stopper.

      - Remove the seal and stopper from each bottle just prior to the addition of the diluent. Do not leave vials unstopped.

      - Reconstitute by pipetting exactly 3ml of distilled water at +20 to +25°C, into the vial.

      - Replace the stopper and hold it firmly in place. Invert the vial gently. DO NOT SHAKE.   Visually verify that all freeze-dried material is dissolved prior to use.

Storage and Stability of Calibrators:

      - Un-reconstituted Calibrator is stable up to the expiry date shown on the side of each individual bottle.

      - Once reconstituted, the components of the Calibration Sera are stable for 8 hours at 25°C, 1 day at 2-8°C.

      - Note: Do not use the reconstitute calibrator beyond stability period.

- Test Procedure

Refer to the appropriate Beckman Coulter AU Analyzer User Guide.

Prepare a work list of registered samples with required investigations and collect listed samples

Sample Processing Using Racks:

      - Place each sample in a rack, as per position number in the list (if volume is not sufficient transfer sample into a validated low volume container).

      - Open the top and/or the front cover of the Rack Input area.

      - Place the rack in the input area with the Rack ID barcode facing the back of the analyzer.

      - Do not leave any empty space between the racks.

      - Close the Rack Input area cover.

      - Select Start button.

To monitor sample status on the Analyzer, select - Result.

Analyzer generated results are reviewed by the technician & only clinically valid results are saved as a precursor to approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1 and Level 2, Two levels of Quality Control sera i.e. normal and pathological range as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity  (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check expiry date of kit and contents.

      - Check instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

- Results of studies conducted to evaluate the susceptibility of the method to interference were as follows:

Lysed serum specimens may give falsely elevated values.

If the reagents are turbid, DO NOT USE as it indicates contamination of the reagent

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

- Adults: < 5 mg/l

- Reportable Interval

Analytical Measurement Range (AMR): 5 - 300 mg/l

Clinical Reportable Range (CRR): 5 - 300 mg/l

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

C-Reactive Protein (CRP) is one of the most sensitive acute-phase reactants for inflammation. Measuring changes in the concentration of CRP provides useful diagnostic information about the level of acuity and severity of a disease. Unlike ESR, CRP levels are not influenced by hematologic conditions such as anemia, polycythemia etc. Increased levels are consistent with an acute inflammatory process. After onset of an acute phase response, the serum CRP concentration rises rapidly (within 6-12 hours and peaks at 24-48 hours) and extensively. Concentrations above 100 mg/l are associated with severe stimuli such as major trauma and severe infection (sepsis).

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of CRP, Beckman Coulter AU680 / AU 480 / DXC700

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)