SOP – 031

PROCEDURE FOR

FSH TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	13

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	16

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CV | Coefficient of Variation |
| 5 | CRR | Clinical Reportable Range |
| 6 | EDTA | Ethylenediaminetetraacetic Acid |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | FSH | Follicle Stimulating Hormone |
| 9 | HOD | Head of Department |
| 10 | IFCC | International Federation for Clinical Chemistry |
| 11 | IQC | Internal Quality Control |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MSP | Management System Procedure |
| 14 | N/A | Not Applicable/Available |
| 15 | pH | Potential of Hydrogen |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | SD | Standard Deviation |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TRF | Test Requisition Form |
| 23 | TDPL | TRUSTlab Diagnostics Private Limited |
| 24 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Follicle stimulating Hormone (FSH) in serum samples using Beckman Coulter DxI 800. This SOP also provides information about the test principle, clinical significance and other test specific details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- FSH, or Follicle-Stimulating Hormone, is a hormone produced by the pituitary gland that plays a crucial role in reproductive health, particularly in women and men. In women, FSH stimulates the growth of ovarian follicles and the production of eggs, while in men, it stimulates sperm production.

- Test Principle

- The DxI hFSH assay is a sequential two-step immunoenzymatic (“sandwich”) assay. A sample is added to a reaction vessel with paramagnetic particles coated with goat anti-mouse: mouse anti-hFSH complexes and TRIS buffered saline with protein. The hFSH binds to the immobilized mouse anti-hFSH on the solid phase. Materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Alkaline phosphatase conjugated goat anti-hFSH is then added and binds to the previously bound hFSH on the particles. A second separation and wash step removes unbound conjugate. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is directly proportional to the concentration of hFSH in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Uncertainty of measurement: CV (%): 11.43

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.2 mIU/ml

      - ULOQ: 200 mIU/ml

Analytical measurement range (AMR): 0.2 – 200 mIU/ml

Sensitivity: 95%

Specificity: N/A.

- Type of Sample

Only non-hemolysed serum/plasma.

Collect blood sample by venipuncture.

For serum, allow samples to clot adequately before centrifugation.

Keep tubes stoppered at all times.

Physically separate serum or plasma from contact with cells as soon as possible.

Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - Follow blood collection tube manufacturer’s recommendations for centrifugation

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

Do not add any preservatives or additives.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800

Reagents

    - **Not provided**

Reagent Storage & Stability

    - .

- Safety

Do not use reagent kits beyond the expiration date. Do not mix reagents from different reagent kits.

Prior to loading the FSH Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

- Calibration Procedure

Calibrator Required:

- N/A.

Calibration Schedule:

    - An active calibration curve is required for all tests. Calibration is required every 28 days. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When quality control results are found to be consistently outside acceptable range.

    - Calibrators run in duplicate.

- Test Procedure

Refer to the appropriate Beckman Coulter DxI 800 User Guide.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use forty-five (45) µL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) µL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

For detailed information on ordering dilutions, refer to the DXI System Operations Manual, Section 5.

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Analyzer generated results are reviewed by the operator & only clinically valid results are saved for approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1 and Level 2, Two levels of Quality Control sera i.e. normal and high are run once daily as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity  (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Storage and Stability of Controls:

      - Unopened: Stable until the expiry date when stored at 2 to 8°C.

      - Reconstituted: Stable for 5 days when stored at 2 to 8°C.

      - Longer storage: -20°C

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

    - Storage and Stability of Controls:

      - Unopened: Stable until the expiry date when stored at 2 to 8 °C.

      - Reconstituted: Stable for 7 days when stored at 2 to 8°C; stable up to 30 days when stored at -20°C.

      - Note: Aliquots are stored at the above temperatures as required.

External Quality Control

- EQAS Participation – Metropolis EQAS

- Interferences & Cross Reactions

Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays. Patients routinely exposed to animals or to animal serum  products can be prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

| - Sample Status | - Insignificant effect up to concentration of: |
| --- | --- |
| - Hemolyzed | - 150 mg/dl of haemoglobin |
| - Lipemic | - 1000 mg/dl of triglycerides |
| - Icteric | - 20 mg/dl of bilirubin |

- Result Calculation

TRUSTlab NRL follows estimation of CV%.

Interpret the results of this test in conjunction with the patient’s clinical presentation.

Handling of repeat test results:

      - In case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever is higher).

      - If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented.

      - If required a fresh sample may be requested (subject to feasibility).

- Biological Reference Interval

| Age | Range (mIU/ml) |
| --- | --- |
| Follicular Phase | 2.5 - 10.2 |
| Ovulation Phase | 3.4 - 33.4 |
| Luteal Phase | 1.5 - 9.1 |
| Postmenopause | 16.0 - 120.0 |

- Reportable Interval

Clinical Reportable Range (CRR): 0.2 – 200 mIU/ml

Values greater than CRR may be reported using further dilution (1:1 dilution) if so requested by the Clinician/doctor, and with the approval of the HOD.

Dilution Procedure:

      - If a sample contains more than the stated value of the highest DXI hFSH Calibrator (S5), report the result as greater than that value [i.e., >200 mIU/ml (IU/l)]. Alternatively, dilute one volume of sample with equal volumes of Access hFSH Calibrator S0 (zero) or Access Sample Diluent A. Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution.

      - The DxI system onboard dilution feature automates the dilution process, using one volume of sample with one volume of Access Sample Diluent A, allowing samples to be quantitated up to approximately 400 mIU/ml (IU/l). The system reports the results adjusted for the dilution.

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

N/A.

- Laboratory Clinical Interpretation

FSH is a glycoprotein produced by the anterior pituitary gland, regulated by hypothalamic gonadotropin-releasing hormone (GnRH).

FSH stimulates follicular growth and stimulates seminiferous tubules and testicular growth.

The FSH test is used in diagnosis of gonadal, pituitary, hypothalamic disorders, and management of infertility.

FSH levels increased in primary hypogonadism, Gonadotropin-secreting pituitary tumors, precocious puberty, complete testicular feminization syndrome.

FSH levels are decreased in secondary hypogonadism, Kallmann syndrome, Pituitary LH or FSH deficiency.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of FSH, Beckman Coulter DxI 800

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)