SOP – 050

PROCEDURE FOR

TSH TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	11

13.	Test Procedure	12

14.	Quality Control Procedure	13

15.	Interferences & Cross Reactions	15

16.	Result Calculation	17

17.	Biological Reference Interval	18

18.	Reportable Interval	18

19.	Procedure when Results are not in Range	18

20.	Critical Alert Values	18

21.	Laboratory Clinical Interpretation	18

22.	Potential Source of Variability	19

23.	References	19

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HAMA | Human Anti-Mouse Antibody |
| 9 | HOD | Head of Department |
| 10 | IFCC | International Federation for Clinical Chemistry |
| 11 | IQC | Internal Quality Control |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MSP | Management System Procedure |
| 14 | N/A | Not Applicable/Available |
| 15 | pH | Potential of Hydrogen |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | SD | Standard Deviation |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TSH | Thyroid Stimulating Hormone |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of TSH in serum and plasma samples using Beckman Coulter CLIA analyzer. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- TSH stands for thyroid-stimulating hormone. It's a hormone produced by the pituitary gland that stimulates the thyroid gland to produce thyroxine (T4) and triiodothyronine (T3), which are essential thyroid hormones. A TSH test measures the amount of this hormone in the blood and is used to assess thyroid function.

- Test Principle

The Access TSH (3rd IS) assay is a two-site immunoenzymatic (“sandwich”) assay.

A sample is added to a reaction vessel with mouse anti-hTSH-alkaline phosphatase conjugate, buffered protein solution and paramagnetic particles coated with immobilized mouse monoclonal anti-hTSH antibody.

The hTSH binds to the immobilized monoclonal anti-hTSH antibody on the solid phase while the mouse anti-hTSH-alkaline phosphatase conjugate reacts with a different antigenic site on the hTSH.

After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away.

Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer.

The light production is directly proportional to the concentration of TSH in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.005 μIU/ml

      - ULOQ: 50.0 μIU/ml

Analytical measurement range (AMR): 0.005 - 50.0 μIU/ml

Clinical Reportable Range (CRR): > 500 μIU/ml (1:10 dilution)

Sensitivity: 95%

Specificity: N/A

- Type of Sample

A minimum of 2 ml of non-hemolysed Serum/Plasma

      - Collect the blood sample by venipuncture observing universal precautions.

      - Allow samples to clot adequately before centrifugation.

      - Keep tubes stoppered and upright at all times.

      - Do not use samples that have been stored at room temperature for longer than 8 hours.

      - Tightly cap and refrigerate specimens at 2 to 8°C if the assay is not completed within 8 hours.

      - Freeze samples at or below -20°C if the sample is not assayed within 48 hours.

      - Freeze samples only once and mix thoroughly after thawing.

      - Before placing samples on the system ensure that:

      - Samples are free of fibrin or other particulate matter.

      - Samples are free of bubbles.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

Heparin – Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

- Beckman Coulter Access TSH (3rd IS) Reagent Pack; Cat. No. B63284: 200 determinations, 2 packs, 100 tests/pack

      - Provided ready to use.

      - Store upright and refrigerate at 2-10°C.

      - Refrigerate at 2-10°C for a minimum of two hours before use on the instrument.

      - Signs of possible deterioration are a broken elastomeric layer on the pack or quality control values out of range.

- If the reagent pack is damaged (e.g., broken elastomer), discard the pack

| Reagent | Ingredients |
| --- | --- |
| R1a | Paramagnetic particles coated with goat anti-mouse IgG: mouse monoclonal anti-hTSH complexes suspended in TRIS buffered saline, with surfactant, bovine serum albumin (BSA), < 0.1% sodium azide, and 0.1% ProClin 300. |
| R1b | TRIS buffered saline with surfactant, BSA, protein (murine, goat), < 0.1% sodium azide, and 0.1% ProClin 300. |
| R1c | Goat anti-hTSH-alkaline phosphatase (bovine) conjugate in TRIS buffered saline, with surfactant, BSA, protein (goat), < 0.1% sodium azide, and 0.1% ProClin 300 |

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-10°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the ACCESS 2 TSH Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

For a detailed discussion of handling precautions during system operation, refer to the ACCESS 2 System Operations Manual.

Storage Instructions

      - The ACCESS 2 TSH Reagent Kit must be stored at 2-8°C and may be used immediately after removal from 2-8°C storage.

      - When stored and handled as directed, reagents are stable until the expiration date.

      - The ACCESS 2 TSH Reagent Kit may be stored on board the ACCESS 2 i System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking onboard time, refer to the ACCESS 2 System Operations Manual, Section 5.

      - Reagents may be stored on or off the ACCESS 2 i System. If reagents are removed from the system, store them at 2-8°C (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright.

      - If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded.

      - After reagents are removed from the system, you must initiate a scan to update the onboard stability timer.

- Calibration Procedure

Calibrator Required

      - The Access TSH (3rd IS) Calibrators are provided at six levels – zero and approximately 0.050, 0.30, 3.0, 15.0, and 50.0 µIU/ml (mIU/l). Assay calibration data are valid up to 28 days.

Calibration Schedule

      - An active calibration curve is required for all tests

      - Recalibrate the assay every 28 days.

      - Additionally, the Beckman Coulter TSH assay requires a two-point calibration when:

- changing lot numbers of primary reagent packs

- replacing critical system components due to service or maintenance issues

- preventive maintenance has been performed

- quality control results are repeatedly out of acceptable range

- LJ plots show positive or negative bias for 10 data points

- Calibration or calibration-verification has failed.

Preparation of Calibrators

      - Calibrator material should be at room temperature before reconstitution. Keep the vials at room temperature for approximately 2 Hours, after taking out from the Deep freezer.

      - Slowly invert the calibrator or diluent bottles several times to mix the contents thoroughly. DO NOT SHAKE.

      - Gently tap the lyophilized calibrator vial on the counter several times to dislodge any material adhering to the stopper.

      - Remove the seal and stopper from each bottle just prior to the addition of the diluent. Do not leave vials unstopped.

      - Reconstitute by pipetting exactly 3ml of distilled water at +20 to +25°C, into the vial.

      - Replace the stopper and hold it firmly in place. Invert the vial gently. DO NOT SHAKE.   Visually verify that all freeze-dried material is dissolved prior to use

Storage and Stability of Calibrators

      - Un-reconstituted Calibrator is stable up to the expiry date shown on the side of each individual bottle.

      - Once reconstituted, the components of the Calibration Sera are stable for 8 hours at +25°C, for 7 days at +4°C and for 1 month at -20°C (frozen).

      - NOTE: Do not use the reconstitute calibrator beyond stability period.

- Test Procedure

- CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Start the entry queue, if required.

CAUTION: Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

Do not use reagents beyond the expiration date.

Handling of repeat test results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10% or 1.5 times the parameter CV% (whichever is higher).

      - If the percentage difference between the values is greater than 10% or 1.5 times the parameter CV% (whichever is higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

- Biorad - Level 1, Level 2, Level 3 any two levels of Quality Control sera i.e. normal, low or high, are run once daily as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity (frothing to be avoided).

- Visually verify that all freeze-dried material is dissolved prior to use.

- Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

Storage & Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2-8°C.

      - Reconstituted: Stable for 5 days when stored, tightly capped, at 2-8°C.

      - Longer storage: -20°C

      - Aliquots are stored at the above temperatures as required.

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce antibodies, e.g. HAMA, that interfere with immunoassays. Additionally, other heterophile antibodies such as human anti-goat antibodies may be present in patient samples.

Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

The Access TSH results should be interpreted in light of the total clinical presentation of the patient, including symptoms, clinical history, data from additional tests and other appropriate information.

Samples containing up to 10 mg/dl (171 mol/l) bilirubin, lipemic samples containing the equivalent of 1800 mg/dl (20.32 mmol/l) triolein, and hemolyzed samples containing up to 1 g/dl (10 g/l) hemoglobin do not affect the concentration of TSH assayed.

The following table describes the cross-reactivity of the antibody used in the TSH assay with substances which are similar in structure to T4. The antibody specificity is determined in a total thyroxine assay to avoid binding displacement of T4 by related compounds. The substances were added to an Access TSH calibrator pool and found to give the following results expressed by weight.

| Substance | Analyte Added (g/dl) | Cross-Reactivity (%) |
| --- | --- | --- |
| L-T4 | 5 | > 100 |
| D-T4 | 10 | 71 |
| L-T3 | 500 | 1.7 |
| R-T3 | 100 | 23 |
| Tetraiodothyroacetic Acid | 25 | 4 |
| D-T3 | 500 | 0.7 |
| 3,3’ L-T2 | 5000 | < 0.1 |
| 3,5 L-T2 | 5000 | < 0.1 |
| 3’5’ L-T2 | 5000 | < 0.1 |
| L-Tyrosine | 5000 | < 0.01 |
| D-Tyrosine | 5000 | < 0.01 |
| Monoiodotyrosine | 5000 | < 0.01 |
| Diiodotyrosine | 5000 | < 0.01 |

The lowest detectable level of TSH distinguishable from zero (Access TSH Calibrator S0) with 95% confidence is 0.25 ng/dl.

Samples containing thyroxine autoantibodies can be assayed in two step procedures such as the Access TSH assay without significant interference.Non-thyrometabolic disorders may cause abnormal TSH levels. Anticonvulsant drug therapy (particularly phenytoin) may result in decreased TSH levels due to an increased hepatic metabolism, and secondarily to displacement of hormone from binding sites. Anti-inflammatory drugs such as salicylate and phenylbutazone also compete for hormone binding sites, but their effect on TSH levels has not been clearly defined. Patients on heparin therapy may have elevated TSH levels due to release of non-esterified fatty acids, which can alter the relationship between free and bound hormones. Determination of thyroid status in patients with non-thyroidal illness (NTI) should be interpreted with caution. In rare conditions, such as Familial Dysalbuminemic Hyperthyroxinemia (FDH), direct free hormone assays may yield erroneous results due to the extreme variations in the albumin-binding capacity for T4.

The following substances were added to euthyroid serum samples. When tested in the TSH assay, the observed mean percent changes in TSH values were as follows:

| Substance | Amount Added (mg/dl) | % Change |
| --- | --- | --- |
| Aspirin | 60 | +6.4 |
| Sodium Salicylate | 50 | +9.1 |
| Phenylbutazone | 7.5 | +8.48 |
| Thiouracil | 5.0 | -0.6 |
| Phenytoin | 5.0 | +8.0 |
| Methimizole | 0.4 | +0.9 |

This assay is not significantly influenced by the presence of thyroxine hormone binding proteins. In one study, the zero calibrator, which contains normal levels of protein, was spiked with the following human source thyroxine hormone binding proteins. Below is the observed change to the assay signal response:

| Substance | Amount Added | % Change |
| --- | --- | --- |
| Albumin | 10.0 g/dl | -7.2 |
| Thyroxine binding globulin (TBG) | 160 µg/ml | -5.6 |
| Prealbumin | 600 µg/ml | -1.3 |

- Result Calculation

- Patient test results are determined automatically by the system software using a weighted four parameter logistic curve (4PLC) math model. The amount of analyte in the sample is determined from the measured light production by means of the stored calibration data. Patient test results can be reviewed using the appropriate screen.

- Biological Reference Interval

- 0.38 – 5.33 μIU/ml

- Reportable Interval

Analytical Measurement Range (AMR): 0.005 - 50.0 μIU/ml

Clinical Reportable Range (CRR): Up to 1:10 dilution

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the clinician/doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Human thyroid-stimulating hormone is a glycoprotein hormone consisting of two noncovalently-bound subunits: an α subunit, which is nearly identical to the α subunits of human luteinizing hormone (hLH), human follicle-stimulating hormone (hFSH), and human chorionic gonadotropin (hCG), and a β subunit, which is responsible for immunological and biological specificity.

TSH, released from the anterior pituitary, is the principal regulator of thyroid function, stimulating the synthesis and release of thyroid hormones thyroxine (T4) and triiodothyronine (T3). T3 and T4 regulate biochemical processes that are essential for normal metabolism. The synthesis and secretion of TSH is stimulated by thyrotropin-releasing hormone (TRH), which is produced by the hypothalamus in response to low levels of circulating T3 and T4. In contrast, elevated levels of T3 and T4 suppress the production of TSH. Collectively, this negative feedback system is referred to as the hypothalamic-pituitary-thyroid axis. Any alteration in the function of this axis can influence the levels of TSH, T4, and T3 in circulation.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition.

      - Product insert of TSH, Beckman Coulter DxI 800.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).