SOP – 052

PROCEDURE FOR

TOTAL T4 TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	11

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | pH | Potential of Hydrogen |
| 15 | QC | Quality Control |
| 16 | QM | Quality Manager |
| 17 | QNS | Quantity Not Sufficient |
| 18 | SD | Standard Deviation |
| 19 | SOP | Standard Operating Procedure |
| 20 | SST | Serum Separator Tube |
| 21 | T4 | Thyroxine |
| 22 | TRF | Test Requisition Form |
| 23 | TDPL | TRUSTlab Diagnostics Private Limited |
| 24 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Total T4 in serum and plasma samples using Beckman Coulter CLIA analyzer. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Total T4, or thyroxine, is a blood test that measures the total amount of thyroxine in the blood, including both bound and free forms. It's a key test used to assess thyroid function.

- Test Principle

The Access Total T4 assay is a competitive binding immunoenzymatic assay.

A sample is added to a reaction vessel with anti-thyroxine antibody, thyroxine-alkaline phosphatase conjugate, and paramagnetic particles coated with goat anti-mouse capture antibody and a stripping agent to dissociate all T4 from binding proteins.

Thyroxine in the sample competes with the thyroxine-alkaline phosphatase conjugate for binding sites on a limited amount of specific anti-thyroxine antibody. Resulting antigen: antibody complexes bind to the capture antibody on the solid phase.

After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away.

Then, the chemiluminescent substrate Lumi-Phos 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is inversely proportional to the concentration of thyroxine in the sample.

The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.50 μg /dl

      - ULOQ: 30.0 μg /dl

Analytical measurement range (AMR): 0.50-30.0 μg/dl

Clinical Reportable Range (CRR): 0.50-30.0 μg/dl

Sensitivity: N/A

Specificity: N/A

- Type of Sample

A minimum of 2ml of non-hemolysed Serum/Plasma

      - Collect the blood sample by venipuncture observing universal precautions.

      - Allow samples to clot adequately before centrifugation.

      - Keep tubes stoppered and upright at all times.

      - Do not use samples that have been stored at room temperature for longer than 8 hours.

      - Tightly cap and refrigerate specimens at 2 to 8°C if the assay is not completed within 8 hours.

      - Freeze samples at or below -20°C if the sample is not assayed within 48 hours.

      - Freeze samples only once and mix thoroughly after thawing.

      - Before placing samples on the system ensure that:

      - Samples are free of fibrin or other particulate matter.

      - Samples are free of bubbles.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

| Reagent Pack | Reagent | Ingredients |
| --- | --- | --- |
| 50 tests/pack | R1a | Mouse monoclonal anti-T4 alkaline phosphatase (bovine) conjugate and Dynabeads® paramagnetic particles coated with streptavidin in a TRIS buffer with protein (aves and murine), surfactant, < 0.1% sodium azide, and 0.1% ProClin 300. |
| 50 tests/pack | R1b | T4 analogue coupled to biotin in a TRIS buffer with protein (aves), surfactant, < 0.1% sodium azide, and 0.1% ProClin 300. |
| 50 tests/pack | R1c | 0.4N Sodium hydroxide (NaOH) solution with 8-Anilino-1-Napthalenesulfonic Acid (ANS) |
| 50 tests/pack | R1d | 0.4N Hydrochloric acid (HCl) solution |

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-10°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

Use this reagent for in vitro diagnostics use only.

Do not pipette by mouth. Exercise the normal precautions required for handling laboratory reagents.

Sodium azide can react with copper and lead plumbing to form explosive metal azides..

- While disposing the reagent, flush with large amounts of water to prevent azide build up as azide reacts with lead and copper plumbing, to form potentially explosive azides.

Exposed metal surfaces should be cleaned with 10% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

Used cuvettes contain human body fluids, and should therefore, be handled with appropriate care to avoid skin contact.

Reagent packs: Sealed cartridges are stable until the Expiry date.

- Calibration Procedure

Calibrator Required:

      - Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

Calibration Schedule:

      - An active calibration curve is required for all tests

      - Recalibrate the assay every 14 days.

      - Additionally, the Beckman Coulter T4 assay requires a two-point calibration when:

- changing lot numbers of primary reagent packs

- replacing system components

- quality control results are repeatedly out of range

Storage and Stability of Calibrators:

      - Un-reconstituted Calibrator is stable up to the expiry date shown on the side of each individual bottle.

      - Once reconstituted, the components of the Calibration Sera are stable for 8 hours at +25°C, for 7 days at +4°C and for 1 month at -20°C (frozen).

      - NOTE: Aliquots are usually stored at -20°C

- Test Procedure

- CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Start the entry queue, if required.

CAUTION: Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

Do not use reagents beyond the expiration date.

Handling of repeat test results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10% or 1.5 times the parameter CV% (whichever is higher).

      - If the percentage difference between the values is greater than 10% or 1.5 times the parameter CV% (whichever is higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

- Biorad - Level 1, Level 2, Level 3, any two levels of Quality Control sera i.e. low, normal and high run daily as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity (frothing to be avoided).

- Visually verify that all freeze-dried material is dissolved prior to use.

- Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

Storage & Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2-8°C.

      - Reconstituted: Stable for 7 days when stored, tightly capped, at 2-8°C.

      - Longer storage: -20°C

      - Aliquots are stored at the above temperatures as required.

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

Samples containing up to 10 mg/dl (171 μmol/l) bilirubin, lipemic samples containing the equivalent of 1800 mg/dl (20.52 mmol/l) triolein or up to 400 mg/dl (10.4 mmol/l) cholesterol, and hemolyzed samples containing up to 1000 mg/dl (10 g/l) hemoglobin do not affect the concentration of total T4 assayed.

In addition, total protein concentrations ranging from 5.0 to 9.0 g/dl (50 to 90 g/l) do not affect the concentration of total T4 assayed.

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

- 4.82 - 15.65 μg/dl

- Reportable Interval

Analytical Measurement Range (AMR): 0.50 - 30.0 μg/dl

Clinical Reportable Range (CRR): 0.50-30.0 μg/dl

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the clinician/doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Triiodothyronine (T4) levels are obtained in suspected cases of hyperthyroidism, either because the patient has typical symptoms or when the thyroid-stimulation hormone (TSH) levels are lower than normal.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition.

      - Product insert of T4, Beckman Coulter DxI 800.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).