SOP – 056

PROCEDURE FOR

FT3 TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	11

12.	Calibration Procedure	11

13.	Test Procedure	12

14.	Quality Control Procedure	13

15.	Interferences & Cross Reactions	15

16.	Result Calculation	16

17.	Biological Reference Interval	16

18.	Reportable Interval	16

19.	Procedure when Results are not in Range	16

20.	Critical Alert Values	17

21.	Laboratory Clinical Interpretation	17

22.	Potential Source of Variability	17

23.	References	17

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CV | Coefficient of Variation |
| 5 | CRR | Clinical Reportable Range |
| 6 | EDTA | Ethylenediaminetetraacetic Acid |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | FT3 | Free triiodothyronine |
| 9 | HOD | Head of Department |
| 10 | IFCC | International Federation for Clinical Chemistry |
| 11 | IQC | Internal Quality Control |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MES | 2-(N-morpholino) ethane sulfonic acid |
| 14 | MSP | Management System Procedure |
| 15 | N/A | Not Applicable/Available |
| 16 | pH | Potential of Hydrogen |
| 17 | QC | Quality Control |
| 18 | QM | Quality Manager |
| 19 | QNS | Quantity Not Sufficient |
| 20 | SD | Standard Deviation |
| 21 | SOP | Standard Operating Procedure |
| 22 | SST | Serum Separator Tube |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of free T3 (FT3) in serum using Beckman Coulter DxI 800. This SOP also provides information about the test principle, clinical significance and other test specific details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Free triiodothyronine (FT3) is the unbound, active form of the thyroid hormone triiodothyronine (T3) that circulates in the bloodstream and is available to enter cells and influence metabolism. It is a key hormone produced by the thyroid gland and plays a vital role in regulating metabolism, energy levels, and other bodily functions.

- Test Principle

The Access Free T3 assay is a competitive binding immunoenzymatic assay. A sample is added to a reaction vessel with an anti-T3 monoclonal antibody conjugated to alkaline phosphatase. During the incubation, the free T3 in the sample reacts with the anti-T3 antibody. Particles coated with streptavidin and biotinylated T3 analog are then added to the mixture. Unoccupied binding sites on the anti-T3 antibody are bridged to the particle through the T3 analog. After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is inversely proportional to the concentration of free T3 in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: 90%

Limit of Quantitation:

      - LLOQ: 0.88 pg/ml

      - ULOQ: 30 pg/ml

Analytical measurement range (AMR): 0.88 pg/ml to 30 pg/ml

Sensitivity: 95%

Specificity: N/A.

- Type of Sample

Only non-hemolysed serum/plasma.

Collect blood sample by venipuncture.

For serum, allow samples to clot adequately before centrifugation.

Keep tubes stoppered at all times.

Physically separate serum or plasma from contact with cells as soon as possible.

Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - Follow blood collection tube manufacturer’s recommendations for centrifugation

      - NOTE: Samples SHOULD NOT be diluted for free T3 determination.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800

Reagents

    - R1: Access Free T3 Reagent Pack

    - Cat. No. A13422: 100 determinations, 2 packs, 50 tests/pack.

    - Provided ready to use.

    - Store upright and refrigerate at 2 to 10°C.

    - Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

    - Stable until the expiration date stated on the label when stored at 2 to 10°C.

    - Stable at 2 to 10°C for 28 days after initial use.

    - Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

    - If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

    - All antisera are polyclonal unless otherwise indicated.

| R1a: | Paramagnetic particles coated with streptavidin in a TRIS buffer with protein (aves), surfactant, < 0.1% sodium azide and 0.1% ProClin 300 |
| --- | --- |
| R1b: | MES buffer and 0.1% ProClin 300 |
| R1c: | T3 analog in a TRIS buffer with protein (aves), surfactant, < 0.1% sodium azide and 0.1% ProClin 300 |
| R1d: | TRIS buffer containing animal protein (goat, bovine, aves), surfactant, < 0.1% sodium azide, and 0.5% ProClin 300 |
| R1e: | Monoclonal antibody-alkaline phophatase conjugate in an ACES buffer with protein (aves), surfactant, < 0.1% sodium azide and 0.1% ProClin 300 |

    - CAUTION:

- Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

    - Do not use reagents beyond the expiration date

Reagent Storage & Stability

    - The COULTER Free T3 Reagent Kit must be stored at 2-8°C and may be used immediately after removal from 2-8°C storage.

    - When stored and handled as directed, reagents are stable until the expiration date.

    - The COULTER Free T3 Reagent Kit may be stored on board the COULTER  System for a maximum of 30 days. After 30 days, the reagent kit must be discarded. For information on tracking onboard time, refer to the COULTER System Operations Manual.

    - Reagents may be stored on or off the COULTER i System. If reagents are removed from the system, store them at 2-8°C (with septa and replacement caps) in an upright position. For reagents stored off the system, it is recommended that they be stored in their original trays and boxes to ensure they remain upright.

    - If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded.

    - After reagents are removed from the system, you must initiate a scan to update the onboard stability timer.

- Safety

Do not use reagent kits beyond the expiration date. Do not mix reagents from different reagent kits.

Prior to loading the COULTER Free T3 Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

- Calibration Procedure

Calibrator Required:

- Cat. No. A13430: S0-S5, 2.5 mL/vial

- Provided ready to use. Store upright and refrigerate at 2 to 10°C. Mix contents by gently inverting before use. Avoid bubble formation. Stable until the expiration date stated on the label when stored at 2 to 10°C. Vial is stable at 2 to 10°C for four months after initial use. Signs of possible deterioration are control values out of range. Refer to calibration card for exact concentrations.

Calibration Schedule:

    - An active calibration curve is required for all tests. Calibration is required every 28 days. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When quality control results are found to be consistently outside acceptable range.

    - The Access Free T3 Calibrators are provided at six levels - zero and approximately 1, 2, 5, 10, and 30 pg/ml. Assay calibration data are valid up to 28 days.

    - Calibrators run in duplicate.

- Test Procedure

Refer to the appropriate Beckman Coulter DxI 800 User Guide.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use ten (10) μL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) μL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate manual system and/or Help system for the minimum sample volume required.

The system default unit of measure for sample results is ng/ml. To change sample reporting units to the International System of Units (SI units) µg/l, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply ng/ml by multiplication factor

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Analyzer generated results are reviewed by the operator & only clinically valid results are saved for approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1, Level 2 and Level 3 any Two levels of Quality Control sera i.e. low, normal and high run daily as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS Participation – Metropolis EQAS

- Interferences & Cross Reactions

Samples containing up to 20 mg/dl (342 µmol/l) bilirubin, 10 mg/dl (119 µmol/l) conjugated bilirubin, lipemic samples containing the equivalent of 3000 mg/dl (33.87 mmol/l) triolein, or up to 500 mg/dl cholesterol and samples containing up to 500 mg/dl (5.0 g/l) of hemoglobin do not effect the concentration of free T3 assayed.

The following substances were added to serum samples. When tested in the Access Free T3 assay, the observed mean percent changes were as follows:

| Substance | Analyte Added (mg/dl) | % Change |
| --- | --- | --- |
| Aspirin | 75 | 1.3 |
| Sodium Salycilate | 75 | 13.9 |
| Ibuprophen | 750 | 17.2 |
| Acetaminophen | 200 | -5.0 |
| Phenylbutazone | 7.5 | 3.9 |
| Thiouracil | 5 | 2.1 |
| Phenytoin | 10 | -0.1 |
| Furosemide | 2 | 15.9 |
| Carbamazepine | 12 | 12.7 |
| Methimazol | 0.4 | -0.1 |
| Oleic Acid | 283 | 11.4 |
| Linoleic Acid | 280 | 0 |
| D-biotin | 0.001 | -7.3 |

The following table describes the cross-reactivity of the assay with substances that are similar in structure to T3. Following NCCLS C45A(12) recommendations the cross-reactants were tested in the absence of serum proteins (i.e. in HEPES pH 7.7 buffer containing 0.1% gelatin). The cross reactivities were calculated as the ratio between the concentration of T3 required to displace 50% of the conjugate (IC50) and the concentration of the cross-reactant required to give the same displacement, calculated on a molarity % basis.

| Substance | Cross-Reactivity (%) |
| --- | --- |
| Reverse T3 | < 0.5 |
| Tetraiodothyroacetic acid | < 0.5 |
| D-thyroxine | < 0.5 |
| L-thyroxine | < 0.1 |
| 3, 5 diiodothyronine | 3.3 |
| Diiodo-L-tyrosine | < 0.01 |
| Monoiodotyrosine | < 0.01 |
| 3-3’, 5-Triiodothyroacetic acid (TRIAC) | 13.0 |

Patients with biotin (vitamin H) treatment may have elevated free T3 levels due to displacement of the biotinylated T3 analog from paramagnetic particles coated with streptavidin.

- Result Calculation

TRUSTlab NRL follows estimation of CV%.

Interpret the results of this test in conjunction with the patient’s clinical presentation.

Serum and plasma values should not be used interchangeably. When monitoring a patient, it is suggested to rebaseline when changing sample types.

- Biological Reference Interval

| Age | Range |
| --- | --- |
| - | 2.5-3.9 pg/ml |

- Reportable Interval

- Clinical Reportable Range (CRR): 0.88 – 30.0 pg/ml

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

N/A.

- Laboratory Clinical Interpretation

Tridothyronine is a hormone synthesized and secreted from the Thyroid gland and formed by pheripheral deiodination of T4.

T3 and T4 are secreted in the circulation in response to Thyroid Stimulating Hormone (TSH) and play an important role in regulating metabolism.

Free T3 levels correlate with T3 secretion and metabolism.

In Hypothyroidism and Hyperthyroidism, free T3 levels parallel changes in total T3 levels.

Measuring free T3 is useful when altered levels of total T3 occur due to changes in T3 binding proteins especially Thyroxine Binding Globulin.

In the circulation 99.7% of T3 is reversibly bound to transport proteins (TBG) and to lesser extent Albumin and Prealbumin.

The unbound fraction of the total T3 is free Triodothyronine or FT3 metabolically active form.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of Free T3, Beckman Coulter DxI 800

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)