SOP – 060

PROCEDURE FOR

VITAMIN B12 TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	12

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	15

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	16

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | pH | Potential of Hydrogen |
| 15 | QC | Quality Control |
| 16 | QM | Quality Manager |
| 17 | QNS | Quantity Not Sufficient |
| 18 | SD | Standard Deviation |
| 19 | SOP | Standard Operating Procedure |
| 20 | SST | Serum Separator Tube |
| 21 | TRF | Test Requisition Form |
| 22 | TDPL | TRUSTlab Diagnostics Private Limited |
| 23 | ULOQ | Upper Limit of Quantitation |
| 24 | VB12 | Vitamin B12 |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Vitamin B12 in serum and plasma samples using Beckman Coulter CLIA analyzer. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Vitamin B12, also known as cobalamin, is an essential nutrient that's crucial for several bodily functions, including DNA synthesis, nerve cell function, and red blood cell formation.

- Test Principle

The Access Vitamin B12 assay is a competitive binding immunoenzymatic assay.

A sample is added to a reaction vessel along with alkaline potassium cyanide and dithiothreitol. This treatment denatures B12 binding proteins and converts all forms of vitamin B12 to the cyanocobalamin form.

After neutralization, intrinsic factor-alkaline phosphatase conjugate and paramagnetic particles coated with goat anti-mouse IgG: mouse monoclonal anti-intrinsic factor are added to the sample.

Vitamin B12 in the sample binds to the intrinsic factor conjugate, preventing the conjugate from binding to the solid phase anti-intrinsic factor.

After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer.

The photon production is inversely proportional to the concentration of vitamin B12 in the sample.

The amount of analyte in the sample is determined by means of a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 50 pg/ml

      - ULOQ: 1500 pg/ml

Analytical measurement range (AMR): 50 – 1500 pg/ml

Clinical Reportable Range (CRR): 50 – 1500 pg/ml

Sensitivity: 95%

Specificity: N/A

- Type of Sample

A minimum of 2 ml of non-hemolysed Serum/Plasma

      - Collect the blood sample by venipuncture observing universal precautions.

      - Allow samples to clot adequately before centrifugation.

      - Keep tubes stoppered and upright at all times.

      - Do not use samples that have been stored at room temperature for longer than 8 hours.

      - Tightly cap and refrigerate specimens at 2 to 8°C if the assay is not completed within 8 hours.

      - Freeze samples at or below -20°C if the sample is not assayed within 48 hours.

      - Freeze samples only once and mix thoroughly after thawing.

      - Before placing samples on the system ensure that:

        - Samples are free of fibrin or other particulate matter.

        - Samples are free of bubbles.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

Do not add any additives/preservatives.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

- Access Vitamin B12 Reagent Pack determinations, 2 packs, 50 tests/pack

- Provided ready to use.

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-10°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the DxI ACCESS 2 Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

For a detailed discussion of handling precautions during system operation, refer to the ACCESS 2 System Operations Manual.

If the reagent pack is damaged (i.e. broken elastomer), discard the pack.

If the microparticle bottle does not remain upright (with a septum installed) while in refrigerated storage off the system, the reagent kit must be discarded.

- Calibration Procedure

Calibrator Required

      - 6 Point Calibrator

- Calibration Schedule

      - An active calibration curve is required for all tests

      - Recalibrate the assay every 21 days.

      - Additionally, the Beckman Coulter Vitamin B12 assay requires a two-point calibration when:

- changing lot numbers of primary reagent packs

- replacing critical system components due to service or maintenance issues

- preventive maintenance has been performed

- quality control results are repeatedly out of acceptable range

- LJ plots show positive or negative bias for 10 data points

- Calibration or calibration-verification has failed.

Preparation of Calibrators

      - Calibrator material should be at room temperature before reconstitution. Keep the vials at room temperature for approximately 2 Hours, after taking out from the Deep freezer.

      - Slowly invert the calibrator or diluent bottles several times to mix the contents thoroughly. DO NOT SHAKE.

      - Gently tap the lyophilized calibrator vial on the counter several times to dislodge any material adhering to the stopper.

      - Remove the seal and stopper from each bottle just prior to the addition of the diluent. Do not leave vials unstopped.

      - Reconstitute by pipetting exactly 3ml of distilled water at +20 to +25°C, into the vial.

      - Replace the stopper and hold it firmly in place. Invert the vial gently. DO NOT SHAKE.   Visually verify that all freeze-dried material is dissolved prior to use

Storage and Stability of Calibrators

      - Un-reconstituted Calibrator is stable up to the expiry date shown on the side of each individual bottle.

      - Once reconstituted, the components of the Calibration Sera are stable for 8 hours at +25°C, for 7 days at +4°C and for 1 month at -20°C (frozen).

      - NOTE: Do not use the reconstitute calibrator beyond stability period.

- Test Procedure

Refer to the appropriate system manuals and/or Help system for a specific description of installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

Mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use forty-five (45) µL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) µL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

For detailed information on ordering dilutions, refer to the DXI System Operations Manual, Section 5.

Handling of repeat test results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10% or 1.5 times the parameter CV% (whichever is higher).

      - If the percentage difference between the values is greater than 10% or 1.5 times the parameter CV% (whichever is higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

- Biorad - Level 1 and Level 2, Two levels of Quality Control sera i.e. normal and high run daily as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity (frothing to be avoided).

- Visually verify that all freeze-dried material is dissolved prior to use.

- Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20°C, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8°C refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

Storage & Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2-8°C.

      - Reconstituted: Stable for 5 days when stored, tightly capped, at 2-8°C.

      - Longer storage: -20°C

      - Aliquots are stored at the above temperatures as required.

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

Samples containing up to 10 mg/dl (171 µmol/l) bilirubin, 9 g/dl (90 g/l) total protein, and lipemic samples containing the equivalent of 1,800 mg/dl (20.32 mmol/l) triglycerides do not affect the concentration of vitamin B12 assayed.

- Result Calculation

- TRUSTlab NRL follows the estimation of %CV.

- Biological Reference Interval

- 120 to 914 pg/ml

- Reportable Interval

Analytical Measurement Range (AMR): 50–1500 pg/ml

Clinical Reportable Range (CRR): 50–1500 pg/ml

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the clinician/doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Vitamin B12 deficiency can occur for one of several reasons. The most common cause is a defect in the secretion of intrinsic factor, resulting in inadequate vitamin B12 absorption from foods. This condition is called pernicious anemia and is most common in people over age 50. Other causes of vitamin B12 deficiency are gastrectomy, malabsorption due to surgical resections, and a variety of bacterial or inflammatory diseases affecting the small intestine. The amount of vitamin B12 absorbed is directly proportional to the length of functional intestine. Vitamin B12 deficiency due to insufficient dietary intake is rare and can occur only after years of abstinence from all animal products.

Elevated levels of vitamin B12 have been associated with pregnancy, the use of oral contraceptives and multivitamins, and in myeloproliferative diseases such as chronic granulocytic leukemia and myelomonocytic leukemia. An elevated vitamin B12 level in itself has not been known to cause clinical problems.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition.

      - Product insert of Vitamin B12 (VB12), Beckman Coulter DxI 800.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).