SOP – 061

PROCEDURE FOR QUALITY ASSURANCE

IN BIOCHEMISTRY

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Input	7

5.	Output	7

6.	Monitoring & Measurement	7

7.	Procedure	8

8.	References	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CM | Chylomicrons |
| 5 | CV | Coefficient of Variation |
| 6 | CRP | C-Reactive Protein |
| 7 | CRR | Clinical Reportable Range |
| 8 | EDTA | Ethylenediaminetetraacetic Acid |
| 9 | EQAS | External Quality Assurance Scheme |
| 10 | HOD | Head of Department |
| 11 | IFCC | International Federation for Clinical Chemistry |
| 12 | IQC | Internal Quality Control |
| 13 | LLOQ | Lower Limit of Quantitation |
| 14 | MSP | Management System Procedure |
| 15 | N/A | Not Applicable/Available |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | RPPHS | Reference Preparation for Proteins in Human Serum |
| 20 | SD | Standard Deviation |
| 21 | SOP | Standard Operating Procedure |
| 22 | SST | Serum Separator Tube |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |
| 26 |  |  |

- Objective

- The purpose of this SOP is to is to define and set up guidelines for the various quality control activities to be carried out by the staff of the Biochemistry department daily & periodically and the various records and registers to be maintained pertaining to Quality control activities of Clinical Biochemistry department.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Input

Internal QC samples

Calibrators/ Standards purchased

QC Samples from the External agency

Test equipment, kits, chemicals and reagents

- Output

Internal QC results

External QC results

L-J charts

- Monitoring & Measurement

Internal & External QC Records.

Standard Deviation, Coefficient of variation, Measurement Uncertainty on Internal QC results.

Feedback on external QC results from the external agency.

- Procedure

The QM ensures that all technicians of the department are knowledgeable about the contents of departmental manual & procedures including changes relevant to the scope of their testing activities.

The QM ensures that only the current procedure manual of the department is available on the worktable and accessible at all times to all technicians of the department.

Reagents, Chemicals & Equipment:

      - Certified & standard glassware of ‘A’ quality are used. Only approved & standard glassware, analytical grade chemicals, approved kits and reagents are used for testing.

      - Whenever reagents are prepared in-house the name & concentration of the reagent, date of preparation of the reagent, storage conditions, e.g temperature, and validity period are clearly indicated on the reagent bottle.

      - If the reagent is not prepared in-house the date of opening is clearly indicated on the reagent/ kit/ product.

      - The chemicals, reagents and kits are stored at the temperature and humidity conditions specified in the product insert or as specified by the manufacturer.

      - Whenever any lyophilized chemical/ reagent is reconstituted, the date of reconstitution is marked on the vial and the reconstituted reagent is not used beyond that mentioned in the kit insert.

      - Only the de-ionized water supplied along with the kit/ reagent or reagent Grade Type I Water is used for reconstitution.

      - The expiry date of the reagents and kits are checked before use and not used them beyond the expiry date.

      - Micropipettes are calibrated once in six months and Temperature measuring devices as mentioned in Equipment calibration Plan of Annexure III and the details are recorded and the documents are maintained.

Pre-analytical Quality Control:

      - Samples are collected in appropriate tubes and containers.

      - Only plain vacutainer tubes are used, if serum is the specimen/ appropriate anti-coagulant containing vacutainer tubes are used, if plasma is the specimen, in appropriate concentration to ensure proper anti-coagulant – blood ratio while collecting blood samples for various tests.

      - The blood is allowed to stand at room temperature for 30 minutes to form a complete clot and then the serum is separated by centrifuging at 3500 rpm for 10 minutes.

      - It is ensured that there are no fibrin clots in the specimen which may block the sample probe.

      - Samples are rejected and not accepted for processing under the following conditions:

        - Samples without a Test request form.

        - Samples, which are clotted or samples containing foreign materials and visible fibrin threads.

        - Samples which are insufficient for testing.

        - Samples without an ID/ barcode labeling/ not accompanied by Requisition Form.

        - Samples collected in improper vacutainer tubes.

        - Wrong specimen type (Example: serum instead of whole blood)

        - Specimens that have not been properly transported and handled.

        - Not received in proper temperature condition.

        - Hemolysed sample / Lipemic / contaminated sample.

        - Samples collected in leaking/ broken tubes

Internal QC testing:

      - Internal QC samples are run daily at the beginning of the day before processing patient samples.

      - Any one level of Normal & Abnormal levels of an assayed QC sample are run daily for all the analytes in the morning before processing patient samples.

      - Lyophilized vial of an assayed QC sample level I & level II for clinical chemistry and Level I, II and III for immunoassay is reconstituted with distilled water and the reconstituted QC sample is aliquoted into vials, with the date of reconstitution marked and stored in the freezer at 0C.

      - The QC aliquot once taken out from the freezer is not reused.

      - The QM or TM ensure that the Internal & External QC specimens are tested in the same manner as patient specimens and analyzed by technicians who routinely perform patient testing.

      - QC is run based on the load of the patient sample. At least two levels of QC are run daily.

        - For Routine Clinical Chemistry –Level I & Level II.

        - For Hormones – Two levels as per IQC programme recommended by the laboratory based on the load.

To induct new lot QC:

      - Calibrate the parameter before running the QC

      - Mean, Standard deviation and Coefficient of Variation of the QC values are calculated once in 30 days based on the kit insert (QC reference).

      - Lab Mean is established after running the QC for 30 days

      - Run QC (New lot) in Parallel with Old Lot.

      - A total allowable error up to 10% is set as the acceptance criteria for the new lot.

Lot to Lot Verification:

      - Parallel testing is to be done when lot Number Changes.

      - Document all the run results.

      - Abnormal samples which are tested on the previous day with the old lot are taken.

      - The new lot reagent is tested with the same sample.

      - The results are noted for accuracy and documented.

      - A total allowable error of up to 10% is set as the acceptance criteria.

Accepting/Rejecting a Quality Control run

      - QC values are accepted if they are within the target range and the standard deviations obey the QC rules mentioned below.

      - The rules followed when one level QC material is used are:

        - QC is rejected if

          - it is outside 3 SD (13s)

          - two consecutive values obtained are outside 2 SD on the same side but within 3 SD (22s)

          - ten consecutive values are above or below the mean, but within 2 SD (10x)

      - The rules followed when 2 level QC materials are used:

        - QC is rejected if

          - either QC value is outside 3 SD (13s)

          - both QC values are outside 2 SD on the same side, but within 3 SD (22s)

          - The difference between both QC values is >4 SD i.e. one level QC is > 2 SD and other level QC is <2SD (R4s).

        - Warning Criteria

          - The 12s rule is usually a warning rule violated when a single control observation is outside the ±2SD limit.

Rejecting a Quality Control run

      - It is checked whether the QC value is within the acceptable range and close to the target mean specified by the manufacturer or product insert.

      - If the QC is not within the acceptable range specified and if recalibration is required, the instrument is calibrated using calibrators specified by manufacturers and the QC is performed again till the QC value is close to the target mean and within the specified range.

Steps to be taken for Out–of–control situation:

      - If the QC is outside the range even after QC rerun, appropriate corrective and preventive actions are taken such as:

        - Checking the Expiry date of the reagent and QC material.

        - Checking the calibration status for the analyte.

        - Changing to a new reagent of the from the same lot.

        - Changing to a new lot of reagent.

        - Changing QC material.

        - Recalibration for the analyte in the instrument using the calibrator.

      - The results are entered into the respective Internal Quality Control Register.

      - The details of out-of-range QC values and the action taken are recorded in the Internal QC Register after approval from the Biochemist/Quality Manager.

      - Patient samples are processed only after the Biochemist approves the QC results.

Levey Jennings Charts & Measurement Uncertainty for IQC:

      - Mean, Standard deviation and Coefficient of Variation of the QC values are calculated once in 30 days.

      - LJ Chart is plotted every month for each parameter and for each level of Internal Quality Control run for parameters in the scope of accreditation.

      - The Measurement Uncertainty (MU) for the testing is calculated by 1.96 ± CV or 1.96 ± SD at a 95% confidence interval.

External Quality Control Testing:

      - External Quality Control/ PT programme is performed once a month by analyzing a sample received from the External agency like CMC or Metropolis.

      - The External QC material is processed along with patient samples and by technicians who normally perform the testing. No special treatment is given to EQAS samples. Also, EQAS/PT samples are introduced along with routine samples received for testing and not run as separate samples.

      - The results are recorded along with the machine printout in the EQAS Register.

      - The results are entered in the prescribed form received from the external agency (online form) and sent to the External Agency for evaluation of results after getting approval from the QM or HOD.

      - A copy of the EQAS sample test results sent by the laboratory is retained.

      - The feedback received from the external agency is filed in the External Quality Control Register after getting signature of the QM or HOD.

      - Actions are taken as per the instructions of the Biochemist in the Remarks column.

Inter-laboratory Comparison Testing:

      - The laboratory participates in Inter-Laboratory Comparison with good medical laboratories for parameters in the scope of testing, which doesn’t have any External Quality Assurance Programme for participation.

      - The laboratory sends its sample to other good laboratories whenever the results are felt to be out of range by the QM or TM or whenever requested by the referring physician.

      - The original results of the laboratory and the results obtained from the Referral laboratory (laboratory where the sample is sent for verification) are filed in the inter-laboratory comparison register maintained in the laboratory.

      - The outcome of the results is checked and verified and if there is any undue deviation between the results of the two laboratories, necessary corrective and preventive action is taken under the guidance of QM.

Split Sample Testing:

      - The laboratory conducts split sample testing to determine the Quality Control for parameters in the scope of testing, which doesn’t have any External Quality Assurance Programme for participation or Inter-laboratory comparison is not possible.

      - The same sample is given to two different technical staff who routinely perform the test.

Criteria for discarding the reagents, control and calibrator

      - Note down the expiry date, as soon as the kit arrives in the laboratory.

      - Kits that exceed the stated date of expiry should be discarded (or) stored for demo purposes to train the beginners, with the purpose of storage labeled on the reagent container.

      - Contaminated kits are strictly disposed.

      - If the reagent, calibrator is not working properly, it is strictly disposed.

      - The date of opening the reagent also should be labeled over the container.

- References

the NABL112, “Specific criteria for accreditation of medical laboratories” appropriate to the examination and interpretation of the examination results.

CLSI Guidelines-2021