SOP – 006

PROCEDURE FOR

URINE BILE SALTS

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	7

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	9

15.	Interferences & Cross Reactions	9

16.	Result Calculation/Interpretation	9

17.	Biological Reference Interval	9

18.	Reportable Interval	9

19.	Procedure when Results are not in Range	10

20.	Critical Alert Values	10

21.	Laboratory Clinical Interpretation	10

22.	Potential Source of Variability	10

23.	References	10

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | CP | Clinical Pathology |
| 3 | CV | Coefficient of Variation |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HBV | Hepatitis B Virus |
| 6 | HIV | Human Immunodeficiency Virus |
| 7 | HOD | Head of Department |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | pH | Potential of Hydrogen |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the manual method of qualitative detection of bile salts by and is called Hay’s test. This SOP also provides information about the test principles, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Clinical Pathology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Clinical Pathology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Bile salts in urine are not normally present, and their presence indicates potential liver problems or issues with bile flow. A urine test for bile salts helps detect these issues.

- Test Principle

- Bile salts, when present in urine, lower the surface tension of urine. When sulfur      powder (flowers of sulfur) is sprinkled on to the surface of urine, the sulfur powder sinks to the bottom of the test tube if the urine sample contains bile salts.  In normal urine the sulfur powder floats on the surface of urine.

- Specific Performance Characteristics

- N/A.

- Type of Sample

Urine is the primary specimen for the test.

Collect a minimum 15 ml of a random sample of urine.

No preservative should be added to the urine.

Process the samples within 1 hour of collection. If delay is anticipated instruct patients to store the sample at 2-8°C in the refrigerator for a maximum of up to 4 hours until transport to the laboratory.

Use fresh and un-centrifuged samples for the test.

Do not accept if the sample volume is inadequate or if samples are received without an ID number.

- Type of Container & Additives

- Sterile screw-capped container.

- Patient Preparation

Instruct patients to collect urine samples in a clean dry container (preferably one provided by TRUSTlab).

Instruct patients to void urine directly into the container and while collecting allow the first portion of urine to escape.

A random sample is sufficient No fasting any other preparation is required.

- Equipment & Reagents

Equipment

- Glass Test tubes

Reagents

- Sulfur powder

- Safety

Follow universal precautions like wearing suitable gloves and eye/face protection while handling all reagents, samples and operating the instrument.

Universal precautions imply that all body fluids are treated as if known to be infectious for Human Immunodeficiency Virus (HIV), Hepatitis B Virus (HBV) or other Blood borne Pathogens.

Dispose of all waste in accordance with applicable national and/or local regulations.

Some reagents contain sodium azide, which may react with copper or lead plumbing to form explosive metal azides. Use caution in disposing of these reagents. If disposing to drain, flush with large volumes of water to prevent azide buildup.

Waste material containing patient samples or biological products should be considered biohazardous when disposing or treating.

Clean up all spills immediately and thoroughly. Disinfect the area for any spills involving biohazardous materials. Dispose of all contaminated materials appropriately.

- Calibration Procedure

- N/A.

- Test Procedure

      - Take 10 ml of clear urine in a test tube. (4 finger breadths)

      - Sprinkle a little dry sulfur powder on to the surface of the urine using a spatula or spoon.

      - Observe the sulfur particles.

- Quality Control Procedure

- As defined in the procedure Urine IQC (TDPL/CP/SOP-002).

- Interferences & Cross Reactions

- N/A.

- Result Calculation/Interpretation

If Sulfur powder sinks to the bottom of the test tube, it indicates the presence of bile salts in the urine sample

If Sulfur powder does not sink to the bottom of the test tube but floats on the surface of urine, it indicates the absence of bile salts.

- Biological Reference Interval

- Negative.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- This test is useful in the differential diagnosis of Jaundice. The presence of Bile salts in urine indicates obstruction to the biliary passages which may be either intra hepatic or extra hepatic commonly due to gall stones anywhere in the bile passages, carcinoma of head of pancreas and stricture of common bile duct.

- Potential Source of Variability

Even a slight air current will cause the sulfur powder to sink. Hence ensure that the fan is switched off while performing the test.

Contamination of the test tube with soap or hypochlorite will cause the sulfur powder to sink even if bile salts are absent.

- References.

      - Handbook of Medical Laboratory Technology 2nd Editon, Edited by Robert H. Caraman, M.D. Laboratory Training Committee, Christian Medical Association of India, Page: 240

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)

      - NABL 112