SOP – 001

PROCEDURE FOR

BLOOD GROUPING AND RH TYPE TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	13

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	14

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | °C | Degree Centigrade |
| 3 | CV | Coefficient of Variation |
| 4 | CRR | Clinical Reportable Range |
| 5 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 6 | EQAS | External Quality Assurance Scheme |
| 7 | HOD | Head of Department |
| 8 | HM | Haematology |
| 9 | IgM | Immunoglobulin M |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | NS | Normal Saline |
| 15 | OH | Hydroxy |
| 16 | pH | Potential of Hydrogen |
| 17 | QC | Quality Control |
| 18 | QM | Quality Manager |
| 19 | QNS | Quantity Not Sufficient |
| 20 | SD | Standard Deviation |
| 21 | SOP | Standard Operating Procedure |
| 22 | SST | Serum Separator Tube |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |
| 26 |  |  |

- Objective

- The purpose of this SOP is to describe the procedure for performing blood grouping and Rh Typing. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Blood grouping refers to the classification of blood into different types based on the presence or absence of specific antigens on the surface of red blood cells. The most common systems used for blood grouping are the ABO system and the Rh system.

- Test Principle

ABO blood groups are determined by phenotyping the patient’s red cells for the presence or absence of A and B antigens and by testing the patient’s serum for the presence or absence of anti-A and anti-B.

Method: Agglutination (Forward grouping for ABO and Rh)

- Specific Performance Characteristics

Limit of Quantitation: N/A

Analytical measurement range (AMR): N/A

Clinical Reportable Range (CRR): N/A

Sensitivity: N/A

Specificity: N/A

- Type of Sample

A minimum of 3 ml of whole blood

      - Collect the blood sample by venipuncture observing universal precautions.

      - Allow samples to clot adequately before centrifugation.

      - Keep tubes stoppered and upright at all times.

      - Do not use samples that have been stored at room temperature for longer than 8 hours.

      - Tightly cap and refrigerate specimens at 2 to 8°C if the assay is not completed within 8 hours.

      - Freeze samples at or below -20°C if the sample is not assayed within 48 hours.

      - Freeze samples only once and mix thoroughly after thawing.

      - Before placing samples on the system ensure that:

        - Samples are free of fibrin or other particulate matter.

        - Samples are free of bubbles.

Sample Stability: 48 hours at 2 – 8℃.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

- K2 EDTA anticoagulant vacutainer.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

      - Microscope

      - Micro-pipette 50-100 µl

Reagents

      - ABO Forward Grouping:

          - Monoclonal antisera (Anti-A, Anti-B, Anti-H)

          - Slides

      - Rh D Phenotyping

          - Monoclonal Anti-D(IgM)

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-10°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

False positive or negative results can be caused by variables such as:

      - Improper technique

      - Contaminated materials

      - Omission of reagents or anti-sera

      - Delays in testing

      - Inadequate incubation time and temperature

      - Inappropriate centrifugation

      - Inappropriate or prolonged storage of red cell

- Calibration Procedure

- N/A.

- Test Procedure

Bench Preparation (for ABO Grouping)

      - Allow all reagents and patient samples to come to Room Temperature.

      - Identify reagents to be used for forward grouping.

      - Separation of Red Blood Cells & Plasma.

      - Centrifuge at 1000 RPM for 1 min, using clean pipette tip, aspirate plasma, gently without disturbing settled cells and transfer to a labelled clean test tube for reverse grouping.

Forward Grouping

      - Add 50 µl each of Anti-A, Anti-B in respective labelled slides.

      - Add 50µl of whole blood to each antiserum in above slides.

Bench Preparation (Rh Typing)

      - Prepare bench for performing the grouping.

      - Allow all reagents and patient samples to come to Room Temperature.

Rh D – Typing & Phenotyping

      - Take 50µl of Anti-D in in respective labelled slide.

      - Read agglutination & note the results in Blood group record.

      - Confirm all negative results under microscope & proceed for Indirect Antiglobulin Test (IAT).

      - Perform the test using Anti-D(IgM) if negative perform using Anti-D (IgM+IgG), still negative, process the sample for DU test as mentioned below.

Du Test procedure (To be performed for all Rh-negative samples)

      - Prepare a 5% suspension of red cells to be tested in isotonic saline.

      - Add one drop of Anti-D(IgM) into the (T)-labelled test tube.

      - Take another test tube and label it as (C) and add one drop of 22% bovine albumin.

      - Add one drop of 5% cell suspension to each of the two test tubes and mix well. Incubate at 37°C for 15 minutes.

      - Wash the contents of both tubes thoroughly, at least three times, with isotonic saline and decant completely after the last wash.

      - Add two drops of AHG reagent and mix well in both tubes.

      - Centrifuge for 1 minute at 1000 rpm.

      - Very gently, re-suspend the cell button, observe for agglutination macroscopically, as well as microscopically a slide.

      - Agglutination indicates the presence of Du antigen. No agglutination indicates the absence of the Du antigen.

- Quality Control Procedure

Internal Quality Control

    - Procedure:

      - All reagents shall be used and controlled according to the manufacturer’s written instructions.

      - All anti-sera must be visually inspected for contamination such as discoloration, cloudiness, turbidity and/or particulate matter.

      - All reagent red cells must be visually inspected for haemolysis and/or discoloration.

      - The results of the visual inspection, reagent lot number, expiry date, date of the inspection and the individual performing the inspection is documented in the Daily IQC Data Monitoring Form (TDPL/GE/FOM-012).

      - The expiry date should be checked of each reagent used. Do not use reagents beyond their expiry date.

      - The reactivity of Blood Grouping Reagents shall be confirmed each day of use by control tests with known antigen positive and negative red cells.

      - Positive control cells should be selected to represent weak expression of the specific antigen.

      - Historical record check- An attempt is made to check for previous results before releasing blood group result. If any discrepancy (Typographical Error) is observed for previous results, investigations are performed and amendment report issued after re-confirmation of present sample.

      - Reagents are to be assessed for:

          - REACTIVITY: Ability of a reagent to react with the corresponding antigen.

          - SPECIFICITY: Test reagent against known positive and negative cells.

          - AVIDITY: Time interval between first rapid mixing of the 2 drops and beginning of macroscopic agglutination.

          - SENSITIVITY: Titration of antisera denotes sensitivity.

    - Avidity And Specificity of Antisera:

| Antisera | Minimum Titre (2-3% Red Cell Suspension) | Minimum Titre (5-minute Incubation and Spin) | Avidity | Reaction Grade (Agglutination) |
| --- | --- | --- | --- | --- |
| Anti A | A1 Cells | 1:256 | 3 - 4 sec | +++ |
| Anti B | B Cells | 1:256 | 3 - 4 sec | +++ |
| Anti AB | A1 Cells | 1:256 | 3 - 4 sec | +++ |
|  | B Cells | 1: 128 | 3 - 4 sec | +++ |
| Anti-D (IgM) Blend Monoclonal | D Cells | 1:128 – 1:256 | 5 - 10 sec | ++++ |

    - Test for Avidity

      - Prepare 50% saline suspension of test cells.

      - Place 2 drops of blood grouping antiserum that is to be checked on a glass slide.

      - Place 1 drop of cell suspension next to it.

      - Mix the 2 drops and spread out in a circle about 15 mm in diameter.

      - Rotate the slide slowly from side to side.

      - Record the time in seconds required for clumps (1 mm) to appear.

      - Perform each test twice and ensure similar results within a limit of 3 seconds.

    - Doubling Dilution Method Titration

      - Principle: Titration is a semi quantitative technique of measuring the concentration of an antibody in a serum which is usually determined by testing two-fold serial dilution of the serum in saline against selected red cells.

      - Method:

        - Arrange & label ten tubes in a tube rack. Label to 10 (serum dilutions 1: 1 of 1:512).

        - Put I-volume (2 drops saline) in all tubes except the first.

        - In the first & second tube put 2 drops serum (I volume)

        - Use clean pipette, mix contents of tube 2.

        - Transfer I-volume (2 drops) from tube 2 to tube 3 to make a dilution 1:4

        - Continue the process till tube 10.

        - Remove I-volume from the last tube. Do not discard. Save it for use if further dilutions are needed.

        - Add I-volume (2-5% saline suspended) red cells according to antisera in each tube (e.g. A cells when anti A is used).

        - Centrifuge all tubes at 1000 rpm for 1 minute.

        - Record results. Gently dislodge the button.

External Quality Control

- EQAS participation –- CMC EQAS

- Interferences & Cross Reactions

- N/A.

- Result Calculation

| Forward Grouping |  |  | Result |
| --- | --- | --- | --- |
| Anti-A | Anti-B | Anti AB |  |
| + | - | + | A |
| - | + | + | B |
| - | - | - | O |
| + | + | + | AB |
| - | - | - | OH or any other irregular antibody |

- Note:  + => Agglutination		- => No Agglutination

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Derives from the development of antibodies against these antigens

Autoantibodies (against self-antigens)

Alloantibodies (against foreign antigens)

These antibodies may cause several pathological processes:

Destruction of native red cells (Autoimmune haemolytic anemia)

      - Destruction of transfused red cells (Haemolytic transfusion reactions)

      - Destruction of foetal red cells by maternal antibodies which traverse the placental barrier (Haemolytic disease of the Newborn).

- Potential Source of Variability

- False positive or negative results can be caused by variables such as:

      - Improper technique

      - Contaminated materials

      - Omission of reagents or anti-sera

      - Delays in testing

      - Inadequate incubation time and temperature

      - Inappropriate centrifugation

      - Inappropriate or prolonged storage of red cell

- References

      - Practical Hematology, Dacie and Lewis, Tenth edition, Page no 501.

      - Manufacturer’s kit inserts.

      - Quality Control of ABO and Rh blood grouping reagents -Document ID   No. National Institute of Biologicals, (Ministry Health & Family Welfare), Government of India BRL /GM/01 Effective Date 30.11.2012.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).