SOP – 006

PROCEDURE FOR

PROTHROMBIN TIME

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	10

16.	Result Calculation	10

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | HOD | Head of Department |
| 5 | HM | Haematology |
| 6 | INR | International Normalized Ratio |
| 7 | IQC | Internal Quality Control |
| 8 | ISI | International Sensitivity Index of the reagent |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | PPP | Platelet Poor Plasma |
| 12 | PT | Prothrombin Time |
| 13 | QC | Quality Control |
| 14 | QM | Quality Manager |
| 15 | QNS | Quantity Not Sufficient |
| 16 | rpm | Revolutions per minute |
| 17 | SOP | Standard Operating Procedure |
| 18 | TRF | Test Requisition Form |
| 19 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the procedure for performing the Prothrombin Time (PT) Test by a conventional routine method using citrated, platelet-poor plasma. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Prothrombin time (PT) is a blood test that measures the time it takes for the liquid portion (plasma) of your blood to clot.

- Test Principle

Tissue Thromboplastin in the presence of calcium activates the extrinsic pathway of human blood coagulation mechanism.

When UNIPLASTIN reagent is added to normal anticoagulated plasma, the clotting mechanism is initiated, forming a solid gel clot within a specified period of time. The time required for clot formation will be prolonged if there is a deficiency of factors/factor activity in the extrinsic pathway of the coagulation mechanism.

Method: Opto-mechanical.

- Specific Performance Characteristics

Linearity: N/A

Measurement Range: Up to 120 sec.

Sensitivity: 0-120 sec.

Clinical Reportable Range: 10-120 sec.

- Type of Sample

Plasma - Trisodium citrate Sample (3.2%)

Draw blood without undue venous stasis and without frothing into a plastic syringe fitted with short needle of 19 to 20 SWG. The vein puncture must be a clean one. Transfer the blood to the tubes, after detaching the needle from the syringe.

Mix exactly nine parts of freshly collected blood with one part of Tri-sodium citrate.

Centrifuge immediately for 15 minutes at 3500 rpm and transfer the plasma into a clean test tube

Sample Stability: Plasma must be tested preferably immediately. The samples are stable only for:

      - 2 hours at 22-24°C

      - 3 hours at 2 -8°C

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

- Sodium Citrate (3.2%) vacutainer.

- Patient Preparation

Patients should not have exercised heavily before blood collection.

Fasting or only non-fatty meals prior to blood collection is preferred as this results in procuring samples with a desirable lower opacity.

- Equipment & Reagents

Equipment

      - Centrifuge

      - Water bath

      - Stopwatch

      - Thermometer

      - Pipette (20 µl – 200 µl)

Reagents

      - Blood in 3.2% Sodium citrate

      - UNIPLASTIN

      - Plastic Test Tube (12x75mm)

- Safety

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

- Calibration Procedure

- N/A.

- Test Procedure

      - Bring the reagent vial to room temperature (20-37°C). Mix the contents of the vial to homogenize the suspension completely.

      - Aspirate form the reagent vial enough reagent in a thoroughly clean and dry test tube.

      - Prewarm the reagent and bring to 37°C before use in the test procedure.

      - Recap the reagent vial and place it immediately in the 2-8°C refrigerator.

      - To a 12x75 mm tube add 0.1 ml plasma (PPP) and place the tube in a prewarmer for 3-5 minutes at 37°C.

      - To the tube forcibly add 0.2 ml of UNIPLASTIN reagent (prewarmed at 37°C for at least 3 minutes) and simultaneously start a stopwatch. Shake the tube gently to mix contents.

      - Gently tilt the tube back and forth and stop the stopwatch as soon as the first strand is visible and the gel/clot formation begins. Record the time in seconds.

      - Repeat steps a-f for a duplicate test on the same sample.

      - Find the average of the duplicate test values. This is the required Prothrombin Time (PT).

      - If a coagulation instrument is being used to perform the test, the instrument manufacturer’s instructions must be strictly adhered to.

- Quality Control Procedure

Internal Quality Control

    - Procedure:

- It is recommended by WHO that MNPT should be established for each lot of PT reagents    by each laboratory, since PT results are dependent on the combination of reagent lot, instrument and technique followed at each laboratory. Usually, plasma from at least 20 normal healthy individuals should be used to establish the MNPT. The average of such PT results in seconds = MNPT.

External Quality Control

- EQAS participation –- CMC EQAS

- Interferences & Cross Reactions

- Turbid, icteric, lipemic or grossly hemolysed samples may generate erroneous PT results.

- Result Calculation

- Manual Method:

- The result may be reported directly in terms of the mean of PT of the test plasma in seconds or as a ratio 'R':

- R=Mean of the patient plasma PT (in seconds)MNPT for the reagent

- or as an International Normalized Ratio (INR):

- INR= RISI

- Where ISI = International Sensitivity Index of the reagent.

- Biological Reference Interval

Normal values using UNIPLASTIN are between 11.0-15.0 seconds. Between manual and Turbo densitometric instrument results a variation of 1-2 seconds may be expected.

Oral Anticoagulant Therapeutic range: INR: 2.0-3.5

- Reportable Interval

- 14, 16, 30 seconds

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- > 70 seconds

- Laboratory Clinical Interpretation

Abnormal Prothrombin time test results doesn’t diagnose any particular disease. It only provides insight about the time it takes for your blood to clot. Multiple diseases and conditions can cause abnormal results.

A prolonged Prothrombin time test result may be due to:

      - Hemophilia A or B

      - Deficiency of Blood Clotting Factors

      - Von Willebrand Disease (A disorder that causes abnormal blood clotting)

      - DIC - Disseminated intravascular Coagulation (a disease in which the proteins responsible for blood clotting are abnormally active)

      - Hypofibrinogenemia (deficiency of the blood clotting factor Fibrinogen)

      - Certain Medications, such as the blood thinners Heparin and Warfarin

      - Nutritional Issues, such as Vitamin K deficiency and malabsorption

      - Antibodies, including Cardiolipin Antibodies

      - Lupus Anticoagulants

      - Leukemia

      - Liver Diseases.

- Potential Source of Variability

- Incorrect mixture of blood and tri-sodium citrate, insufficient prewarming of plasma and reagent, contaminated reagents, glassware etc. are potential sources of errors.

- References

      - Practical Hematology, J. Dacie and M. Lewis, Tenth edition.

      - Uniplastin Kit insert - Tulip Diagnostics (P) Ltd.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).