SOP – 011

PROCEDURE FOR

LEISHMAN STAINING

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	9

15.	Interferences & Cross Reactions	10

16.	Result Calculation	10

17.	Biological Reference Interval	10

18.	Reportable Interval	10

19.	Procedure when Results are not in Range	10

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	11

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | ESR | Erythrocyte Sedimentation Rate |
| 5 | HOD | Head of Department |
| 6 | HM | Haematology |
| 7 | INR | International Normalized Ratio |
| 8 | IQC | Internal Quality Control |
| 9 | ISI | International Sensitivity Index of the reagent |
| 10 | MSP | Management System Procedure |
| 11 | N/A | Not Applicable/Available |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | rpm | Revolutions per minute |
| 16 | SOP | Standard Operating Procedure |
| 17 | TRF | Test Requisition Form |
| 18 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the procedure for performing the microscopic examination of blood smear after Leishman’s stain. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Leishman stain is a microscopy stain primarily used for blood smears to differentiate and identify various blood cells, malaria parasites, and trypanosomes. It is a Romanowsky-type neutral dye stain based on a methanolic mixture of polychromed methylene blue and eosin, also known as eosin-polychrome methylene blue.

- Test Principle

Leishman Stain is a mixture of Methylene Blue and derivatives complexed by Eosin. Leishman Solution fixes blood smears by the methyl alcohol content. As a neutral stain in an alcoholic medium, Leishman Solution releases its staining action only after the addition of buffered solution. They include the following:

Method: Staining and Microscopic examination

- Specific Performance Characteristics

- N/A.

- Type of Sample

3 ml Whole Blood

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

K2 EDTA Vacutainer

- Patient Preparation

No special preparation.

- Equipment & Reagents

Equipment

      - Microscope

      - Glass slides

Reagents

      - Ready to use Leishman stain, along with commercially available buffer sachet.

      - Distilled water

- Safety

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

Store Leishman’s Stain solution at RT in tightly sealed bottles, protected from light and heat. Do not store the solution near bottles containing acid. If stored correctly Leishman’s Stain solution is stable for approximately 3 months.

- Calibration Procedure

- N/A.

- Test Procedure

      - Use smears that are as thin as possible and air-dried.

      - Prepare the smear and label the smeared area with patient Vial ID (Use Lead pencil for labeling)

      - Fully cover the smears with Leishman Stain solution. Stain for 2 minutes.

      - Add twice the amount of distilled water and mix by swirling. Keep it for at least 10 min.

      - Rinse thoroughly with distilled water.

      - Dry the slides using blotting paper and air-dry.

      - Label the smear with patient name (Use marker pen).

      - Buffer Preparation:

        - Add one sachet of buffer to 1 litre of Distilled water.

        - Mix by gentle swirling.

        - Label the bottle of buffered water and document the procedure in the Leishman Stain QC Register (TDPL/HM/REG-002).

- Quality Control Procedure

Internal Quality Control

      - Prepare one slide daily using a patient sample. Stain the slide as indicated in procedure Section 13. Review the slide for Colour, Precipitation and Contamination.

      - If the colour does not meet the specifications identified in the SOP as indicated above or precipitation and/or contamination are present, the quality is determined to be unsatisfactory.   Indicate unsatisfactory in the Leishman Stain QC Register (TDPL/HM/REG-002) and replace the stain.  Document the problem and corrective action in the appropriate section of the QC form.  If the stain is determined to be satisfactory, indicate satisfactory on the QC Register.

      - Leishman stain: Appearance- Blue coloured solution. Clarity- Clear without any particles.

      - Microscopic Examination Blood staining is carried out and staining characteristics are observed under microscope by using oil immersion lens.

External Quality Control

- N/A.

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- Results:

          - Erythrocytes				Reddish

          - Cytoplasm of lymphocytes		Blue to Grey-blue

          - Nuclei				Purple-pink

          - Neutrophilic granulations		Pink-violet

          - Eosinophil granulations		Orange-red

          - Basophilic granulations		Dark Purple

          - Platelets				Purplish-purple

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

Sources of Error:

      - Fixation:

        - Inadequate fixation may result in a detached specimen, indistinct nuclear detail or poor granule staining.  This is especially evident in basophils and primary myeloid granules.

        - Dry freshly prepared smears for at least fifteen (15) minutes to prevent drying artifacts.  Use only pre-cleaned glass slides.  Slides other than pre-cleaned may result in improper fixation and the presence of haze over the slide due to acid-base disturbance. Pre-cleaned slides are necessary for quality stain appearance.

        - Use only anhydrous methanol for fixative.  Keep it tightly closed and away from moisture or chemical fumes.

        - If fixation problems occur or increased granularity visualization is necessary, increase the fixation period to one (1) minute.

      - Staining:

        - Excessively blue stain appearance may be due to improper smear preparation, prolonged staining time or increased stain alkalinity.	Excessively pink stain appearance may be due to insufficient or decreased staining time, prolonged washing time or increased stain acidity.

      - Rinse:

        - Use only distilled water or phosphate buffered saline for the final rinse. Tap water is never acceptable as a substitute for the rinse since it can disturb the acid base balance and may contain chlorine.

        - Vigorous or prolonged washing may dislodge cells from under-fixed smears, cause nuclear clumping or under-staining.

        - Insufficient rinsing may result in a blue smear or excessive precipitate formation.

      - Drying:

        - Air drying is the preferred method.  Forced rapid drying may alter the color intensities by shortening the exposure time to the wash water because the red spectrum of colors continues to develop as long as the cellular elements are wet.

- References

      - Practical Hematology, J. Dacie and M. Lewis, Tenth edition, 2006. Page No 19-38,55.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).