SOP – 013

PROCEDURE FOR

CBC PARAMETERS

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Type of Container & Additives	10

9.	Patient Preparation	10

10.	Equipment & Reagents	10

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	12

14.	Quality Control Procedure	14

15.	Interferences & Cross Reactions	14

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	21

19.	Procedure when Results are not in Range	21

20.	Critical Alert Values	22

21.	Laboratory Clinical Interpretation	22

22.	Potential Source of Variability	23

23.	References	23

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | CV | Coefficient of Variation |
| 3 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HCT | Hematocrit |
| 6 | HOD | Head of Department |
| 7 | HM | Haematology |
| 8 | INR | International Normalized Ratio |
| 9 | IQC | Internal Quality Control |
| 10 | LJ | Levey Jennings |
| 11 | MCH | Mean Corpuscular Haemoglobin |
| 12 | MCHC | Mean Corpuscular Hemoglobin Concentration |
| 13 | MCV | Mean Corpuscular Volume |
| 14 | MSP | Management System Procedure |
| 15 | N/A | Not Applicable/Available |
| 16 | Nm | Namometer |
| 17 | PLT | Platelet |
| 18 | QC | Quality Control |
| 19 | QM | Quality Manager |
| 20 | QNS | Quantity Not Sufficient |
| 21 | RBC | Red Blood Cells |
| 22 | RDW | Red cell Distribution Width |
| 23 | SOP | Standard Operating Procedure |
| 24 | TRF | Test Requisition Form |
| 25 | TDPL | TRUSTlab Diagnostics Private Limited |
| 26 | WBC | White Blood Cells |

- Objective

- The purpose of this SOP is to describe the procedure for the Quantitative enumeration of different formed elements of blood and Hemoglobin in whole blood by automated cell counter by the principle of Impedance Variation. This test includes Total WBC count, Total RBC count, Haemoglobin, Hematocrit, MCV, MCH, MCHC, Platelet count & Differential count. It is the enumeration of cellular elements of blood, evaluation of red cells indices and determination of cell morphology by means of stained smears. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Complete Blood Count (CBC) is a common blood test that assesses your overall health by analyzing the different cells in your blood. It measures red blood cells, white blood cells, and platelets, providing insights into potential health issues, infections, and other conditions.

- Test Principle

White blood cells, red blood cells and platelets are counted by impedance variation generated by the passage of cells through a calibrated micro aperture.

It is based on the principle that blood cells are less conductive to electric current than the cell diluting fluid.

The number of cells is quantitatively determined by passing a metered volume of a cell suspension (by diluting the sample in electrolytic diluents) between two electrodes placed on either side of the aperture. The current passes between the two electrodes through the diluents. When the cell passes through the aperture, electric resistance (or impedance) between the two electrodes increases proportionately with the cell volume which is registered as a pulse which is amplified and allows the cell to be counted.

The magnitude of each pulse is proportional to the size of the particle. The distribution of magnitude of pulses according to particle size of the cell population is printed as a Histogram.

Only electronic impulses of a size falling within a threshold limit are registered as cell particles while pulses from particles with a size beyond the fixed limit are not counted.

WBCs are counted after lysing the RBCs by a detergent (ionic or non-ionic) which does not lyse the WBCs themselves. After dilution (1: 50 for WBC count and Hb 1: 50,000 for RBC count Indices and platelet count) the diluted sample is directed to two sides of the instrument and particles on the red cells side that are larger than 36 fl are counted as RBCs.

Cells from 2- 20 fl are counted as platelets.

For WBC counting lysing reagent is added to the sample in the system to lyse RBCs and causes differential shrinking of WBCs, thereby allowing them to be separated into a partial differential (3-part differential) and white cells are shown in 3 groups Lymphocytes (35- 90 fl) mononuclear cells (90-160 fl) and granulocytes (160 -450 fl)

HCT is calculated based on RBC pulse height value. It is measured by the height of impulse generated by a passage of cells through a micro aperture, which is directly proportional to the volume of analysed cells.

RBCs are lysed and the freed haemoglobin combines with potassium cyanide to form cyanmethemoglobin, which is measured by Spectrophotometry at 450 nm.

MCV, MCH and MCHC are calculated from RBC count Hb and HCT.

MCV: calculated from RBC and HCT

MCH: Calculated from RBC and HGB

MCHC: Calculated from HGB and HCT

RDW: Calculated from deviation of RBC distribution width.

- Specific Performance Characteristics

Linearity:

| WBC | 0 – 300000 (cells/cu mm) |
| --- | --- |
| RBC | 0 – 8 (millions/cu mm) |
| Platelets | 0 – 25.00 (lakhs/cu mm) |
| HGB | 0-240 (gm%) |

Measurement Range:

| WBC | 0.0-999999 (cells/cu mm) |
| --- | --- |
| RBC | 0.00-9.99 (millions/cu mm) |
| Platelets | 0 – 29.99 (lakhs/cu mm) |
| HGB | 0-300(gm%) |
| MCV | 0.0-250 (fl) |

Sensitivity: Not applicable.

Clinical Reportable Range:

| WBC (cells/cu mm) | 500-2,00,000 |
| --- | --- |
| RBC (millions/cu mm) | 0.5-8 |
| Platelets (lakhs/cu mm) | 0.05-10.00 |
| HGB (gm%) | 2.0-25 |
| MCV (fl) | 40-120 |

- Type of Sample

3 ml Whole Blood

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

K2 EDTA Vacutainer

- Patient Preparation

No special preparation.

- Equipment & Reagents

Equipment

- Horiba Yumizen H500 – Automated Haematology Analyzer

Reagents

      - ABX Diluent

      - ABX Cleaner

      - White diff

      - ABX Minoclair

- Safety

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

- Calibration Procedure

The purpose of the calibration is to maintain system accuracy. Calibration program should be run if it is the first time the analyzer has been used, or certain major components of the analyzer have been changed or the quality control results indicate there may be a problem. Calibration can be done by 3 calibration programs: manual calibration, auto calibration using commercial calibrators and auto calibration using fresh blood samples.

Before calibration do the background check. Enter the auto calibration from Calibration from the main menu. Select the count mode from the sample mode pull down list.

Press [ENTER] to activate the edit boxes. Enter lot number and expiry date in the respective boxes. Enter the expected results and limits (range) into the respective boxes of the parameters to be included in the calibration. After finishing the settings, press [ENTER] deactivate the edit boxes. To run a calibrator, check the system status is ready and the count mode is ‘Whole’. Present vial of mixed calibrator to the sample probe and press the aspirate key. When the analysis is finished, the result will be displayed on the screen. To run in predilute mode select the count mode ‘Predilute’. Present diluent and a message box will pop up to press aspirate key to add diluent. Present a clean sample tube to the sample probe. Press the Aspirate key to dispense 0.7 mL of diluent into the tube. The dispensing volume is controlled by the analyzer. When the dispensing is finished, press ‘Enter’ to close the box. Add 20 µL of calibrator to the diluent and shake the tube to mix the sample. Present vial of mixed calibrator to the sample probe and press the aspirate key. When the analysis is finished, the result will be displayed on the screen.

If nonnumeric parameter values (***) are obtained, a message box will pop up warning ‘Invalid Result’. Click ‘Enter’ to clear the results. If all the parameter values obtained are numeric, a message box will pop up to confirm the obtained calibration result is valid. Click ‘Enter’ to save the results or Click ‘Cancel’ to abort the result.

Run the calibrator 3-5 times and the analyzer will automatically calculate the CVs and calibration factors. The CVs should meet the requirements of the following table.

- Table 1

| Parameter | Expected Range | CV % |
| --- | --- | --- |
| WBC | 7000 – 15000 (cells/cumm) | ≤ 2.5 |
| RBC | 3.50 – 6.00 (million/cumm) | ≤ 2.0 |
| HEMOGLOBIN | 100 – 180 (gm%) | ≤ 1.5 |
| MCV | 70.0 – 110.0 (Fl) | ≤ 0.5 |
| PLT | 1.50 – 5.00 (lakhs/cumm) | ≤ 5.0 |

The calculated calibration factor should be within 75% - 125%. Any calculated value that falls between 0% - 75% or 125%-9999% will be flagged with a ‘*’. Other values will not be displayed. Press [Main] to enter the ‘Count’ screen, a message box will pop up to confirm the new calibration factors. Click ‘Enter’ to save the new calibration factors. To calibrate with fresh blood press [MENU], select ‘Calibration’ and ‘Fresh Blood’. Check the system status is ready and select the mode as ‘Whole’ or ‘Predilute’ from the pull-down list. Prepare 3 – 5 normal fresh blood samples for the calibration. Press [1]…[5] to switch among ‘Sample 1’ to ‘Sample 5’ to select the sample for calibration. Press [Enter] to enter the editing state of the expected value of sample 1. Enter the expected result into the ‘Mean’ edit box. After finishing editing the calibration settings run the sample 1 to perform the fresh blood calibration. If nonnumeric parameter values (***) are obtained, a message box will pop up warning ‘Invalid Result’. Click ‘Enter’ to clear the results. If all the parameter values obtained are numeric, a message box will pop up to confirm the obtained calibration result is valid. Click ‘Enter’ to save the results or Click ‘Cancel’ to abort the result.

Run the ‘Sample 1’ 3-5 times and the analyzer will automatically calculate the CVs and calibration factors. The CVs should meet the requirements of table 1.

Enter the “Auto – Fresh Blood” screen of ‘Sample 2’…’Sample 5’ by pressing [2] …[5]. Follow the calibration steps of sample 1 to run at least another two fresh blood samples. After getting calibration factors for at least 3 fresh blood samples, press [6] to enter the ‘Calculate’ screen. The calibration factors for sample 1 to sample 5 are displayed on the screen. The ‘Calculate’ screen can display a maximum of 5 sets of calibration factors. The calculated calibration factor should be within 75% - 125%. Any calculated value 	that falls between 0% - 75% or 125%-9999% will be flagged with a ‘*’. Other values will not be displayed. Press [Main] to enter the ‘Count’ screen, a message box will pop up to confirm the new calibration factors. Click ‘Enter’ to save the new calibration factors.

- Test Procedure

Press ‘Sample Analysis”.

Enter the sample information by pressing ‘ID’ at the count screen. Enter the sample information by one of the two ways, ID or All Info. After entering all the desired information. Click the ‘Enter’ button to save the return to the ‘Count’ Screen. Press ‘Cancel’ button if you don’t want to save the entered information.

At the count screen be sure the system status area displays ‘Ready’.

Present the sample to the sample probe so that the tip is well into the tube and press the aspirate key. The system status area will display ‘Running’ and the analyzer start aspirating sample. After the beep sound the sample probe is out of the tube, remove the sample tube. The sample probe will retract into the analyzer and the analysis progress will be displayed on the screen.

When the analysis is finished, the result will be displayed on the screen and the sample ID will automatically increase by 1 and the sample probe will be repositioned. If auto print function is enabled, the analysis result will be automatically printed out.

The analyzer detects WBC/RBC clogging or bubbles during the analysis, the corresponding error messages will be displayed in the upper left corner of the screen.

To remove the error list, go to the maintenance screen. Zap and flush the apertures.

The analyzer automatically saves a maximum of 35,000 sample results. When the maximum numbers have been reached, the newest result will overwrite the oldest.

If the analysis result is followed by an “H” or “L”, the analysis result has exceeded the upper or lower limit of the reference range.

If the *** as opposed to the result, it means the result is either unreliable or out of the operating range. If the WBC result is less than 0.5 x 109/l, the analyzer will not perform the differential analysis, and all the related parameter values will be numeric (***).

Abnormal WBC histograms will be flagged by one of the markings R1, R2, R3, R4, Rm. R1 indicates abnormality on the left side of the lymphocyte hump and possible presence of platelet clumps, giant platelets, nucleated red cells, insolvable red cells, protein and lipoid debris in sample, or electrical noise. R2 indicates abnormality between the lymphocyte hump and the mid-sized cell area and possible presence of abnormal lymphocyte, plasma cell, atypical lymphocyte, original granulocytes in the sample and eosinophilia or basophilia. R3 indicates abnormality between the mid-sized cellaria and the granulocytes, abnormal sub population in the sample, or eosinophilia. R4indicates abnormality on the right side of the granulocytes hump and neutrophilia. Rm indicates at least two R flags.

Abnormal Platelet histograms will be flagged by one of the markings Pm, Ps and PL. Pm indicates blue demarcation between the platelet and red blood cell area and possible presence of large platelet, platelet coagulation, small red blood cell, cell debris or fibrin. Ps indicates excessive small platelets. PL indicates excessive large platelets.

When the blood count has been flagged by the analyzer – e.g. because of microcytosis, agglutination, cell fragments, platelet clumps etc., blood smear has to be prepared and examined. For smear preparation, examination and reporting refer TDPL/HM/SOP-14.

The analyzer will enter the screen saver if it has been idle at the count screen for 10 minutes. When it happens, the sample probe will retract into the analyzer, the LCD will turn dark, and the power indicator will be flickering. Press any key to resume the display and reposition the sample probe.

- Quality Control Procedure

Internal Quality Control

      - Run all three levels of an assayed QC sample (Three Level Quality Control samples - Low, Normal and High) and verify if the values are close to the Mean and within the reference range mentioned in the QC product insert.

      - If the QC is not within the reference range even for any one level of QC, perform another cleaning procedure and rerun that level of QC.

      - Three level controls are used on a daily basis.  The lab Mean is calculated after 20 runs and used as reference range still the new lot receives.

External Quality Control

- EQAS – AIIMS EQAS.

- Interferences & Cross Reactions

Abnormal BUN, glucose, or sodium levels could affect the MCV.

Abnormal WBCs could affect lymphocytes, monocytes, and granulocytes.

Abnormally small WBCs could affect white count, lymphocytes, monocytes, and granulocytes.

Clumped platelets could affect white count, lymphocytes, monocytes, granulocytes, RBC, MCV, RDW, platelet count, and MPV.

- Result Calculation

- The Red Blood Cells constants (Mean Cell Volume, Mean Cell Haemoglobin and Mean Cell Haemoglobin Concentration) are calculated from the RBC, HGB and HCT.

      - Mean Cell Volume (MCV): The MCV is calculated from the RBC and HCT, using the following equation:

- MCV fl= HCT %×10RBC (1012l)

      - Mean Cell Haemoglobin (MCH): The MCH is calculated from the RBC and HGB, using the following equation:

- MCH pg= HGB×10RBC (1012l)

      - Mean Cell Haemoglobin Concentration (MCHC): The MCHC is calculated from the HCT and HGB, using the following equation:

- MCHC g/l= HGB×100HCT

- Biological Reference Interval

| REFERENCE INTERVALS – HAEMOGLOBIN |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Haemoglobin | Female/ Male | 0 | Year | To | 2 | Day | gm% | 14-22 |
|  |  |  | 3 | Day | To | 6 | Day | gm% | 8-11 |
|  |  |  | 7 | Day | To | 7 | Day | gm% | 13.5-21.5 |
|  |  |  | 8 | Day | To | 14 | Day | gm% | 12.5-20.5 |
|  |  |  | 15 | Day | To | 1 | Month | gm% | 12.5-16.5 |
|  |  |  | 2 | Month | To | 2 | Month | gm% | 9.2-13 |
|  |  |  | 3 | Month | To | 6 | Month | gm% | 11.1-14.1 |
|  |  |  | 7 | Month | To | 1 | Year | gm% | 11.1-14.1 |
|  |  |  | 2 | Year | To | 6 | Year | gm% | 11-14.5 |
|  |  |  | 7 | Year | To | 12 | Year | gm% | 11.5-15.5 |
|  |  | Female | 13 | Year | To | 100 | Year | gm% | 12-15 |
|  |  | Male | 13 | Year | To | 100 | Year | gm% | 13-17 |

| REFERENCE INTERVALS – RBC COUNT |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  |  |  | Units | Range |
| HM014 | R.B.C. count | Female/ Male | 0 | Year |  |  | To | 2 | Day | million/mm3 | 5.0-7.0 |
|  |  |  | 3 | Day |  |  | To | 6 | Day | million/mm3 | 4.0-6.6 |
|  |  |  | 7 | Day |  |  | To | 7 | Day | million/mm3 | 3.9-6.3 |
|  |  |  | 8 | Day |  |  | To | 14 | Day | million/mm3 | 3.6-6.2 |
|  |  |  | 15 | Day |  |  | To | 1 | Month | million/mm3 | 3.0-5.4 |
|  |  |  | 2 | Month |  |  | To | 2 | Month | million/mm3 | 3.1-4.3 |
|  |  |  | 3 | Month |  |  | To | 6 | Month | million/mm3 | 4.1-5.3 |
|  |  |  | 7 |  | Month |  | To | 1 | Year | million/mm3 | 3.9-5.1 |
|  |  |  | 2 |  |  | Year | To | 12 | Year | million/mm3 | 4.0-5.2 |
|  |  | Female | 13 |  |  | Year | To | 100 | Year | million/mm3 | 3.8-4.8 |
|  |  | Male | 13 |  |  | Year | To | 100 | Year | million/mm3 | 4.5-5.5 |

| REFERENCE INTERVALS – HCT |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Hematocrit (HCT) | Female/ Male | 0 | Day | To | 2 | Day | % | 50-75 |
|  |  |  | 3 | Day | To | 6 | Day | % | 45-67 |
|  |  |  | 7 | Day | To | 7 | Day | % | 42-66 |
|  |  |  | 8 | Day | To | 14 | Day | % | 49-53 |
|  |  |  | 15 | Day | To | 1 | Month | % | 33-53 |
|  |  |  | 2 | Month | To | 2 | Month | % | 28-42 |
|  |  |  | 3 | Month | To | 6 | Month | % | 30-40 |
|  |  |  | 7 | Month | To | 1 | Year | % | 30-38 |
|  |  |  | 2 | Year | To | 6 | Year | % | 34-40 |
|  |  |  | 7 | Year | To | 12 | Year | % | 35-45 |
|  |  | Female | 13 | Year | To | 100 | Year | % | 36-46 |
|  |  | Male | 13 | Year | To | 100 | Year | % | 40-50 |

| REFERENCE INTERVALS – MCHC |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | MCHC | Female/ Male | 0 | Day | To | 2 | Day | gm% | 30-36 |
|  |  |  | 3 | Day | To | 6 | Day | gm% | 29-37 |
|  |  |  | 7 | Day | To | 14 | Day | gm% | 28-38 |
|  |  |  | 15 | Day | To | 1 | Month | gm% | 29-37 |
|  |  |  | 2 | Month | To | 2 | Month | gm% | 28.5-35.5 |
|  |  |  | 3 | Month | To | 6 | Month | gm% | 30-36 |
|  |  |  | 7 | Month | To | 1 | Year | gm% | 32-36 |
|  |  |  | 7 | Month | To | 1 | Year | gm% | 32-36 |
|  |  |  | 2 | Year | To | 6 | Year | gm% | 31-37 |
|  |  |  | 7 | Year | To | 12 | Year | gm% | 31-37 |
|  |  |  | 13 | Year | To | 100 | Year | gm% | 32.5-34.5 |

| REFERENCE INTERVALS – MCH |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | MCH | Female/ Male | 0 | Day | To | 14 | Day | pg | 31-37 |
|  |  |  | 15 | Day | To | 1 | Month | pg | 30-36 |
|  |  |  | 2 | Month | To | 2 | Month | pg | 27-33 |
|  |  |  | 3 | Month | To | 6 | Month | pg | 24-30 |
|  |  |  | 7 | Month | To | 1 | Year | pg | 25-29 |
|  |  |  | 2 | Year | To | 6 | Year | pg | 24-30 |
|  |  |  | 7 | Year | To | 12 | Year | pg | 25-33 |
|  |  |  | 13 | Year | To | 100 | Year | pg | 27 - 32 |

| REFERENCE INTERVALS – MCV |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | MCV | Female/ Male | 0 | Day | To | 2 | Day | fL | 100 - 120 |
|  |  |  | 3 | Day | To | 6 | Day | fL | 92 - 118 |
|  |  |  | 7 | Day | To | 7 | Day | fL | 88-126 |
|  |  |  | 8 | Day | To | 14 | Day | fL | 89-124 |
|  |  |  | 15 | Day | To | 1 | Month | fL | 92-116 |
|  |  |  | 2 | Month | To | 2 | Month | fL | 97-103 |
|  |  |  | 3 | Month | To | 6 | Month | fL | 68-84 |
|  |  |  | 7 | Month | To | 1 | Year | fL | 72-84 |
|  |  |  | 2 | Year | To | 6 | Year | fL | 77-87 |
|  |  |  | 7 | Year | To | 12 | Year | fL | 77-95 |
|  |  |  | 13 | Year | To | 100 | Year | fL | 81-101 |

| REFERENCE INTERVALS – RDV-CV |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM014 | RDW-CV | Female/ Male | 0 | Year | To | 100 | Year | % | 11.6 - 14.0 |

| REFERENCE INTERVALS – TOTAL WBC |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Total WBC Count | Female/ Male | 0 | Day | To | 2 | Day | cells/mm3 | 10000 - 26000 |
|  |  |  | 3 | Day | To | 6 | Day | cells/mm3 | 7000 - 23000 |
|  |  |  | 7 | Day | To | 7 | Day | cells/mm3 | 6000-22000 |
|  |  |  | 8 | Day | To | 14 | Day | cells/mm3 | 6000-22000 |
|  |  |  | 15 | Day | To | 1 | Month | cells/mm3 | 5000-19000 |
|  |  |  | 2 | Month | To | 2 | Month | cells/mm3 | 5000-15000 |
|  |  |  | 3 | Month | To | 6 | Year | cells/mm3 | 6000-18000 |
|  |  |  | 7 | Month | To | 1 | Year | cells/mm3 | 6000-16000 |
|  |  |  | 1 | Year | To | 12 | Year | cells/mm3 | 6000 - 13000 |
|  |  |  | 2 | Year | To | 6 | Year | cells/mm3 | 5000-15000 |
|  |  |  | 7 | Year | To | 12 | Year | cells/mm3 | 5000-13000 |
|  |  |  | 13 | Year | To | 100 | Year | cells/mm3 | 4000 - 10000 |

| REFERENCE INTERVALS – GRANULOCYTES (NEUTROPHILS) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Granulocytes (Neutrophils) | Female/ Male | 0 | Day | To | 2 | Day | cells/mm3 | 4000-14000 |
|  |  |  | 3 | Day | To | 6 | Day | cells/mm3 | 3000-5000 |
|  |  |  | 7 | Day | To | 7 | Day | cells/mm3 | 3000-7000 |
|  |  |  | 8 | Day | To | 14 | Day | cells/mm3 | 3000-7000 |
|  |  |  | 15 | Day | To | 1 | Month | cells/mm3 | 3000-9000 |
|  |  |  | 2 | Month | To | 2 | Month | cells/mm3 | 1000-5000 |
|  |  |  | 3 | Month | To | 6 | Month | cells/mm3 | 1000-6000 |
|  |  |  | 7 | Month | To | 1 | Year | cells/mm3 | 1000-7000 |
|  |  |  | 2 | Year | To | 6 | Year | cells/mm3 | 1500-8000 |
|  |  |  | 7 | Year | To | 12 | Year | cells/mm3 | 2000-8000 |
|  |  |  | 13 | Year | To | 100 | Year | % | 40-80 |

| REFERENCE INTERVALS - MONOCYTES |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  |  | Units | Range |
| HM012 | Monocytes | Female/ Male | 0 | Day | To |  | 2 | Day | cells/mm3 | 500-2000 |
|  |  |  | 3 | Day | To |  | 6 | Day | cells/mm3 | 500-1000 |
|  |  |  | 7 | Day | To |  | 7 | Day | cells/mm3 | 100-1700 |
|  |  |  | 8 | Day | To | 14 |  | Day | cells/mm3 | 100-1700 |
|  |  |  | 15 | Day | To |  | 1 | Month | cells/mm3 | 300-1000 |
|  |  |  | 2 | Month | To |  | 2 | Month | cells/mm3 | 400-1200 |
|  |  |  | 3 | Month | To |  | 6 | Month | cells/mm3 | 200-1200 |
|  |  |  | 7 | Month | To |  | 1 | Year | cells/mm3 | 200-1000 |
|  |  |  | 13 | Year | To |  | 1 | Year | % | 02-10 |

| REFERENCE INTERVALS - EOSINOPHILS |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Eosinophil Count | Female/ Male | 0 | Day | To | 2 | Day | cells/mm3 | 100-1000 |
|  |  |  | 3 | Day | To | 6 | Day | cells/mm3 | 100-2000 |
|  |  |  | 7 | Day | To | 7 | Day | cells/mm3 | 100-800 |
|  |  |  | 8 | Day | To | 14 | Day | cells/mm3 | 100-900 |
|  |  |  | 15 | Day | To | 1 | Month | cells/mm3 | 200-1000 |
|  |  |  | 2 | Month | To | 12 | Year | cells/mm3 | 100-1000 |
|  |  |  | 13 | Year | To | 100 | Year | % | 01-06 |

| REFERENCE INTERVALS - BASOPHILS |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM012 | Basophil Count | Female/ Male | 0 | Day | To | 1 | Day | cells/mm3 | 00-700 |
|  |  |  | 2 | Day | To | 6 | Month | cells/mm3 | 000-400 |
|  |  |  | 7 | Month | To | 2 | Year | cells/mm3 | 000-1200 |
|  |  |  | 2 | Year | To | 12 | Year | cells/mm3 | 000-300 |
|  |  |  | 13 | Year | To | 100 | Year | % | <1-2 |

| REFERENCE INTERVALS – ABSOLUTE ESINOPHIL COUNT (AEC) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Code | Parameter | Sex | Age |  |  |  |  | Units | Range |
| HM056 | Absolute Eosinophil Count | Female/ Male | 0 | Day | To | 2 | Day | cells/mm3 | 100-1000 |
|  |  |  | 3 | Day | To | 6 | Day | cells/mm3 | 100-2000 |
|  |  |  | 7 | Day | To | 7 | Day | cells/mm3 | 100-800 |
|  |  |  | 8 | Day | To | 14 | Day | cells/mm3 | 100-900 |
|  |  |  | 15 | Day | To | 1 | Month | cells/mm3 | 200-1000 |
|  |  |  | 2 | Month | To | 12 | Year | cells/mm3 | 100-1000 |
|  |  |  | 13 | Year | To | 100 | Year | cells/mm3 | 20-500 |

- Reportable Interval

- Clinical Decision Values

| Parameter | Decision Levels |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 |  |
| Hemoglobin | 4.5 | 10.5 | 17 | 23 |  |  |
| Mean Corpuscular Volume | 80 | 100 |  |  |  |  |
| Platelet count | 0.1 | 0.5 | 1 | 6 | 10 |  |
| White blood cell count | 500 | 3000 | 12000 | 30000 |  |  |

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

| Test |  | Units | Lower Limit | Upper Limit | Comments |
| --- | --- | --- | --- | --- | --- |
| Hematocrit | Adult | % | 20 | 60 | First report only |
|  | Newborn | % | 33 | 71 |  |
| Haemoglobin | Adult | gm% | 7 | 20 | First report only |
|  | Newborn | gm% | 10 | 22 |  |
| WBC | Adult | cells/mm3 | 2000 | 30000 | First report only |
|  | Children | cells/mm3 | 2000 | 43000 |  |
| Platelets |  | cells/mm3 | 0.4 | 1000 |  |
| Blasts |  | Any seen (first report only) |  |  |  |

- Laboratory Clinical Interpretation

Using CBC to evaluate anemia

Differential diagnosis of microcytic anemia

White Blood Cell Information

      - Leucocytosis = >11,000 WBC cells/mm3

      - Myeloproliferative disorders

      - Infection (especially bacterial)

      - Absent in 15-30% of cases

      - Tissue necrosis

      - Leukopenia = < 4000 cells/cumm

      - Viral infection usually leukopenia or normal WBC

Drugs

      - Examples: Clozaril, chemotherapeutic agents

Autoimmune disease

Aplastic anemia

Platelet count

Thrombocytopenia < 1.5 lakhs/cumm

Thrombocytosis > 4 lakhs/cumm

Mean platelet volume (MPV)

Indicates size (reference range 7.4-10.4 fl)

Thrombocytopenia and:

      - Large platelets → peripheral destruction

      - Small platelets → production defect

- Potential Source of Variability

      - WBC: False High Leukocyte count due to the presence of lyses resistant and nucleated erythrocytes.

      - RBC: False low erythrocyte counts due to the presence of cold agglutinins and fragmented erythrocytes.

      - Haemoglobin: False high haemoglobin measurement due to Lipemia and Abnormal protein.

      - HCT: False low haematocrit measurement due to cold agglutinins and fragmented erythrocytes. False high haematocrit measurement due to leucocytosis.

      - Platelet: False low platelet counts due to platelet aggregation and giant platelets. False high platelet counts due to microcytic and fragmented erythrocytes.

- References

      - User Manual of Horiba Yumizen H500.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).