SOP – 014

PROCEDURE FOR

PERIPHERAL BLOOD SMEAR

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	17

16.	Result Calculation	17

17.	Biological Reference Interval	17

18.	Reportable Interval	17

19.	Procedure when Results are not in Range	18

20.	Critical Alert Values	18

21.	Laboratory Clinical Interpretation	18

22.	Potential Source of Variability	18

23.	References	18

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | HOD | Head of Department |
| 5 | HM | Haematology |
| 6 | IQC | Internal Quality Control |
| 7 | MSP | Management System Procedure |
| 8 | N/A | Not Applicable/Available |
| 9 | PLT | Platelet |
| 10 | QC | Quality Control |
| 11 | QM | Quality Manager |
| 12 | QNS | Quantity Not Sufficient |
| 13 | RBC | Red Blood Cell |
| 14 | rpm | Revolutions per minute |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |
| 18 | WBC | White Blood Cell |

- Objective

- The purpose of this SOP is to describe the procedure to perform a stained blood smear by the Leishman staining method and observing it under a light microscope. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- A peripheral blood smear is a diagnostic tool where a small blood sample is spread on a slide, stained, and examined under a microscope to evaluate the characteristics of blood cells. This allows healthcare professionals to assess the size, shape, and appearance of red blood cells, white blood cells, and platelets, which can reveal clues about various health conditions.

- Test Principle

The blood smear examination provides a visual confirmation of automated instrument results, and in some instances, may provide morphologic clues for a specific case finding. A properly made peripheral smear is examined to identify RBC, WBC, and PLT morphology, estimate hematologic counts, perform WBC differentials and confirm the accuracy of automated differentials.

Method: Staining and microscopic examination

- Specific Performance Characteristics

- N/A.

- Type of Sample

3 ml Whole Blood

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

K2 EDTA Vacutainer

- Patient Preparation

No special preparation.

- Equipment & Reagents

Equipment

      - Microscope with Oil Immersion objective

      - Slides

      - Marker

Reagents

      - Leishman stain

      - Phosphate Buffer

- Safety

Blood should always be tested using Universal Precautions.

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

Refer to each reagents MSDS sheet for specific information regarding handling and hazards.

Reagents used for CBC and individual parameters are not for invitro testing.

- Calibration Procedure

- N/A.

- Test Procedure

General Criteria for Review:

      - The following criteria or flags from the automated haematology instrumentation may trigger peripheral blood smear (PBS) review critical values, suspect messages, instrument flags.

      - Review histogram for any interpretative messages and/or flags.

      - Refer to the section below to evaluate results for the need for a peripheral smear scan or differential.

      - If the analysis result is followed by an “H” or “L”, the analysis result has exceeded the upper or lower limit of the reference range.

      - If the *** as opposed to the result, it means the result is either unreliable or out of the operating range. If the WBC result is less than 0.5 x 109/l, the analyser will not perform the differential analysis and all the related parameter values will be numeric (***).

      - Abnormal WBC histograms will be flagged by one of the markings R1, R2, R3, R4, Rm. R1 indicates abnormality on the left side of the lymphocyte hump and possible presence of platelet clumps, giant platelets, nucleated red cell, insolvable red cell, protein and lipoid debris in sample, or electrical noise. R2 indicates abnormality between the lymphocyte hump and the midsized cell area and possible presence of abnormal lymphocyte, plasma cell, atypical lymphocyte, original granulocytes in the sample and eosinophilia or basophilia. R3 indicates abnormality between the mid sized cell area and the granulocytes, abnormal sub population in the sample, or eosinophilia. R4 indicates abnormality on the right side of the granulocytes hump and neutrophilia. Rm indicates at least two R flags.

      - Abnormal Platelet histograms will be flagged by one of the markings Pm, Ps and PL. Pm indicates blue demarcation between the platelet and red blood cell area and possible presence of large platelet, platelet coagulation, small red blood cell, cell debris or fibrin. Ps indicates excessive small platelets. PL indicates excessive large platelets.

      - When the blood count has been flagged by the analyser – e.g. because of microcytosis, agglutination, cell fragments, platelet clumps etc., blood smear has to be prepared and examined.

Slide Examination

      - Low Power (10X)

- Examine for adequate staining and evaluate the slide for an even distribution of leukocytes, erythrocytes, and platelets. Find a suitable area representative of the various cell types that will be counted. Scan for immature cells, infrequent cell types, NRBC’s, platelet clumps, and Rouleaux, while assessing cellular distribution

- Perform a WBC Estimate:

          - Count the number of WBCs on low power where the red cells form a uniform monolayer. Divide the number by 4 to obtain the number of WBC/lPF. Then multiply the number by 1,000 to obtain the estimated total WBC count.

          - If there is a marked discrepancy between the estimate and the automated WBC count, recheck the smear for abnormal distribution of cells. If this does not explain the discrepancy, repeat the WBC, make a new smear, and check for possible slide/specimen/report mix up.

      - Oil Immersion (100X):

- Perform a Platelet Estimate:

          - Scan the smear around the edges on high power to make sure there are no platelet clumps, giant platelets, platelet satellitism, microcytic erythrocytes, and/or small red cell fragments.

          - View the slide under oil immersion where the red cells form a uniform monolayer. Count the number of platelets in 10 fields and then calculate the average number per field. Take the average number of platelets/OIF and multiply by 20,000 to obtain the estimated total platelet count. If the manual platelet count agrees with ± 25% of the automated count, the automated count will be reported.

      - RBC Morphology Examination:

- Review a minimum of 10 oil immersion fields (OIF) when determining how many abnormal cells seen per field.

- Quality Control Procedure

Internal Quality Control

      - Prepare one slide daily using a normal sample. Stain the slide as indicated in previous sections.  Review the slide for Color, Precipitation and Contamination.

      - If the color does not meet the specifications identified in the STP as indicated above or precipitation and/or contamination are present, the quality is determined to be unsatisfactory.   Indicate unsatisfactory on the Internal Quality Control (IQC) record and replace the stain as indicated in the following section.  Document the problem and corrective action in the appropriate section of the QC form.  If the stain is determined to be satisfactory, indicate satisfactory on the Leishman Stain QC Register (TDPL/HM/REG-002).

      - Technicians are evaluated annually and quarterly on proficiency testing specimens.

      - Differential Quality Control:

- The following is a list of measures undertaken by our laboratory to ensure the quality of differential results reported on patients:

          - All the flagged reports are reviewed by peripheral smear examination.

          - Refer to Blood Film review criteria.

          - Differentials must be repeated if each cell does not meet the limits set out in the table below. If the repeat is still not within the established limits, a second technician should repeat the differential.

          - Whenever the tech1 has concerns with a differential the rest of the CBC can be released with an annotation “Differential to follow”. The smear should be evaluated as soon as possible by the pathologist.

          - Leukocyte abnormalities seen during a smear review require a manual differential completed regardless of the protocol for when to perform a manual differential.

          - Perform the manual differential and compare it to the automated differential using the 95% confidence limits table. Each cell should compare within the range set.

          - Differentials between technicians should also fall within the established limits (95% confidence limits).

          - If the manual differential performed by two technicians agrees within the established limits, but is not in agreement with the automated differential, then the manual differential should be reported out instead of the automated differential.

          - It is good laboratory practice to circulate unknown QA slides quarterly. The results will be compared with peers and should be within the 95% confidence limits as set by the following table. Any problem areas will be covered between the technician and the supervisor.

      - Procedure to use 95% confidence limits table:

          - Look up each cell number you want to compare to in column “a”. E.g. 25% neutrophils. If a 100-cell differential was performed go to the column “n=100” to determine the acceptable range. If a 200-cell differential was done refer to column “n=200”. E.g. you counted 32 neutrophils in a 100-cell differential.

          - Determine acceptability of the differential by checking if the number you counted for a certain cell falls within the stated range. E.g. Stated range is 16-35%; you are within this range.

          - Repeat for each cell type.

          - Determine acceptability of each cell line by comparing automated to manual differential. For the neutrophils/granulocytes, segmented neutrophils, band neutrophils and other neutrophil precursors must be added together and for lymphocytes, reactive lymphocytes and lymphocytes must be added together.E.g. 77 segs and 15 bands = 92%

          - You must be within this range, or the differential must be repeated.

          - Example: The automated or technician differential indicated the following differential and the acceptable range for each number was looked up in “n=100” column:

            - Neutrophil: 70 acceptable range 60-79

            - Lymphocyte: 15 acceptable range 8-24

            - Monocyte: 7 acceptable range 1-12

            - Eosinophil: 3 acceptable range 0-9

            - A 100-cell differential is completed by another technologist and based on the acceptable range the results are as follows:

            - Neu: 52 not within limits

            - Lymph: 27 not within limits

            - Mono: 12

            - Eos: 7 within limits

            - Basso: 2 within limits

            - This manual differential would have to be repeated.

      - Rumke Binomial Distribution Table

- 95% Confidence limits for various percentages of blood cells of a given type as determined by differential counts

| a | n = 100 | n = 200 | n = 500 | n = 1000 |
| --- | --- | --- | --- | --- |
| 0 | 0 – 2 | 0 - 2 | 0 – 1 | 0 - 1 |
| 1 | 0 – 6 | 0 - 4 | 0 – 3 | 0 - 2 |
| 2 | 0 – 8 | 0 - 6 | 0 – 4 | 1 - 4 |
| 3 | 0 – 9 | 1 - 7 | 1 – 5 | 2 - 5 |
| 4 | 1 – 10 | 1 - 8 | 2 – 7 | 2 - 6 |
| 5 | 1 – 12 | 2 - 10 | 3 – 8 | 3 - 7 |
| 6 | 2 – 13 | 3 - 11 | 4 – 9 | 4 - 8 |
| 7 | 2 – 14 | 3 - 12 | 4 – 10 | 5 - 9 |
| 8 | 3 – 16 | 4 - 13 | 5 – 11 | 6 - 10 |
| 9 | 4 – 17 | 5 - 14 | 6 – 12 | 7 - 11 |
| 10 | 4 – 18 | 6 - 16 | 7 – 13 | 8 - 13 |
| 15 | 8 – 24 | 10 - 21 | 11 – 19 | 12 - 18 |
| 20 | 12 – 30 | 14 - 27 | 16 – 24 | 17 - 23 |
| 25 | 16 – 35 | 19 - 32 | 21 – 30 | 22 - 28 |
| 30 | 21 – 40 | 23 - 37 | 26 – 35 | 27 - 33 |
| 35 | 25 – 46 | 28 - 43 | 30 – 40 | 32 - 39 |
| 40 | 30 – 51 | 33 - 48 | 35 – 45 | 36 - 44 |
| 45 | 35 – 56 | 37 - 53 | 40 – 50 | 41 - 49 |
| 50 | 39 – 61 | 42 - 58 | 45 – 55 | 46 - 54 |
| 60 | 49 – 70 | 52 - 67 | 55 – 65 | 56 - 64 |
| 70 | 60 – 79 | 63 - 77 | 65 – 74 | 67 - 73 |
| 75 | 65 – 84 | 68 - 81 | 70 – 79 | 72 - 78 |
| 80 | 70 – 88 | 73 - 86 | 76 – 84 | 77 - 83 |
| 90 | 82 – 96 | 84 - 94 | 87 – 93 | 87 - 92 |
| 100 | 96 – 100 | 98 - 100 | 99 – 100 | 99 |

      - Establishing Conversion Factor for WBC Estimation:

          - Perform electronic WBC count on 30 consecutive fresh patient blood samples. Alternatively, if unable to perform 30 consecutive samples, it is acceptable to perform 10 samples per day on 3 consecutive days.

          - Prepare and stain one peripheral blood film for each sample.

          - For each film, under hpf microscopy, find an area where 50% of the red cells are overlapping doubles or triplets. Then count WBC in 10 consecutive fields.

          - Divide by 10 the total number of WBC found to obtain the average number per hpf.

          - Divide the electronic WBC count by the average number of WBC per hpf.

          - Add the numbers obtained in step 5 and divide by 30 (number of observations in this analysis) hpf.

          - Round the number calculated in step 6 to the nearest whole number to obtain an estimation factor.

          - Use the WBC Estimation Conversion Factor Log Sheet.

      - Procedure For Platelet Estimates:

          - Platelet estimation is to be used for cases in which instrument flags identify invalid platelet data, platelet clumps and for verifying platelet counts below 100x109/l. Scan the slide for clumps under hpf. If platelet clumps are present, they cause a false decrease in platelet count. Edit out the platelet count by replacing it with the message ―Platelets clumped, results invalid. Do not report the MPV. NOTE: For this purpose, platelet clumping is defined as at least 5 clumps, with at least 5 platelets per clump. Rare or occasional smaller clumps should be ignored.

          - Switch to 100x objective. Total the number of platelets in at least 10 fields. Determine the average number per field. Multiply by the conversion factor established for your microscope. The estimate should agree with the analyser count +/- 25%

          - Example: Average of 6.3 cells per field x conversion factor of 11 = estimate of 69 x 10^9/l. Analyzer count of 88. Multiply by 0.6 or 1.4 for +/- 25% range of 53 to 123. Platelet estimate agrees with analyser.

          - If your estimate disagrees with the instrument platelet count by more than 25% in the absence of clumps: Have another technologist do an estimate, when possible, to confirm the discrepancy.

          - If the patient is anaemic or polycythaemia (RBC below 4.0 or over 6.0), the estimate may be inaccurate. A second blood film should be prepared, (see Indirect Platelet Count Procedure) and performed by two technologists, this may require referral.

          - Examine the platelet histograms and the blood smear for the following sources of error:

            - RBC fragments: may be small enough to be counted as platelets.

            - Mega thrombocytes: if many of the platelets on the film appear abnormally large, i.e. approximately the size of RBC, the analyser may not be including them in the count. If your estimate varies from the analyser by more than 40%, add the message ―Platelet estimate appears higher on film; mega thrombocytes may be excluded from count. ‖ Do not comment anything if occasional large platelets are present, and your estimate agrees within 25%.

            - Leukocyte cytoplasmic fragments, if present in large numbers, may falsely elevate the instrument platelet count. They appear as particles similar in size to platelets, but with a homogeneous, rather than granular, structure.

          - If you are reporting the platelet estimate instead of the analyser count, append an appropriate comment:

- Platelet estimate ≤ 0.2 lakhs/cumm

- Platelet estimate > 0.2 lakhs/cumm

- Platelet estimate > 0.5 lakhs/cumm

- Platelet estimate > 1 lakhs/cumm

- Note: DO NOT USE this method if RBC is normal.

          - Send to Senior Technologist/Pathologist for review on first occurrence for each patient for whom the instrument count is determined to be invalid due to RBC fragments, leukocyte fragments, or mega thrombocytes.

          - Procedure:

            - Perform electronic platelet count on 30 consecutive fresh patient blood samples. Alternatively, if unable to perform 30 consecutive samples it is acceptable to perform 10 samples per day on 3 consecutive days.

            - Prepare and stain one peripheral blood film for each sample.

            - For each film, under oil immersion (100x) microscopy, find an area where 50% of the red cells are overlapping in doubles or triplets. Then count platelets in 10 consecutive fields.

            - Divide by 10 the total number of platelets found to obtain the average number per oil immersion field.

            - Divide the electronic platelet count by the average number of platelets per oil immersion field.

            - Add the numbers obtained in step 5 and divide by 30 (number of observations in this analysis) to obtain the average ratio of the platelet count to platelets per oil immersion field.

            - Round the number calculated in step 6 to the nearest whole number to obtain an estimation factor, the number of peripheral blood platelets represented by one platelet in an oil immersion field.

            - Use the platelet estimation conversion factor log sheet.

      - Criteria for Blood Smear Review by Pathologist: Significant morphologic abnormalities:

        - RBC:

          - Anisocytosis> 3+, Poikilocytosis> 3+, Hypochromia> 3+, Polychromasia> 3+

          - Basophilic stippling > 3+, Elliptocytes> 3+, Stomatocytes> 3+

          - Microcytes> 2+, Macrocytes> 2+, Target cells > 2+, Rouleaux> 2+

          - Tear drop cells > 1+, Schistocytes> 1+, Spherocytes> 1+, Acanthocytes> 1+

          - Sickle cells any, Howell jolly bodies any, Pappenheimer bodies any, agglutination any

        - WBC:

          - Dohle bodies > 3+, Hyposegmented neutrophils > 2+, Hyper segmented neutrophils > 1+

          - Hypo granular granulocytes any, Auer rods any

        - PLT:

          - Giant platelets > 2+, platelet satellitosis> 1+

External Quality Control

- EQAS – AIIMS EQAS.

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- Presence of blasts cells, any abnormal cells, low platelets.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

| Test |  | Units | Lower Limit | Upper Limit | Comments |
| --- | --- | --- | --- | --- | --- |
| Adult | % | 20 | 60 | First report only |  |
| Newborn | % | 33 | 71 |  |  |
| Adult | gm% | 7 | 20 | First report only |  |
| Newborn | gm% | 10 | 22 |  |  |
| Adult | Cells/cumm | 2000 | 30000 | First report only |  |
| Children | Cells/cumm | 2000 | 43000 |  |  |
| Platelets |  | Lakhs/cumm | 0.4 | 10 |  |
| Blasts |  | Any seen (first report only) |  |  |  |
| Drepanocytes (Sickle Cells) |  | Presence of sickle cells or aplastic crisis |  |  |  |

- Laboratory Clinical Interpretation

      - Typing of anemia is done by the peripheral blood smear examination.

      - Malaria diagnosis and identification of species

      - Other hemiparasites identification

      - Morphological typing of leukemia

- Potential Source of Variability

      - The estimated WBC count varies if done from the peripheral smear examination.

      - Parasitaemia Index varies.

      - Reticulocyte count varies.

- References

      - Practical Hematology, J. Dacie and M. Lewis, Tenth edition, 2006. Page No 19-38, 62.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).