SOP – 016

PROCEDURE FOR

Hb ELECTROPHORESIS

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	9

15.	Interferences & Cross Reactions	10

16.	Result Calculation	10

17.	Biological Reference Interval	10

18.	Reportable Interval	10

19.	Procedure when Results are not in Range	10

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	11

23.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | HOD | Head of Department |
| 5 | Hb | Haemoglobin |
| 6 | HM | Haematology |
| 7 | HPLC | High-Performance Liquid Chromatography |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | PLT | Platelet |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | RBC | Red Blood Cell |
| 16 | rpm | Revolutions per minute |
| 17 | SOP | Standard Operating Procedure |
| 18 | TRF | Test Requisition Form |
|  | TDPL | TRUSTlab Diagnostics Private Limited |
|  | WBC | White Blood Cell |

- Objective

- The purpose of this SOP is to describe the procedure to perform Hb Electrophoresis (Haemoglobin Variant Analysis) test by the HPLC method. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Hemoglobin electrophoresis is a laboratory test that separates different types of hemoglobin in a blood sample using an electric field. It also checks for hemoglobinopathy - disorders involving abnormal types of hemoglobin. The Hb types are:

Hemoglobin Types A11 and A2 (HgbA1 and A2): Hemoglobin A11 accounts for most of the hemoglobin in your normal red blood cells. Hemoglobin A2 accounts for about 2% to 3 % of your total hemoglobin.

Hemoglobin Type F (HgbF): Normally, this hemoglobin type accounts for the largest part of red blood cells in fetuses, babies and children up to age 3. A high type F hemoglobin level in adults is considered an abnormal hemoglobin level.

Hemoglobin Type S (HgbS): This hemoglobin type is a symptom of sickle cell anemia. Sickle cell anemia is an example of sickle cell disease. Sickle cell anemia is a serious illness that most commonly affects people who are Black.

- Hemoglobin C (HgbC): This hemoglobin type is linked to hemolytic anemia that develops when your red blood cells are destroyed more easily than normal red blood cells or have a shorter life span than normal red blood cells. Like sickle cell anemia, hemoglobin C most commonly affects people who are Black

- Test Principle

The D-10 Dual Program is based on chromatographic separation of the analytes by ion exchange high-performance liquid chromatography (HPLC). The samples are automatically diluted on the D-10 and injected into the analytical cartridge.

The D-10 delivers a programmed buffer gradient of increasing ionic strength to the cartridge, where the hemoglobins are separated based on their ionic interactions with the cartridge material. The separated hemoglobins then pass through the flow cell of the filter photometer, where changes in the absorbance at 415 nm are measured.

The D-10 software performs reduction of raw data collected from each analysis. Two-level calibration is used for quantitation of the HbA2/F/A1c values. A sample report and a chromatogram are generated for each sample. The A2, F, and A1c peaks are shaded.

- Specific Performance Characteristics

- N/A.

- Type of Sample

3 ml Whole Blood

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

K3 EDTA – Lavender Top Vacutainer

- Patient Preparation

No special preparation.

- Equipment & Reagents

Equipment

      - BIO-RAD D-10

      - Rack Holder

Reagents

      - Elution Buffer 1: 2 bottles containing 2000ml of a Bis-Tris/Phosphate buffer (pH 6.0)

      - Elution Buffer 2: 1 bottle containing 1000 ml of a Bis-Tris/Phosphate buffer (pH 6.7)

      - Diluent

- Safety

Blood should always be tested using Universal Precautions.

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

Refer to each reagents MSDS sheet for specific information regarding handling and hazards.

Reagents used for CBC and individual parameters are not for invitro testing.

- Calibration Procedure

Calibrator Required

      - HbA2/F/A1c Calibrator

      - Use deionized water as zero calibrator

Procedure

      - Place the following barcoded vial adapters in positions 1 through 4 of the sample rack: Calibrator 1, Calibrator 2, A1C Low Control, A1C High Control. (Can include patient samples).

      - Make sure barcodes are facing the back of the rack.

      - Pipette 1 ml of each sample type into the designated vial. Calibrators do not require pre-dilution. Controls require a 1:300 dilution with wash/diluent.

      - Insert rack into system.

      - If the calibrator and/or control barcodes are not available or if they don’t read correctly, manually type the following info into the worklist (Run/Edit screen): New barcodes: Calibrator 1, Calibrator 2, A1C Low Control, A1C High Control

Calibration Frequency

- Re-calibration must be done every 30 days or when the following occurs:

      - Change in reagent lot or significant shift in control values.

      - Major preventive maintenance was performed on the analyzer or a critical component was replaced.

- Test Procedure

Run Low and High controls in positions 1 and 2 of rack

Fill remaining rack positions with patient samples: -Pre-diluted patient samples (1:300) are placed in non-barcoded tube adapters

Insert rack into system

Barcodes are read

Manually enter the sample ID for any non-barcoded patient sample and/or mis-read barcodes in the Run/Edit screen.

Select the appropriate program (HbA2/F).

Select “Start” to begin run.

The D-10 system automatically analyzes the samples, separating hemoglobins based on their properties and detecting them using a filter photometer, according to Bio-Rad

D-10: Select “Eject” from the Run screen to remove rack when done.

When the run is complete, the system will remain in Standby for 30 minutes during which time more samples can be run. After 30 minutes of idleness, the system goes into the Sleep State.

- Quality Control Procedure

Internal Quality Control

- Run quality control samples as per the D-10 system's recommendations to ensure accurate results

External Quality Control

- ILC – Metropolis, Hyderabad.

- Interferences & Cross Reactions

| - Substance | - Limits |
| --- | --- |
| - Icterus | - No significant interference up to 42 mg/dl for conjugated bilirubin and 60 mg/dl for unconjugated bilirubin. |
| - Hemolysis | - No significant interference up to 250 mg/dl. |
| - Drugs | - No interference was found at therapeutic concentrations using common drug panels |

- Result Calculation

- TRUSTlab follows the estimation of CV %.

- The D-10 system provides results, including HbA1c, HbF, and HbA2 values, with clearly labeled peaks and common hemoglobin variants.

- Biological Reference Interval

- N/A.

- Reportable Interval

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

      - Blood transfusions may obscure abnormal haemoglobins. If there is a history of blood transfusion, repeat testing is recommended 3 months after the last transfusion.

      - Iron deficiency reduces HbA2 levels. HbA2 values between 1.5 % to 2 % constitute a grey zone which may be seen in both iron deficiency and in some thalassemias (e.g. Alpha thalassemia). Correction of iron deficiency is advised before re-testing to ensure accurate results.

      - In patients with HbA2 values between 3.6 % to 4 % which also constitute a grey zone, molecular studies are recommended to rule out thalassemias.

      - In neonates and infants repeat testing is advised after attaining one year of age for accurate results.

      - Mild to moderate increase in fetal hemoglobin can be seen in acquired conditions such as pregnancy, megaloblastic anemia, thyrotoxicosis, chronic kidney disease and medications.

      - This test is only a screening test for detection of Beta thalassemia and hemoglobinopathies. Molecular/genetic studies are recommended for confirmation of these or to rule out alpha thalassemia and silent carriers.

      - The results should be considered in conjunction with family history and laboratory findings, including RBC parameters, iron studies, ferritin and peripheral smear examination.

- Potential Source of Variability

      - In very rare cases, gammopathy may cause unreliable results, especially type IgM (Waldenstrom’s macroglobulinemia).

- References

      - Practical Hematology, J. Dacie and M. Lewis, Tenth edition, 2006. Page No 19-38, 62.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).