SOP – 003

PROCEDURE FOR

MICROTOMY

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration	10

13.	Procedure	10

14.	Quality Control Procedure	14

15.	Interferences & Cross Reactions	15

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	16

19.	Troubleshooting	16

20.	Critical Alert Values	18

21.	Laboratory Clinical Interpretation	18

22.	Potential Source of Variability	18

23.	References	19

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | °C | Degree Centigrade |
| 3 | CRR | Clinical Reportable Range |
| 4 | EQAS | External Quality Assurance Scheme |
|  | H&E | Hematoxylin and Eosin |
| 5 | HOD | Head of Department |
| 6 | IQC | Internal Quality Control |
| 7 | LOD | Limit of Detection |
| 8 | LOQ | Limit of Quantitation |
| 9 | MSP | Management System Procedure |
| 10 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 11 | N/A | Not Applicable/Available |
| 12 | pH | Potential of Hydrogen |
| 13 | QC | Quality Control |
| 14 | QM | Quality Manager |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- This document provides steps for microtomy (section cutting) from the paraffin blocks. These thin sections can then be subjected to routine H&E staining/ special staining /or immunohistochemistry procedures and observed under microscope to give definitive diagnosis.

- Scope

- This procedure applies to all staff of the Histopathology Department of TRUSTlab, tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Histopathology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

Any other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definitions

Microtomy: It is the process of cutting the tissues from the paraffin block into uniform thin sections for examination under microscope.

Trimming: It is the process in which the paraffin on the surface of the tissue in the prepared paraffin block is cut so that the tissue surface is exposed for thin sectioning.

- Principle

- Microtomes used in this procedure are precision instruments that will advance an object (paraffin block) for a predetermined distance, then slide the object to the cutting tool knife or blade) and then force the object through the knife thus producing a uniform thin section.

- Specific Performance Characteristics

Uncertainty of measurement: NA

Accuracy: NA

Limit of quantitation (LOQ): NA

Limit of detection (LOD): NA

Analytical measurement range (AMR): NA

Clinical reportable range (CRR): NA

Sensitivity: NA

Specificity: NA

- Type of Sample

Paraffin blocks of embedded tissue and paraffin blocks submitted for second opinion or for immunohistochemistry.

Sample Stability: The paraffin blocks are stable at RT till the archival period (10 years).

Sample Retention: 10 years.

Sample Rejection Criteria:

      - No tissue present in the paraffin block.

- Type of Container & Additives

- Paraffin Blocks in block Mailers.

- Patient Preparation

N/A.

- Equipment & Reagents

Equipment

- Microtome

      - Name: Leica

      - Environmental conditions: Stable ambient temperature

Reagents

      - Gelatin

      - Isopropyl Alcohol

Reagent Storage Condition Stability

- All reagents mentioned above are stable at Room Temperature until the expiry date mentioned on the reagent package.

Additional Requirements

      - disposable microtome blades

      - floatation bath

      - scalpel

      - forceps

      - clean glass slides

      - diamond pencil

      - tissue paper

- Safety

- Safety precautions as mentioned in SOP should be followed.

Be very careful when handling microtome knives or blades. The knife/blade edges are extremely sharp and can cause severe injuries.

Always remove the knife/blade before detaching the knife holder from the instrument.

Knives that are not in use should always be stored in the knife-case.

Never place a knife anywhere with the cutting edge facing upwards and never try to catch a falling knife.

Always clamp the specimen block BEFORE clamping the knife.

Lock the hand wheel and cover the knife edge with the knife-guard prior to any manipulation of knife or specimen, prior to changing the specimen block and during all work breaks.

Always wear safety goggles when sectioning brittle specimens. Specimen may splinter.

Only authorized and qualified service personnel may repair the instrument and access the instrument’s internal components.

Lock the hand wheel (see below) before you start to clean the instrument.

Do not use any solvents containing acetone or xylene for cleaning.

No fluid must enter the interior of the instrument during work or cleaning.

Hand wheel locking mechanism:

      - To activate the locking mechanism, push the handle to the left. The hand wheel locks in place in the 12 o’ clock position and cannot be rotated any further.

      - To unlock the hand wheel, pull the handle outwards (to the right).

Hand wheel brake:

- With the lever on the right side of the microtome base plate, the hand wheel brake can be activated with the handle in any position. When the lever is pushed upward, the hand wheel is locked in position. Two symbols on the microtome base plate indicate the ‘locked’ and ‘unlocked’ position.

In case one’s hand is injured by the blade, inform the pathologist. Use the first aid kit available in the department.

- Calibration

Calibration Schedule

- Calibration is done annually and also when there is a major equipment repair by the authorized service agency of the manufacturer.

- Procedure

Pre-maintenance Checklist

      - Always wear disposable laboratory gloves

      - Move specimen clamp to the upper limit position and lock the hand wheel

      - Remove the blade from the blade holder and put it into the receptacle for used blades

      - Remove specimen from specimen clamp

      - Remove section waste with a dry brush

Cleaning of the Outside Surface of the Microtome (if necessary)

      - The varnished outside surfaces are cleaned with non-commercial household cleaner or soap water and then dried with a cloth/wipe.

Cleaning of the Knife-Holder

      - Remove the front pressure plate for cleaning.

      - Fold knife guard downward.

      - Rotate blade clamping lever downward.

      - Remove the blade carefully.

      - Pull out the clamping lever sideways.

      - Remove the pressure plate.

      - Clean with a wipe dipped in mild detergent.

      - Dry the knife holder and reassemble.

      - When re-assembling, ensure that the upper edge of the pressure plate is parallel to and at the same height as the upper edge of the rear pressure plate.

Lubricate the clamps and levers, monthly or earlier, when required.

Procedure

- Before starting the procedure, paraffin blocks are kept on a cold plate for sufficient hardening of the paraffin wax.

      - Clamping the Paraffin Block

        - Always clamp the paraffin block BEFORE clamping the knife.

        - Lock the hand wheel (see 4.16.13) and cover the knife edge with the knife guard prior to any manipulation of knife or paraffin block, prior to changing the paraffin block and during all work breaks.

        - Rotate handwheel until paraffin block clamp is in the uppermost position.

        - Lock the handle of the handwheel in the 12 o’clock position.

        - Insert a paraffin block into the paraffin block clamp as follows:

        - Turn knurled screw anticlockwise to lower the movable jaw.

        - Insert the block as required so that the cutting surface is in vertical plane.

        - Turn knurled screw clockwise to raise the movable jaw until the block is securely clamped.

      - Inserting the Disposable Blade

        - Be very careful when handling microtome knives or blades. The knife or blade edges are extremely sharp and can cause severe injuries.

        - Carefully insert the disposable blade into the knife holder and clamp as follows:

        - Fold knife guard downward if knife is used.

        - Rotate clamping lever forward to insert the blade.

        - Carefully insert the blade from one side.

        - Rotate clamping lever back upwards to clamp the blade.

      - Adjusting the Clearance Angle

        - The index marks (0°, 5° and 10°) for adjustment of the clearance angle are located on the right side of the knife holder.

        - There is also an index mark on the right side of the knife holder basis which serves as a reference point when adjusting the clearance angle.

        - Rotate lever near the index marks anticlockwise to release.

        - Move the knife holder until the index mark of the desired clearance angle coincides with the reference line on the knife holder base.

        - Hold the knife holder in the selected position and rotate lever clockwise to clamp.

      - Trimming the Paraffin Block

- For trimming, the older used blades can be used. Trimming can be done by the following methods:

        - Trimming the specimen with the coarse feed:

          - Unlock the hand wheel (push handle inward or release brake mechanism via lever, as mentioned in 4.16.14).

          - Rotate coarse feed wheel to bring the paraffin block gradually closer to the knife. At the same time, turn hand wheel to trim the paraffin block to the desired level at which the cutting surface of tissue is completely exposed from where sectioning can be started.

        - Trimming the block via the section thickness selection:

          - Select a section thickness appropriate for trimming (e.g. 50 µm) with the section thickness selection knob. The thickness selected is seen through a window.

          - Rotate hand wheel until the paraffin block has been trimmed down to the desired level at which the cutting surface of tissue is completely exposed from where sectioning can be started.

        - Trimming with the mechanical trimming function:

          - The Leica microtome is equipped with a mechanical trimming function which is activated via de trimming lever.

          - The trimming lever can be locked in 3 positions: 0 µm, 10 µm, and 50 µm.

          - The single dot mark corresponds to 10 µm (• = 10 µm), the double dot mark corresponds to 50 µm (•• = 50 µm).

          - Turn the coarse feed wheel to bring the block close to the knife.

          - Select desired trimming value by pushing the trimming lever to the desired position and hold.

          - With every turn of the hand wheel the paraffin block advances by 10 µm or 50 µm depending on the position of the trimming lever.

          - Rotate the hand wheel to trim the paraffin block down to the desired level at which the cutting surface of tissue is completely exposed from where sectioning can be started.

          - Release trimming lever. When the lever is released, trimming is deactivated (lever flips back to zero position).

      - Sectioning

        - Always rotate the hand wheel evenly with the rotation speed of the hand wheel depending on the hardness of the paraffin block being sectioned. Harder specimens (paraffin blocks) should be sectioned at slower speed.

        - Use the adjusting knob on the right side of the microtome front panel to select the desired section thickness (3 to 5 µm) unless the appropriate value is already seen in the window.

        - For sectioning, use another blade or different cutting edge of the blade, than the one used for trimming.

        - For that purpose, move the knife holder to the desired position or move the knife or blade itself.

        - Start sectioning by rotating the hand wheel evenly in clockwise direction.

        - Pick up the sections and mount them on glass slides.

      - Changing Paraffin Block

        - Lock the handwheel and cover the knife-edge with the knife-guard prior to any manipulation of knife or specimen, prior to changing the specimen block and during all work breaks.

          - Bring the specimen to the uppermost position (via handwheel) and lock the hand wheel in 12 O’clock position.

        - Cover the blade edge with the knife guard.

        - Remove the specimen from the specimen clamp and insert a new specimen block.

        - Use the coarse feed wheel to move the specimen clamp backwards far enough to be able to start trimming the new specimen block

        - NOTE: Alternatively, the required number of blocks can be trimmed together and sectioning is done later.

      - Picking up the Sections

        - During sectioning, a ribbon of sections is formed. Hold the sections with either a blunt forceps or fingers.

        - Place the selected sections in a container filled with water with few drops of alcohol added (alcohol reduces the surface tension of the water and this helps in flattening of the sections). Place the sections smoothly with trailing end of the ribbon making contact with the water first. Remove the folds of the sections with the help of forceps by gently pressing the section with the ends of closed forceps.

        - Select the thin, uniform section without holes from the above ribbon. To take the section on the slide bring a clean slide up to the section in almost vertical position. When the slide touches the section, it is lifted vertically out of water.

        - Transfer the selected sections to the floatation water bath. A floatation bath temperature of 58-60°C is satisfactory for floating cut sections.  Tissue floats and spreads evenly by the temperature of the floatation bath, thus removing any residual folds. Adding a pinch of gelatin to the water in floatation works as section adhesive.

        - Label a slide with the help of a diamond pencil with the HPE number and the year as mentioned on the paraffin block.

        - Take the corresponding section on the slide. Keep the slide on a hot plate for drying.

        - Document daily the total number of paraffin blocks cut in the Microtome log - Histopathology

- Quality Control Procedure

Internal Quality Control

      - Floatation Bath: Change the water of the floatation bath daily to prevent growth of organisms.  After changing the water, add a pinch of gelatin to it. Adding the gelatin helps as a section adhesive for slide. Adjust the temperature of the water bath to approximately 58-60°C. Monitor the temperature of the floatation bath daily. Document this in Tissue Floatation Bath log.

      - After taking the section on slide, remove any section fragments left in the floatation bath with the help of tissue paper. This prevents the ‘floats’ seen on the subsequent slides.

      - Microtome blades are changed after cutting approximately 50 to 60 paraffin blocks. Change the cutting area of the blade after cutting approximately 15 to 20 paraffin blocks.

      - Quality of microtomy is assessed by the Pathologist during reporting on the stained tissue section slides and the same is entered in the IQC-Histopathology Form (TDPL/HP/FOM-003) in histopathology.

External Quality Control

- EQAS participation –- RML EQAS

- Interferences & Cross Reactions

Improper exposure of the section due to less trimming.

Excessive trimming may cause loss of diagnostic material thereby affecting the diagnosis. e.g. tiny endoscopic biopsies where trimming may cause loss of the tissue.

Too many ribbons taken on a glass slide.

Fragments of the tissue sections other than the block being cut can be taken on the slide and thus form a ‘float’. So, remove all such fragments with the help of tissue paper before transferring a new section to the water bath.

Rule of one block cutting one ribbon in the water and the section on the slide should be followed. Too many sections in the floatation bath results in taking the wrong section on the wrongly labeled slide and affects reporting.

Improper spreading of the section results in folds in the sections.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A

- Troubleshooting

Faults due to poor processing

| Fault | Cause | Remedy |
| --- | --- | --- |
| Tissue is shrunken away from the wax | Insufficient dehydration | Reprocess the tissue |
| The tissue is quite soft when the block is trimmed | Insufficient fixation | Reprocess the tissue |
| The tissue appears full of holes | Insufficient infiltration | Re-infiltrate and re-embed the tissue |
| On trimming, wax appears crystalline | -Contaminated wax -Block not cooled rapidly enough | -Re-infiltrate, re-embed in fresh wax -Re-embed and cool rapidly |
| On sectioning the tissue compresses but the paraffin does not | -Insufficient infiltration in wax | -Re-infiltrate and re-embed in wax |
| Section crumbles and drops out of wax leaving a rim of wax as a section | -Insufficient infiltration in wax -Overheated paraffin bath causing tissue hard and brittle | -Re-infiltrate and re-embed in wax -Embed in a hard wax |
| Sections crumble instead of flattening on water bath | Insufficient infiltration in wax | -Re-infiltrate and re-embed in wax |

Faults occurring during section cutting

| Fault | Cause | Remedy |
| --- | --- | --- |
| Ribbon and consecutive sections are curved | -Leading and trail edges of block are not parallel. -Blade blunt in one area. -Surplus area of wax on one side. -Tissue varying in consistency. | -Trim with sharp scalpel until parallel. -Use different part of blade or replace. -Trim away excess wax. -Reorient block by turning through 90o. Cool block with ice. Mount individual sections. |
| Alternate sections are thick and thin | -Block or blade loose -Knife tilt is too high -Wax too soft for tissue or conditions -Insufficient clearance angle--Mechanism of microtome faulty | -Tighten block and blade -Adjust the tilt -Cool block in ice or re-embed in higher melting point wax. -Slightly increase the clearance angle. -Check for any obvious faults, e.g. pawls may be worn. |
| Section shows thick and thin horizontal lines (chatters) | -Knife or block loose in holder -Blunt knife -Excessively steep knife angle -Tissue or wax too hard for sectioning -Calcified areas in tissue. | -Tighten -Change the cutting part of knife or use new disposable knife -Use sharp, heavy-duty knife, reduce slant angle of knife -Rehydrate and decalcify or surface decalcify. |
| Tear or scratch across the section or splitting of the ribbon | -A jagged knife edge -Dirt on the knife edge | -Change the cutting area of the knife or change the knife (disposable) -Clean the knife edge |
| Sections do not join to form ribbon | -Wax too hard for sectioning conditions -Debris on blade edge -Knife angle too steep | -Breath on block to warm or re-embed in lower melting point wax -Clean with a moistened cloth -Adjust to optimum angle |
| Sections expand and disintegrate on water surface | -Poor impregnation of tissue -Water temperature too high | -Return tissue to impregnating bath for few hours -Cool. |
| Sections roll in to a tight coil instead of remaining flat on knife | -Blade is blunt -Too little rake angle -Section thickness too great for wax. | -Use a different part of the blade or replace it. -Reduce blade tilt if clearance angle is excessive -Reduce section thickness or use slightly higher melting point wax. Breath on block as sections are being cut. |
| Section becomes attached to block on return stroke | -Insufficient clearance angle -Wax debris on blade edge -Debris on block edge -Static electric charge on ribbon | -Increase clearance angle -Clean with a moistened cloth -Trim edge with sharp scalpel -Earth microtome to water tap. Humidify air with moist bath near knife. |
| Areas of tissue in block not present in sections | -Incomplete impregnation of tissue -Wax blocks become detached. | -Return tissue to impregnating bath for few hours or reprocess if fault is excessive -Re attach with hot spatula |

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Histopathology technician assesses the quality of the microtomy during the procedure by looking for the faults as mentioned in Trouble shooting.

Pathologist examines the slides made from the tissue sections under microscope for the quality of microtomy and the same is entered in the IQC-Histopathology Form (TDPL/HP/FOM-003) in histopathology.

- Potential Source of Variability

- N/A.

- References

      - Bancroft JD, Gamble M: Theory and Practice of Histological Techniques, 5th edition 2002, Churchill Livingstone U.K., 85-108.

      - K. Lakshminarayanan: Histological Techniques- A Practical Manual, Bhalani Publishing House, Mumbai, 22-31.

      - Leica Microtome: Operating instructions

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)