SOP - 003

PROCEDURE FOR

HIV QUANTITATIVE BY PCR

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	12

15.	Interfering Substances	12

16.	Result Calculation	13

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ART | Antiretroviral Therapy |
| 2 | Ct | Cycle Threshold |
| 3 | DNA | Deoxyribonucleic Acid |
| 4 | EDTA | Ethylenediaminetetraacetic Acid |
| 5 | EQAS | External Quality Assurance Scheme |
| 6 | HIV | Human Immunodeficiency Virus |
| 7 | IC | Internal Control |
| 8 | MI | Microbiology |
| 9 | MIC | Minimum Inhibitory Concentration |
| 10 | NA | Nucleic Acid |
| 11 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 12 | NTC | No-template Control |
| 13 | PCR | Polymerase Chain Reaction |
| 14 | QC | Quality Control |
| 15 | QS | Quality Standard |
| 16 | RNA | Ribonucleic Acid |
| 17 | rpm | Revolutions per minute |
| 18 | RT-PCR | Reverse Transcriptase PCR |
| 19 | SOP | Standard Operating Procedure |
| 20 | STAT | Special Turnaround Time |
| 21 | TAT | Turnaround Time |
| 22 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - The purpose and scope of this document is to describe the procedure for the detection of Human Immunodeficiency Virus - 1 RNA in plasma, serum, whole blood specimen by Real Time PCR.

- Scope

This procedure applies to all staff of the Molecular Biology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- An HIV quantitative test, also known as a viral load test, measures the amount of HIV RNA in a person's blood. This test is crucial for monitoring the effectiveness of antiretroviral therapy (ART) and assessing the progression of the disease. It's a key tool in managing HIV infection, helping healthcare providers make informed decisions about treatment and patient care.

- Test Principle

Pathogen detection by the polymerase chain reaction (PCR) is based on the amplification of specific regions of the pathogen genome. In real-time PCR the amplified product is detected via fluorescent dyes. These are usually linked to oligonucleotide probes that bind specifically to the amplified product. Monitoring the fluorescence intensities during the PCR run (i.e., in real-time) allows the detection and quantitation of the accumulating product without having to re-open the reaction tubes after the PCR run.

- Specific Performance Characteristics

- HIV-1 Quantitative RT PCR kit – NeoDx Biotech Labs.

      - Linearity: 20.98 IU/ml to 107 IU/ml

      - Measurement Range: 20.98 IU/ml to 107 IU/ml

      - Sensitivity: 99.8%

      - Specificity: 99.8%

- Type of Sample

Whole Blood EDTA / Plasma / Serum.

Collect 2.0 ml blood in Lavender top (K2- EDTA). Thoroughly mix the blood with the additives by gently inverting the tube four or five times. Whole blood should be separated into plasma and cellular components by centrifugation for 20 minutes at 3500 rpm within 4 hours.

- Note: Do not use heparin as an anticoagulant as it is an inhibitor of PCR.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

RNA Extraction kit

PCR Amplification Kit

Biosafety Cabinet

Laminar Air flow

Calibrated variable micropipettes, Sterile pipette filter tips (aerosol free)

Vortex mixer

Mini centrifuge with rotor for 1.5 ml /2 ml reaction tubes

Micro Centrifuge 0.2 ml /1.5 ml

Real Time PCR machine

PCR Tubes / strips (0.2 ml)

Cooling block (96 x 0.2 ml tubes)

1.5 ml / 2ml Centrifuge tubes

1.5 ml / 2ml Centrifuge tubes stand

Cooling block (1.5 ml tubes)

Sterile powder free gloves

Lab coats, Facemask, Head cap

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Equipment calibrations are done on an annual basis by an NABL accredited calibration agency or by specific guidelines of equipment manufacturers.

- Test Procedure

Pre-Assay Instructions

- Nucleic Acid Extraction (Isolation): Viral Nucleic Acid purification is carried out according to the instructions in NucleoDx Blood Nucleic Acid Extraction Kit.

    - Recommended sample volume for extraction and elution are as follows:

| Name of the Nucleic Acid Extraction Kit | Recommended Sample volume | Recommended Final Elution volume |
| --- | --- | --- |
| NucleoDx Blood Nucleic Acid Extraction Kit | 200 µl | 60 µl |

    - NeoDx HIV Real time RT-PCR Kit Contents (50 Reactions)

| Contents | 100 Rx | Storage |
| --- | --- | --- |
| Master Mix | 2*750 µl | -20°C |
| Probe and Primer Mix | 150 µl | -20°C |
| Positive Test Control | 100 µl | -20°C |
| Quantitation Standard 1 – QS1 - 104 IU/µl | 350 µl | -20°C |
| Quantitation Standard 2 – QS2 – 103 IU/µl | 350 µl | -20°C |
| Quantitation Standard 3 – QS3 – 102 IU/µl | 350 µl | -20°C |
| Quantitation Standard 4 – QS4 - 10 IU/µl | 350 µl | -20°C |
| Quantitation Standard 5 – QS5 - 1 IU/µl | 350 µl | -20°C |
| Nuclease Free Water | 1000 ml | -20°C |

    - Fluorescent Probe Details

| Target | Reporter | Quencher |
| --- | --- | --- |
| HIV-1 gag protein gene | CY5 | BHQ2 |
| Human Beta Actin | TEXAS RED / ROX | BHQ2 |

    - Reaction Mixture Preparation

| Reagents | Volume per Reaction (35 µl) | Volume per 20 Reaction (17 µl) |
| --- | --- | --- |
| Master Mix | 15 µl | 300 µl |
| Primer and Probe Mix | 1.5 µl | 30 µl |
| Nuclease Free Water | 3.5 µl | 70 µl |
| Sample/ Control/ QS | 15 µl | 15 µl/ reaction |
| Volume | 35 µl | 20 µl Reaction Mix + 15 µl of sample |

    - Cycling Conditions

| Steps | No. of Cycles | Temperature (°C) | Time |
| --- | --- | --- | --- |
| Reverse Transcriptase | 1 | 45 | 20 min |
|  | 1 | 95 | 2 min |
| PCR cycling | 40 | 95 | 15 sec |
|  |  | 58 | 30 sec |

- Note: Fluorescence data should be collected during the 58°C/30sec extension step.

Preparation

      - Proper disinfection practices of the Biosafety Cabinet, Laminar Flow, Pipettes and other minor supplies used in the process.

      - Bring all the reagents and specimens to the required temperature before beginning the test. The reagents should not be subjected to frequent temperature fluctuations.

      - Make sure that the cooling block is pre-cooled to 2-8°C.

      - Do not use kit components beyond the expiration date.

Procedure

- The procedural sequence of reagent addition should be strictly adhered to, to avoid any discrepant results.

      - Before each use, all reagents need to be thawed completely, mixed (by repeated pipetting), and centrifuged briefly.

      - Calculate the number of reactions for each experiment including all controls with one excess reaction volume in the reaction cocktail to accommodate pipetting errors.

      - Assay should be run along with positive and negative controls and standards.

      - Prepare the reaction mix in a 1.5/2ml tube for the calculated number of samples referring to the Reaction Mixture table in Master Mix Preparation room.

      - Dispense 20 µl of reaction mix in each tube or strip.

      - Add 15 µl of nuclease free water to NTC well.

      - Carefully add 15 µl of nucleic acid extract to the wells in template addition room.

      - Add 15 µl of Quantitative Standard in the designated wells in Positive Control addition area in the designated wells.

      - Add 2 µl of Positive Test Control and make up the volume by adding 13 µl of nuclease free water.

      - Seal the plate carefully and spin down and set the program in RT PCR instrument and run the PCR reaction in the PCR Equipment – Quant Studio 5.

- Analysis should be performed independently using threshold settings after analysis of the Standard Curve for number of IU/µl. Each sample is analysed along with an internal control and NTC to verify the sample run. After the completion of the run the samples are analysed and the values are compared to standard/ positive and negative control for qualitative result. The threshold should be set manually just above the background signal of the negative controls and negative samples. The threshold value will usually be in the range of 0.05–0.25.

- Quality Control Procedure

Internal Quality Control Procedure

      - Ideally in HIV Viral Load testing five levels of control as provided by the manufacturer are required to be incorporated for Kit and standard calibration, A negative control/NTC and a positive template control is also included in each run. This ensures that the laboratory is capable of detecting the entire range of values of the analyte in the clinical samples. QC must be acceptable (in control) before the testing and reporting of results. Any test results obtained when QC is unacceptable or not performed are invalid and must be repeated. The person performing testing must evaluate QC results and accordingly determine validity of the test run.

      - All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – MHL EQAS

- Interfering Substances

- Several factors may interfere with and affect the Real-Time PCR amplification:

Elevated levels of bilirubin (≤15 mg/dl) and lipids (≤800 mg/dl) and hemolytic samples do not influence the system.

Heparin (≤10 IU/ml) affects the PCR. Samples that have been collected in tubes containing heparin as an anticoagulant should not to be used. Also, samples of heparinized patients must not be used.

Samples from delayed transport conditions.

Contamination of reagents – can show amplification of non-template control (NTC) which is usually molecular grade water added in place of sample.

- Result Calculation

- After the completion of the run, go to the analysis window. The run profile shows positive and negative samples based on the Ct values for HIV RNA target (CY5) Flourophore and IC (ROX) Flourophore. Make sure that IC for each sample is picked almost at same Ct which indicates high fidelity of real time run.

Interpretation of Qualitative Results

| Observation | Interpretation |
| --- | --- |
| - Presence of sigmoid curve and Ct value in the result column of the sample for both CY5 and ROX channel | HIV RNA DETECTED |
| - Presence of sigmoid curve and Ct value in the result column only for IC ROX channel but not for CY5 channel. | HBV RNA Less than detection limit of 20 IU/ml |
| - The negative control should not show any Ct value in the result column for CY5 channel. - The standard/positive control should show sigmoid and Ct value in the result column for CY5. | Valid Control |

- NOTE: Absence of sigmoid curve and Ct for Sample/Positive control/Negative control – If both target and IC are not amplified – There could be problem in the extraction of nucleic acid from sample or there could be problem in the PCR real time reaction reagents. This can be cross checked by re extraction of sample or setting up real time with another valid real time kit.

Interpretation of Quantitative Results

- For the analysis of samples in each run include four standards and one NTC along with unknown samples. Each sample is analysed along with an internal control and NTC to verify the sample run. After the completion of the run the samples are analysed. Interpret the values for unknown samples, only if the Slope of Standards is between -3.1 to -3.6 and PCR efficiency is between 90%-110% (0.9 – 1.1) and there should be no amplification in negative control.

Result IU/ml=Result IU/µl ×Elution Volume (µl - 60) Sample Volume (ml-0.2)

To convert the results from IU/µl to IU/ml, use the following formula:

- Conversion Factor

- The conversion factor 1 IU/ml = 1.74 HIV copies/ml

- Standard Curve Control Parameter

| Control Parameters | Valid Value |
| --- | --- |
| Slope | -3.0 / 3.74 |
| PCR Efficiency | 85% / 115% |
| R Square | 0.95-1 |

- Biological Reference Interval

- Not detected.

- Reportable Interval

- The reporting of results of HIV RNA detection and quantification is as per turnaround time (TAT). The turnaround time for reporting is 72 hours from the time of receiving samples to lab. In case of urgent requests from clinician or referral lab, the sample is considered under STAT (Special Turnaround time).

- Procedure when Results are not in Range

| Observation | Possible cause | Recommendation |
| --- | --- | --- |
| Standard curve does not show expected parameters of PCR efficiency and slope | Improper mixing of calibrators | While putting PCR run, properly thaw, and mix the viral calibrators before use. |
|  | Inappropriate storage of reagents | Store the reagents at recommended temperature for its optimal performance. Avoid repetitive freezing and thawing. Check expiry of reagents. In that case use new lot of reagents. |
| HIV target growth curve in extraction control | Cross contamination of samples occurred during RNA extraction procedures or assay setup. | Repeat the assay with stricter adherence to good laboratory guidelines. |
| No signal for IC target in sample | Presence of RT-PCR inhibitors from clinical specimens | Repeat the test from fresh sample following good laboratory practices and strict adherence to assay protocol |
|  | Improper assay set up and execution |  |

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Interpret the values for unknown samples, only if the R2 > 0.98.

Slope of calibrators is between -3.1 to -3.6.

PCR efficiency is between 90% - 110% and no amplification in CY5 channel of negative control.

Based on the Standard Curve, the target in the sample is quantified.

- Potential Source of Variability

If the DNA recovery in extraction procedure is poor, the Ct value of HIV RNA target would be high as it picks up at late Cts.

Haemolysed samples and delayed transport samples often give low yield or no yield of HIV RNA and often yield false negative results.

It is important to pipet the indicated quantities and mix well after each reagent addition. Check pipettes regularly.

Strict compliance with the user manual is required for optimal PCR results.

Attention should be paid to expiration dates printed on the kit box and labels of all components. Do not use expired components.

Although rare, mutations within the highly conserved regions of the viral genome covered by the kit’s primers and/or probe may result in under quantitation or failure to detect the presence of the virus in these cases. Validity and performance of the assay design are revised at regular intervals.

The presence of PCR inhibitors may cause false negative or invalid results.

- References

- NeoDx HIV-1 Quantitative PCR kit Handbook