SOP - 004

PROCEDURE FOR

MEDIA PREPARATION & STERILIZATION

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	7

8.	Patient Preparation	7

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	8

13.	Test Procedure	9

14.	Quality Control Procedure	21

15.	Interfering Substances	21

16.	Result Calculation	21

17.	Biological Reference Interval	21

18.	Reportable Interval	21

19.	Procedure when Results are not in Range	21

20.	Critical Alert Values	21

21.	Laboratory Clinical Interpretation	21

22.	Potential Source of Variability	21

23.	References	22

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ATCC | American Type Culture Collection |
| 2 | CLSI | Clinical & Laboratory Standards Institute |
| 3 | Eh | Electron-hydrogen potential |
| 4 | MI | Microbiology |
| 5 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 6 | pH | Potential of Hydrogen |
| 7 | rpm | Revolutions per minute |
| 8 | SOP | Standard Operating Procedure |
| 9 | TDPL | TRUSTlab Diagnostics Private Limited |
| 10 | UV | Ultraviolet |
| 11 | UTI | Urinary Tract Infection |
| 12 | WHO | World Health Organization |

- Objective

  - This SOP describes the preparation and sterilization of various culture media used for bacteriology testing process.

- Scope

This procedure applies to all staff of the Microbiology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Culture media in microbiology are specialized substrates designed to support the growth of microorganisms in a laboratory setting.

- Test Principle

Cultivation of bacteria in vitro requires a nutritious medium with all essential minerals, adequate PH and environment in a solid, liquid or semisolid phase culture medium. These cultures are presterilized in an autoclave to avoid growth of environmental contaminant bacteria or fungi.

- Specific Performance Characteristics

- N/A.

- Type of Sample

- N/A.

- Patient Preparation

- N/A.

- Type of Container & Additives

- N/A.

- Equipment & Reagents

Dehydrated Media

Disposable Petri dishes

Digital balance

Distilled Water

pH Paper

Sterile Glass ware

Cotton roll

Craft paper

Autoclave

Laminar Air Flow Chamber

Autoclave indicator

Marker/Sticker

- Safety

The use of disposable gloves and proper protective clothing should be used for media preparation.

Do not pipette by mouth.

Materials used should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent before and after media preparation. Consult a physician immediately in case of accidental contact with mucosa or skin.

All precautionary measures for autoclave operation to be followed.

- Calibration Procedure

- Calibration of Laminar flow & weighing balance is done on an annual basis by an NABL accredited calibration agency. Necessary maintenance measures are undertaken for all the equipment.

- Test Procedure

General Instructions

      - Prepare media in a separate area of the microbiology lab.

      - Observe strict aseptic precautions while preparing media.

      - Wear face mask buttoned lab coats & disposable heat protective gloves while preparing media.

      - Use Sterilized Petri dishes & sterile glassware while preparing the media.

      - Wipe the worktable with 70% ethanol prior to preparing and pouring the media onto plates.

      - Measure the requisite amount of water and pour it in a heat-resistant beaker or a big size conical flask. Ensure that the capacity of the container is larger than the volume of medium to be prepared.

      - Use only type 1 reagent grade water for dissolving/reconstituting the dehydrated media.

      - Weigh the exact quantity of the dehydrated media/ingredients in the digital balance according to the volume of media and dissolve it in of distilled water in a clean conical flask.

      - Boil, to dissolve the medium completely.

      - Check the pH of the media using pH paper.

      - Sterilize the media by autoclaving at 15 lbs pressure (121˚C) for 15 minutes in an autoclave.

      - Attach a strip of thermal indicator with each batch of media being autoclaved.

      - Sterilize the laminar flow cabinet by using UV lamp for 10-15 minutes prior to pouring the autoclaved media into the sterile petri dishes (20-25ml) or sterile test tubes (1-5ml) as butts or slants covering ¾ of the test tube after cooling to 45-50°C.

      - Label the plates (name of the media, date of preparation, expiry date 3 weeks from date of preparation) and store the media in the refrigerator at 2 - 8˚C in appropriate sealed manner to avoid contamination.

      - Check the sterility of the media prepared by incubating one plate from the lot prepared for 24-48 hours.

      - Check the quality of the media prepared for ability to support growth on appropriate strain of bacteria (ATCC or isolate from a clinical specimen).

Mac Conkey Agar

| Description | Dehydrated commercially available media |
| --- | --- |
| Appearance and Composition grams/litre | Appearance Light yellow to pink homogeneous free flowing powder. Peptone 17.0; Proteose peptone 3.0; Lactose 10.0; Bile salts 1.5; Sodium chloride 5.0; Neutral red 0.03; Agar 15.0; Final pH (at 25°C) 7.1±0.2 |
| Preparation | - 5.5 g of Mac Conkey Agar powder in 100 ml of Distilled Water. - Prepared as plate. |
| Principle | Mac Conkey agar is a selective and differential medium. Lactose-utilizing bacteria utilize lactose and produce acid. At the high level of acid production neutral red indicator present in media turns to pink color so the lactose fermenting colonies appear pink in colour. Non-lactose fermenting bacteria do not ferment lactose hence no acid is produced and no change in media pH or color of the colony is observed. |
| Application | Used for differentiating lactose fermenting and non-lactose fermenting bacteria in clinical samples. |
| Quality Check | Escherichia coli (ATCC 25922) – Lactose fermenter Pseudomonas aeruginosa (ATCC 27853) – Non-Lactose fermenter |

Muller Hinton Agar

| Description | Dehydrated commercially available media |
| --- | --- |
| Appearance and Composition grams/litre | Cream to yellow homogeneous free flowing powder HM infusion B 300.0; Acicase 17.5; Starch 1.5; Agar 17.0 Final pH (at 25°C) 7.3±0.1 |
| Preparation | - 38.0 grams in 1000 ml purified/ distilled water. - Prepared as plate. |
| Principle | Mueller Hinton Agar is recommended for the diffusion of antimicrobial agents impregnated on paper disc through an agar gel as described in CLSI Approved Standard. Mueller Hinton Agar has been selected by the CLSI for several reasons: - It demonstrates good batch-to-batch reproducibility for susceptible testing. - It is low in sulfonamide, trimethoprim and tetracycline inhibitors. - It supports the growth of most non-fastidious bacterial pathogens. - A lot of data and much experience regarding its performance have been recorded. Mueller Hinton Agar is not appropriate for assay by disc diffusion method with slow growing organisms, anaerobes and capnophiles. With slow growing organisms, increased incubation may cause deterioration of diffusing antibiotic and produce unprecise readings. |
| Application | Used as a test medium for antimicrobial susceptibility testing. |
| Quality Check | Enterococcus faecalis (ATCC 29212) tested for Cotrimoxazole (25mcg) sensitivity – 24-32 mm zone |

Michrom UTI Agar

| Description | Enriched media dehydrated commercially available. |
| --- | --- |
| Appearance and Composition grams/litre | Cream to homogeneous free flowing powder. Peptone special 15.0; Chromogenous mixture 2.45; Sodium chloride 5.0; Agar 15.0 Final pH (at 25°C) 6.8±0.2 |
| Preparation | - Suspend 32.45 grams in 1000 ml purified/distilled water. - Prepared as plate. |
| Principle | MICHROM UTI agar is a widely used differential medium, which supports the growth of a wide variety of organisms causing urinary tract infections. |
| Application | Used for isolation of urinary pathogens with preliminary identification based on the colour of the colonies. Pink – Escherichia coli; Blue – Klebsiella/ Enterobacter/ Citrobacter; Yellow – Proteus; White – Candida, Staphylococcus; Tiny blue - Enterococcus. |
| Quality Check | Escherichia coli (ATCC 25922) Enterococcus faecalis (ATCC 29212) |

Peptone Water

| Description | Dehydrated commercially available powder. |
| --- | --- |
| Appearance and Composition grams/litre | Cream to homogeneous free flowing powder. Peptone 10.0; Sodium chloride 5.0 Final pH (at 25°C) 7.2±0.2 |
| Preparation | - Suspend 15.0 grams in 1000 ml purified/distilled water. - Stored in sterile test tubes, 3-5ml per tube. |
| Principle | Peptone Water is particularly suitable for preparation of young culture broths for various biochemical reactions and antimicrobial sensitivity testing and as a substrate in the study of indole production by bacteria using tryptophan. Peptone Water with pH adjusted to 8.4 by addition of sodium chloride is suitable for the cultivation and enrichment of Vibrio species. |
| Application | Peptone Water is used as a growth medium and as a base for carbohydrate fermentation media. |
| Quality Check | Escherichia coli (ATCC 25922) – Indole Positive Pseudomonas aeruginosa (ATCC 27853) – Indole Negative |

Nutrient Broth

| Description | Enriched media dehydrated commercially available. |
| --- | --- |
| Appearance and Composition grams/litre | Cream to yellow homogeneous free flowing powder Peptone 5.0; Sodium chloride 5.0; HM peptone B 1.50; Yeast extract 1.5. Final pH (at 25°C) 7.4±0.2 |
| Preparation | - Suspend 13.0 grams in 1000 ml purified / distilled water. - Stored in sterile test tubes, 3-5ml per tube. |
| Principle | Nutrient broth is particularly suitable for preparation of young culture broths for various biochemical reactions and antimicrobial sensitivity testing and as a substrate in the study of indole production by bacteria. |
| Application | Nutrient Broth is used for the general cultivation of less fastidious microorganisms. |
| Quality Check | Escherichia coli (ATCC 25922) – Turbidity |

Urea Agar

| Description: | Enriched media dehydrated commercially available. |
| --- | --- |
| Appearance and Composition grams/litre | Light yellow to light pink homogeneous free flowing powder Peptone 1.0; Dextrose (Glucose) 1.0; Sodium chloride 5.0; Disodium hydrogen phosphate 1.20; Potassium dihydrogen phosphate 0.8; Phenol red 0.012; Agar 15.0. Final pH (at 25°C) 6.8±0.2 |
| Preparation | - Suspend 24.01 grams in 950 ml purified/distilled water. - Cool to 45-50°C. - Aseptically add 50 ml of sterile 40% Urea Solution and mix well. - Prepared as slants in sterile test tubes, 3-5ml per tube. |
| Principle | Urea agar is used to detect urease production by all rapidly urease-positive Proteus organisms and also by other members of Enterobacteriaceae. Urea is hydrolyzed to liberate ammonia. Phenol red indicator detects the alkalinity generated by a visible colour change from orange to pink. Prolonged incubation may cause alkaline reaction in the medium. All urea test media rely on the alkalinity formation and so they are not specific for determining the absolute rate of urease activity The utilization of proteins may raise the pH to alkalinity due to protein hydrolysis and excess of amino acids liberation results in false positive reaction |
| Application | This test can be used as part of the identiﬁcation of several genera and species of Enterobacteriaceae – Proteus, Klebsiella, some Yersinia and Citrobacter species, as well as some Corynebacterium species. It is also useful to identify Cryptococcus spp., Brucella, Helicobacter pylori, and many other bacteria that produce the urease enzyme. Directly, this test is performed on gastric biopsy samples to detect the presence of H. pylori. |
| Quality Check | Escherischia coli (ATCC 25922) – Urease Negative Klebsiella pneumoniae (ATCC 700603) – Urease Positive |

Citrate Agar

| Description | Enriched media dehydrated commercially. |
| --- | --- |
| Appearance and Composition grams/litre | Light yellow to light pink homogeneous free flowing powder Magnesium sulphate 0.2; Ammonium dihydrogen phosphate 1.0; Dipotassium phosphate 1.0; Sodium citrate 2.0; Sodium chloride 5.0; Bromothymol blue 0.08; Agar 15.0. Final pH (at 25°C) 6.8±0 Final pH (at 25°C) 6.8±0.2 |
| Preparation | - Suspend 24.28 grams in 1000 ml purified/ distilled water. - Before using water, ensure the pH of water is 6.5 to 7.0. - Initial colour of the medium may deviate from expected colour, if the above precaution is ignored. - Prepared as slants in sterile test tubes, 3-5ml per tube. |
| Principle | These media are used for the differentiation of bacteria on the basis of citrate utilization as sole carbon source. It is recommended that APHA Ammonium dihydrogen phosphate and sodium citrate serve as the sole nitrogen and carbon source respectively. Microorganisms also use inorganic ammonium salts as their sole nitrogen source. Metabolism of these salts causes the medium to become alkaline, indicated by a change in colour of the pH indicator from green to blue. Bromothymol blue is the pH indicator. Colour change along with growth on the slant should be checked for citrate utilization. The medium should be freshly prepared because in dry conditions, changes in colour may appear even before inoculation, especially at the bottom of the slant. |
| Application | Recommended for differentiation the members of Enterobacteriaceae on the basis of citrate utilization. |
| Quality Check | Escherischia coli (ATCC 25922) – Citrate Negative Klebsiella pneumoniae (ATCC 700603) – Citrate Positive |

Triple Sugar Iron Agar (T.S.I)

| Description: | Enriched media dehydrated commercially available |
| --- | --- |
| Appearance and Composition grams/litre | Light yellow to light pink homogeneous free flowing powder Beef extract 3.0; Peptone 20.0; Yeast extract 3.0; Lactose 10.0 Sucrose 10.0; Dextrose monohydrate 1.0; Ferrous sulphate 0.2; Sodium chloride 5.0; Sodium thiosulphate 0.3; Phenol red 0.024; Agar 15.0. Allow the medium to set in sloped form with a butt about 2.5cm long. to dissolve the medium completely. Mix well and distribute into test tubes and sterilize by maintaining at 10lbs pressure (115°C) 12.000 |
| Preparation | - Suspend 64.42 grams (the equivalent weight of dehydrated medium per litre) in 1000 ml purified/ distilled water. - Allow the medium to set in sloped form with a butt about 2.5cm long and a slant with slant to butt ratio of 1:2. |
| Principle | Organisms that ferment dextrose monohydrate produce a variety of acids, varying the colour of the medium from red to yellow. More amounts of acids are liberated in butt region (fermentation) than in the slant (respiration). Growing bacteria also form alkaline products from the oxidative decarboxylation of peptone and these alkaline products neutralize the large amounts of acid present in the butt changing the medium to pink. Gas production (CO2) is detected by the presence of cracks or bubbles in the medium, when the accumulated gas escapes. Thiosulphate is reduced to hydrogen sulphide by several species of bacteria and H2S combines with ferric ions of ferric salts to produce the insoluble black precipitate of ferrous sulphide. Reduction of thiosulphate proceeds only in an acid environment and blackening usually occurs in the butt of the tube. Yellow butt/ yellow slant: Bacteria that ferment lactose or sucrose (or both), in addition to dextrose Yellow butt/ Pink slant: Organism is dextrose fermenter but does not ferment lactose and/or sucrose. Pink Butt / Pink slant: Non-fermenter of sugars, facultative anaerobe No change in butt/ Pink slant: Non-fermenter of sugars, strict aerobe |
| Application | Triple Sugar Iron Agar Medium is recommended for identification of gram-negative enteric bacilli on the basis of dextrose, lactose and sucrose fermentation and hydrogen sulphide production |
| Quality Check | Escherischia coli (ATCC 25922) – Acid / Acid Negative Pseudomonas aeruginosa (ATCC 27853) – Alkali / No change |

XLD Agar

| Description | Enriched media dehydrated commercially available |
| --- | --- |
| Appearance and Composition grams/litre | Yeast extract 3.0; L-Lysine 5.0; Lactose 7.5; Sucrose 7.5 Xylose 3.50; Sodium chloride 5.0; Sodium deoxycholate 2.5 Sodium thiosulphate 6.8; Ferric ammonium citrate 0.8; Phenol red 0.08; Agar 15.0 Final pH (at 25°C) 7.4±0.2 |
| Preparation | - Suspend 56.68 grams in 1000 ml purified/distilled water. - Heat with frequent agitation until the medium boils. DO NOT AUTOCLAVE OR OVERHEAT. - Transfer immediately to a water bath at 50°C. After cooling, pour into sterile Petri plates. It is advisable not to prepare large volumes that will require prolonged heating, thereby producing precipitate. Prepared as plates. |
| Principle | The media formulation does not allow the overgrowth of other organisms over Salmonella and Shigella. The medium contains yeast extract, which provides nitrogen and vitamins required for growth. The sugars xylose, lactose and sucrose provide sources of fermentable carbohydrates, but xylose is mainly incorporated into the medium since it is not fermented by Shigellae but practically by all enterics. This helps in the differentiation of Shigella species. Sodium chloride maintains the osmotic balance of the medium. Lysine is included to differentiate the Salmonella group from the non-pathogens. Salmonellae rapidly ferment xylose and exhaust the supply. Subsequently lysine is decarboxylated by the enzyme lysine decarboxylase to form amines with reversion to an alkaline pH that mimics the Shigella reaction. However, to prevent this reaction by lysine-positive coliforms, lactose and sucrose are added to produce acid in excess. Degradation of xylose, lactose and sucrose to acid causes phenol red indicator to change its colour to yellow. Bacteria that decarboxylate lysine to cadaverine can be recognized by the appearance of a red colouration around the colonies due to an increase in pH. |
| Application | Used for the isolation and enumeration of Salmonella Typhi and other Salmonella species from clinical and nonclinical samples. |
| Quality Check | Escherichia coli (ATCC 25922) – Yellow |

Selenit F Broth

| Description: | Enriched media dehydrated commercially available |
| --- | --- |
| Appearance and Composition grams/litre | A cream to yellow coloured homogeneous free flowing powder. Colour and Clarity of prepared medium: Light to medium amber-coloured, clear solution with minimum precipitate. Part A - Tryptone 5.0; Lactose 4.0; Sodium phosphate 10.0 Part B - Sodium hydrogen selenite 4.0 Final pH (at 25°C) 7.0±0.2 |
| Preparation: | - Suspend 4.0 grams of Part B in 1000 ml distilled water. - Add 19.0 grams of Part A. - Mix well. - Warm to dissolve the medium completely. - Distribute in sterile test tubes. - Sterilize in a boiling water bath or free flowing steam for 10 minutes. - DO NOT AUTOCLAVE. Excessive heating is detrimental. - Discard the prepared medium if large amount of selenite is reduced (indicated by red precipitate at the bottom of tube/bottle). |
| Principle: | Selenite F Broth is used as enrichment medium for isolation of stool pathogens like Salmonella, Shigella as they are present in a very small number in the intestinal flora. Tryptone provides nitrogenous substances. Lactose maintains the pH of medium. Selenite is reduced by bacterial growth and alkali is produced. An increase in pH lessens the toxicity of the selenite and results in overgrowth of other bacteria. The acid produced by bacteria due to lactose fermentation serves to maintain a neutral pH. Sodium phosphate maintains a stable pH and also lessens the toxicity of selenite |
| Application: | Used for the propagation and isolation of stool pathogens. |
| Quality Check | Escherichia coli (ATCC 25922) – Growth inhibited |

Fluid Thioglycollate Medium

| Description | Enriched media dehydrated commercially available |
| --- | --- |
| Appearance and Composition grams/litre | A cream to yellow homogeneous free flowing powder. Colour and Clarity of prepared medium: Light to medium amber coloured, clear solution without any precipitate Sterile clear BHI broth in bottles. Tryptone 15.0; Yeast Extract 5.0; Dextrose 5.5; L-Cystine 0.5; Sodium chloride 2.500; Sodium Thioglycollate 0.5; Resazurin Sodium 0.001; Agar 0.75. Final pH (at 25°C) 7.1±0.2 |
| Preparation: | - Suspend 29.75 grams in 1000 ml purified/distilled water. - Prepared as test tube broths. |
| Principle: | Fluid thioglycolate medium is used as enriched medium for isolation of fastidious pathogens with anaerobic requirements. Glucose monohydrate, tryptone and yeast extract provide the necessary nutrients and growth factors for bacterial multiplication. Sodium Thioglycollate & L-cystine lower the oxidation-reduction potential of the medium by removing oxygen to maintain a low Eh. By creating an environment with a low Eh, the reducing agents prevent the accumulation of peroxides which can be toxic to some organisms. The sulphahydryl groups (-SH) of these compounds neutralizes the antibacterial effect of mercurial preservatives and other heavy metal compounds which exert a bacteriostatic effect in the materials under examination. Any increase in the oxygen content is indicated by a colour change of redox indicator, resazurin to pink. The small amount of agar helps in maintaining low redox potential for stabilizing the medium. This medium is nutritious and well buffered to support the growth of a wide variety of organisms. |
| Application: | Used for the propagation of pathogenic cocci and other fastidious organisms. Can also be used for preservation of organism as stock. |
| Quality Check | Escherichia coli (ATCC 25922) - Turbidity |

Bile Aesculin Agar

| Description | Enriched media dehydrated commercially |
| --- | --- |
| Appearance and Composition gram/litre | Light yellow to light brown homogeneous free flowing powder Peptic digest of animal tissue 5.0; Beef extract 3.0; Esculin 1.0; Bile salts 40.0; Ferric citrate 0.5; Agar 15.0 Final pH (at 25°C) 6.8±0 Final pH (at 25°C) 6.6±0.2 |
| Preparation | - Suspend 64.5 grams in 1000 ml purified/ distilled water. - Prepared as slants or butts in sterile test tubes, 3-5ml per tube. |
| Principle | Organisms hydrolyze esculin to esculetin and dextrose. Esculetin further reacts with ferric citrate to form a dark brown or black complex. Used for identification of Enterococcus genera and Yersinia enterocolitica |
| Application | Recommended for differentiation the members of Enterobacteriaceae on the basis of citrate utilization. |
| Quality Check | Enterococcus faecalis (ATCC 29212) – Positive |

Sterilization Procedure by Autoclave

- Autoclaving can sterilize anything that can withstand a temperature of 121°C for 20 minutes and pressure of 15lbs. A pressure cooker used in homes for cooking purposes can also be used as a makeshift autoclave.

- Operating Instructions

      - Ensure that there is sufficient water inside the chamber.

      - Load the autoclave and fasten the lid keeping the discharge tap open.

      - Adjust the safety valve to the required temperature and turn the heat on.

      - Allow the mixture of air and steam to pass out freely till air has been discharged.

      - Close the air discharge tap and let the steam pressure rise within the chamber till it attains a temperature of 121°C at 15lbs, for 15-20 minutes.

      - Hold on the pressure for 15 minutes.

      - Turn off the heat and let the autoclave cool.

      - Slowly open the air and steam discharge taps after the pressure gauge has reached zero.

      - Allow the material to cool before these are handled (usually agar bottles take hours before these become safe to handle.

- Quality Control Procedure

Quality check for media is specified in respective media preparation.

Control for Media Autoclave: Steam indicator tapes are used with every batch of media preparation.

Control for Discard Autoclave: Autoclaves can be used for disinfection of infectious waste before it is discarded.

Steam indicator tapes are used with every batch and strips impregnated with spores of Geobacillus stearothermophilus are run weekly to check for the efficacy of autoclave.

- Interfering Substances

- Proper preparation, maintenance of pH, autoclaving, appropriate volume of media in plates or tubes at right temperature, storage, drying of media before use and proper sealing during incubation is required for proper performance of media.

- Result Calculation

Performance of media is specified in respective media preparation.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- The media whose performance is not appropriate are discarded and fresh media is prepared and used.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

Clinical Microbiology Standard Operating Procedures from W.H.O

District Laboratory Practices in Tropical countries by Monica Cheeseborough Part – II

Practical Medical Microbiology by Mackie &Mc Cartney 14th edition, Page no: 815-823