SOP - 007

PROCEDURE FOR CONVENTIONAL BACTERIAL IDENTIFICATION

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	21

15.	Interfering Substances	21

16.	Result Calculation	22

17.	Biological Reference Interval	22

18.	Reportable Interval	22

19.	Procedure when Results are not in Range	22

20.	Critical Alert Values	22

21.	Laboratory Clinical Interpretation	22

22.	Potential Source of Variability	22

23.	References	23

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ATCC | American Type Culture Collection |
| 2 | BAL | Bronchoalveolar Lavage |
| 3 | ET | Endotracheal |
| 4 | MI | Microbiology |
| 5 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 6 | SOP | Standard Operating Procedure |
| 7 | TDPL | TRUSTlab Diagnostics Private Limited |
| 8 | TSI | Triple Sugar Ion |

- Objective

  - This SOP describes the procedure for identification of possible pathogens in various sample inoculated in the culture media after incubation for appropriate time and temperature.

- Scope

This procedure applies to all staff of the Microbiology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Conventional bacterial identification relies on a combination of microscopic examination, cultural techniques, and biochemical tests to determine the characteristics of a bacterium. These methods help differentiate bacteria based on their morphology, staining properties, colony characteristics, and biochemical reactions.

- Test Principle

The growth properties and biochemical properties of potential pathogens are employed to identify causative pathogens and reported as a pathogen based on the sample type.

The pathogen is presumptively identified by its colony characteristics, culture media changes, staining properties, biochemical properties.

- Specific Performance Characteristics

- Refer to Test Procedure.

- Type of Sample

- Bacterial colony isolated on specific culture medium.

- Patient Preparation

- N/A.

- Type of Container & Additives

- N/A.

- Equipment & Reagents

Calibrated Inoculation loop

Microscope

Glass Slides

Acid Fast stain kit

Bunsen burner

Culture Media

Marker

Incubator

Alcohol spirit

Cotton/Tissue paper

Biosafety Cabinet

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

Suitable infection prevention precautions to be followed.

- Calibration Procedure

Biosafety Cabinet, Incubator and autoclave calibrations are done on an annual basis by an NABL accredited calibration agency. Necessary maintenance measures are undertaken for all the equipment.

- Test Procedure

Precautions

      - Liquid media may be inoculated first when processing fluid specimens. This reduces the chances of carry-over from contaminated solid media. However, liquid media should be inoculated after the solid media when swabs and faeces are examined, to avoid diluting the organisms contained on or in the sample and to avoid any non-viable organisms present in a liquid medium being transferred to other liquid media, solid media or to slides.

      - Smears for staining are usually made after all culture media have been inoculated to avoid carry-over of contaminants that may be on the surface of the slide. However, there may be occasions where the smear is required prior to culture and then a sterile slide should be used.

      - For the isolation of individual colonies, the inoculum should be spread, usually by using a sterile loop, taking care to avoid the edges of the plate where contaminants are more likely to be located.

      - Antimicrobial discs for identification (e.g. Optochin, Bacitracin) may be added as appropriate. Discs should be placed near the edge of the plate, between the areas covered by the first and second spread, to avoid total inhibition of very susceptible organisms.

      - As a minimum requirement, all culture plates and containers must be labeled to identify the patient’s name or laboratory number or barcode. Additional labeling, including date of culture or sub-culture will be necessary for selected specimens, such as those requiring prolonged incubation or sub-culture from enrichment broth.

      - To work safely and minimize risks of cross contamination suitable racks should also be used when inoculating, incubating or storing liquid cultures or culture plates.

Aseptic Techniques

      - When handling specimens or cultures, aseptic technique is important to avoid their contamination and to protect the worker from infection from the sample.

      - The following points should be observed when culturing specimens or performing subcultures:

        - Caps and lids from culture media should not be placed on the workbench but retained in the hand or placed in such a way while the sample is being processed, taking care not to contaminate the hand or cap. Caps and lids should be replaced as soon as possible.

        - If the work is being carried out on the open bench, a Bunsen burner/loop sterilizer should be in close proximity to sterilize loops or wires.

        - Keep samples away from the face when opening culture containers.

        - Aerosol production should be minimized by:

          - Opening caps slowly as the contents of containers are sometimes under pressure.

          - Avoiding vigorous swirling or shaking of the sample prior to opening.

          - Cooling loops that have been heated before use.

          - Avoiding expelling the last drop from a pipette.

          - If forceps or scissors are used when handling specimens, they should be heated in a Bunsen flame/loop sterilizer and allowed to cool before use.

          - Loops and wires should be heated in the Bunsen flame/loop sterilizer with the loop end downwards until red-hot, then removed and allowed to cool. Care should be taken that the loop does not contain fluid or large particles of matter that may ‘splatter’ when placed in the flame, and that the entire loop\wire up to the loop holder is heated to red-hot state.

Procedure

    - Grams Stain

- Colonies may be identified by using the Gram’s staining as Gram-Positive or Gram-Negative – coccus, coccobacillus, bacillus.

    - Colony Characteristics

Growth on Blood Agar

Growth on Blood Agar

Growth Characteristics

Growth on Urichrom Agar

    - Motility Test

- A peptone water emulsion of a colony of bacteria is prepared and incubated for 30 min to 2 hours. 1-2 drops of subculture emulsion is placed on a slide and covered with a cover slip and observed under microscope with 10 x, 40x. Motile organisms and nonmotile organisms are identified. Using a concavity slide would give a better result. Examples of motile organisms are – Escherichia coli, Enterobacter species, Pseudomonas. Examples of nonmotile organisms are – Klebsiella, Acinetobacter.

    - Biochemical Reactions

      - Catalase Test

- Detects the presence of enzyme catalase in the bacteria which helps in breakdown of hydrogen peroxide into water and oxygen. The Catalase is most commonly used for the differentiation of members of Micrococcacae from those of Streptococcacae.

- Materials Required: Hydrogen peroxide (3%), 18-24 hours culture of the organism to be tested, Tubes and slides, Positive control and Negative control organisms and Capillary tubes.

- Control Organisms:

- Positive control: Staphylococcus aureus (ATCC 25923)

- Negative control: Enterococcus faecalis (ATCC 29212)

- Test Procedure

- Tube Test: Take three sterile clean tubes and 3 % hydrogen peroxide (1ml) to the tube. Pick-up a small portion of the colony with the help of toothpick or plastic stirrer of the positive and negative control strains and the test colony and drop the stick into the respective tubes of hydrogen peroxide. Check for reaction within 10 seconds.

- Slide Test: Take a slide and clean it so as to make it grease free. Add a drop of 3% H2O2 at three places on the slide. Pick up a small portion of the colony with the help of a toothpick or plastic stirrer of the positive and negative control strains and the test colony and add to the three designated drops of hydrogen peroxide. Wait for 10 seconds for the reaction to occur.

- Results & Interpretation

- Sustained appearance of bubbles or effervescence in the slide or in the tube is a positive Catalase test as seen with positive control. No appearance of bubbles is a negative catalase test as seen with negative control. Interpret the result of the test colony by the presence of bubble - positive for catalase or absence of bubbles – negative for catalase test.

- Precautions

- The reading should be taken within 10 seconds as the test is very time sensitive.

- A few organisms could give a false positive reaction due to the presence of enzymes other than Catalase which could decompose H2O2. This could lead to the formation of a few bubbles at the bottom of the tube or on the slide after 20 seconds. This should be read as a negative reaction.

- Red blood cells are known to have a positive Catalase activity. Care should be taken while picking colonies from blood agar to avoid picking up of the media with applicator stick. The Hydrogen peroxide bottle should be labeled and stored away from all other reagents. If the bottle is left open it could lead to the deterioration of the reagent, and this again interferes with the test result. This fact could be made out when the standard strain of Staphylococcus aureus does not give a positive result or gives a weak reaction.

- Do not use a Nichrome or a stainless-steel wire for the test as oxidation products at the top of the loop could interfere with the test and give wrong results.

      - Oxidase Test

- Detects cytochrome oxidase enzyme that helps in transfer of electrons to oxygen from cytochrome c as the last link in the chain of aerobic respiration with the formation of water. The dye (tetra methyl para phenyl diamine dihydrochloride 1%) in the test as mentioned above acts as a substitute for oxygen as the artificial electron acceptor in the reduced state. The dye is colorless, however, in the oxidized state, in the presence of oxidase enzyme and oxygen, the dye forms a coloured indophenol blue which is deep blue in colour. Enterobacteriaceae members give negative oxidase test results, while Pseudomonads give positive oxidase results.

- Materials Required

- Oxidase discs (commercially available which are impregnated with the reagent i.e. Tetra methyl para phenyl diamine dihydrochloride 1%), plastic applicator sticks or toothpicks, glass slide, Petri dish.

- Control Organisms

  - Positive control: Pseudomonas aeruginosa (ATCC 27853)

  - Negative control: Escherichia coli (ATCC 25922)

- Test Procedure

- Place three oxidase reagent disc is placed on a clean glass slide and a colony of control strains are applied respectively on each disc and results observed within 10 seconds. If the controls trains give accurate results, the test strain is also checked.

- Results and Interpretation: Organism having the oxidase enzyme will give a deep blue colour within 10 seconds, which is given by the positive control strain and organisms that lack the oxidase enzyme give a colourless result or no change, which is given by the negative control strain. Interpretation of the test strain should be done accordingly.

- Precautions

  - A few organisms such as Salmonella typhi may give a false positive result after 10 seconds and they are considered as oxidase negative. Care should be exercised in these cases, not to interpret these organisms as oxidase positive.

  - The discs should always be stored at 4°C in the refrigerator. The temperature fluctuations can affect the quality of the discs.

  - A heavy inoculum of organisms could delay the reaction beyond 10 seconds.  Hence the inoculum should be representative and at the same time shouldn’t be too heavy.

  - A nichrome wire or a stainless-steel wire loop should not be used for the test as these wires due to repeated use could produce oxidized products at the top of the loop which in turn could lead to false positive reactions.

  - Do not use the discs beyond the expiration date.

  - Use a single disc for one test organism. Using a single disc for more than one organism could lead to wrong results due to over smearing of organisms.

      - Indole Test

- Certain bacteria have the enzyme tryptophanase that is capable of breaking down the substrate tryptophan to indole, ammonia and pyruvic acid. The product indole is one of the main metabolic products of the amino acid tryotophan. This indole production is an important characteristic feature in the presumptive identification of the genera in the family Enterobacteriaceae, especially to differentiate between Esch.coli from the members of the Klebsiella and Enterobacter group. The test involves the addition of the reagent p-dimethyl benzaldehyde which reacts with the indole. The indole is formed due to the breakdown of the tryptophan by the organisms in a medium of peptone water. This indole reacts with the reagent and forms a red coloured complex or a ring called nitrosoindole. This constitutes a positive result.

- Materials Required

- Test tubes, Peptone water, 18-24 hours culture of the test organism, Control organisms (the details of the same are mentioned below), p-dimethyl amino benzaldehyde (5%) or Kovac’s reagent.

- Positive Control: Escherichia coli (ATCC 25922)

- Negative control: Klebsiella pneumoniae (ATCC 700603)

- Test Procedure

- A few drops of indole reagent are added to 18-24 hours liquid subculture in peptone water or nutrient broth tube along the walls of a colony to be identified.

- Results & Interpretation

- The development of a bright fuchsin red colour within seconds of adding the reagent at the interface between the reagent and the peptone water/nutrient broth is an indication of a positive test and the formation of indole from tryptophan as seen for the Positive Control. A colourless to yellow ring indicates absence of indole production as seen for negative control. The reagent is checked with controls whenever a new reagent bottle is opened.

- Precautions

  - Do not add the reagent directly into the tube as the reagent may sink down to the bottom of the tube and fail to give a ring at the interface of the two liquids.

  - The reagent should be stored in the refrigerator and should be shaken gently before use.

  - The bottle of Kovac’s reagent should be tightly closed at all times and never left open. Failure to comply with this requirement may lead to deterioration of the reagent as the alcohol tends to evaporate leaving a concentrated deposit at the bottom of the bottle.

  - If the positive control strain does not give the desired reaction with the reagent, the reagent should be checked, and the peptone water needs to be changed after two successive failed reactions as that batch of peptone water may not contain an adequate amount of tryptophan.

      - Citrate Utilization Test

- This test is used to determine the ability of the bacteria to utilize the citrate as the sole source of carbon and energy. Sodium citrate is a sodium salt of citric acid, a simple organic compound found as one of the metabolites in the Krebs cycle. It also tests the ability of the organisms to utilize ammonium salts as the sole source of nitrogen. The organisms which are capable of extracting nitrogen from ammonium salts with the production of ammonia leading to alkaline bye products. These products cause alkalinisation of the medium which in turn lead to a change in the colour of the medium from green to blue (as bromothymol blue is the indicator).

- Materials Required

- Simmons citrate agar tubes, Control organisms, liquid subculture of the organism or primary plate with colonies

  - Positive control: Klebsiella pneumoniae (ATCC 700603)

  - Negative control: Escherichia coli (ATCC 25922

- Test Procedure

- Pick up a single well isolated colony from the primary medium using the inoculating loop or a 2–4-hour liquid subculture and inoculate as single streak on the slant surface of the Simmons citrate agar test tube. Incubate the inoculated tube for 18-24 hours at 37°C in the incubator. Observe for growth or colour change in the medium. Repeat the same for the control organisms.

- Results & Interpretation

- Change of medium colour to blue colour or growth in the medium indicates the test is Positive, the bacteria is capable of utilising citrate as seen by Positive Control test organism. Original green colour of the medium without growth indicates the test to be negative. The reagent is checked with controls whenever a new batch of media is prepared.

- Precautions

  - The pH of the medium should be maintained at 6.9 & the test involves the production of alkaline products. This is important to avoid false positive or false negative results.

  - It is very important that the medium used for the test should be free of any source of proteins and carbohydrates

      - Urease Production Test

- This test is used to determine the ability of the bacteria to produce urease for utilization of urea. When urea is utilized, ammonia is formed during incubation which makes the medium alkaline. showing a pink – red colour by change in the phenol red indicator tested.

- Materials Required

- Urea agar medium, Control organisms, liquid subculture of the organism or primary plate with colonies.

  - Positive control: Klebsiella pneumoniae (ATCC 700603)

  - Negative control: Escherichia coli (ATCC 25922)

- Test Procedure

- Heavy inoculum of the test organism as a colony or a loopful of liquid subculture of the organism is inoculated on the surface of the slant and incubated in incubator at 37ºC for 18-24 hrs. The next day the slant is observed for the change in the colour of the medium.

- Results & Interpretation

- Medium changes to pink with production of urease by the organism and the test is considered positive, as seen with positive control. If there is no colour change it indicates the test to be Negative, as seen for negative control. The reagent is checked with controls whenever a new batch of media is prepared.

      - Coagulase Test

- Coagulase test identifies whether an organism produces the exoenzyme coagulase, which causes plasma to form a clot. Organisms that produce coagulase can form protective barriers of fibrin around themselves, making themselves highly resistant to phagocytosis, other immune responses, and some other antimicrobial agents. Staphylococcus aureus produces the enzyme coagulase which causes plasma to clot by converting soluble fibrinogen to insoluble fibrin.

- Materials Required

- Citrated plasma, peptone / nutrient broth, test organism and control organism.

  - Positive control: Staphylococcus aureus (ATCC 25923)

  - Negative control: Staphylococcus saprophyticus (ATCC BAA750)

- Slide Test: The slide coagulase test is used to identify the presence of bound coagulase or clumping factor, which is attached to the cell walls of the bacteria. Bound coagulase reacts with the fibrinogen in plasma, causing the fibrinogen to precipitate. This causes the cells to agglutinate, which creates the “lumpy” look of a positive test.

  - Procedure: Place a drop of normal saline on each end of a slide. Emulsify a colony of test organisms in each of the drops to make two thick suspensions. Add a loopful of plasma to one of the suspensions and mix gently. Look for clumps of organism within 10 seconds. No plasma is added to the second suspension as a control for the saline and also to rule out non emulsifiable colonies. Control strains are tested to check for the stability of plasma.

- Tube Test: The tube coagulase test is used to identify the presence of either bound coagulase or free coagulase, which is an extracellular enzyme. Free coagulase reacts with a component of plasma called coagulase-reacting factor to form a coagulum.

  - Procedure: Take three small tubes and label T-test broth culture, P-positive control N-negative control. Pipette 0.2 ml of plasma into each tube. Add 0.8 ml of the test broth culture, Positive Control and negative control into each tube. After mixing gently, incubate the 3 tubes at 37˚C. Examine for clotting after 2-8 hours.

- Results & Interpretation

- The results should be read as follows:

  - If clumping is seen within 10 seconds in the slide coagulase - indicates the test is Positive.

  - If there is no clumping the test is Negative for coagulase.

  - If the clot formation is observed in the tube- indicates the test to be Positive

  - If no clot formation is observed in the tube, then the test is said to be Negative.

      - Triple Sugar Ion Agar Test

- TSI is a composite medium that incorporated glucose, lactose and sucrose in the medium. The concentration of lactose is 10 times that of glucose. Organisms that follow the fermentative pathway of metabolism break down the sugars contained in the medium with the production of acid / gas. Phenol red being the indicator, the color of the medium changes to a deep yellow if there is a fermentation of the sugars present, as the mixed acids produced would cause this colour change. The amount of gas produced would determine the disintegration of the medium due to the gas bubbles. At times the entire medium could be lifted to the top of the tube. Ferrous sulphate is present as the indicator for production of hydrogen sulphide and this leads to blackening of the medium. Utilisation of proteins in the medium makes it alkaline making the medium pink in colour.

- Materials Required

- TSI agar poured in the form of butt and slants, 4-6 hours liquid broth culture of test organism.

  - Escherichia coli (ATCC25922) - Acid Slant / Acid butt

  - Pseudomonas aeruginosa (ATCC 27853) - Alkaline slant / Alkaline butt

  - Proteus mirabilis - Alkaline slant and acid butt with H2S between the two portions of medium.

- Test Procedure

- Inoculate the test broth, the charged wire as a stab inoculum into the butt of the TSI tube. Without touching the bottom of the tube, withdraw the straight wire and streak the slant portion of the tube with the wire or with a sterile loop. Incubate aerobically for 24 hours at 37°C. Repeat the same for the control organisms.

- Results & Interpretation

  - Acid slant and acid butt (Yellow slant with a yellow butt): When the slant and the butt both are yellow in color then glucose, lactose/sucrose is fermented. This is indicative of lactose fermenting coli forms like Escherichia coli, Klebsiella and Enterobacter.

  - Alkaline Slant and Acid butt (red slant with a yellow butt): The organism is capable of utilizing only glucose, but not sucrose or lactose. Indicative of non-lactose fermenting genus like Shigella species and Salmonella Paratyphi A.

  - Alkaline slant and alkaline butt: (red slant with a red butt) When the slant and the butt both are red in color then there is no utilization of any sugar in the medium. This could be indicative of a non-fermenter such as Acinetobacter or Pseudomonas species. When there is black coloration of the medium then there is production of the H2S in the medium by the organism. The presence of H2S may obscure other details of fermentation, acid production and gas production. However, a close inspection of both the slant and the butt portions of the tube would lead to an accurate interpretation of the results.

- Precautions

  - The slant and the butt portions of the medium should be equal and approximately 3cms and this ratio should be maintained so that the slab inoculum of the test organism is uniform throughout the medium.

  - A straight wire should be used instead of a loop as using the latter could produce a huge dent in the butt portion of the medium, causing a distortion of the medium. This could adversely affect the interpretation of gas production.

  - Care should be taken to ensure that the wire does not touch the bottom of the tube as that would also create an artificial vent in the medium, leading to problems in the interpretation of gas production.

- Quality Control Procedure

QC of culture media is checked with American Type Culture Collection – Microbial strains.

Sample collection containers are checked for their sterility and sample stability as per lot change.

- Interfering Substances

- Sample collection process and transportation should be done appropriate to the sample to avoid sample contamination with the commensal flora and environmental contaminants. Adequate aseptic precautions to be taken to avoid contamination during culturing process.

- Result Calculation

- Refer to Section 13.

- Biological Reference Interval

- No growth is observed for all samples from sterile body sites unless there is an infection. Commensal growth may be observed in sputum, BAL, ET/ tracheal secretions, semen, stool, urine, specimens from – oral cavity, throat swab, eye, ear, vagina, endocervix, stool, open wounds.

- Reportable Interval

- 72 Hours

- Procedure when Results are not in Range

- When the results are found inappropriate to the sample type, the sample is re-stained or checked with a repeat sample as per individual sample and growth discretion.

- Critical Alert Values

- All significant growth observed in samples from sterile body sites is informed as critical alert.

- Laboratory Clinical Interpretation

- Specific clinical interpretation is given as per sample type and pathogen isolated.

- Potential Source of Variability

The following factors play a crucial role in the result variations and appropriate care should be taken to eliminate them:

      - prior antibiotic usage by patient

      - sample type

      - sample collection

      - sample transport temperature

      - time to sample collection and inoculation

      - portion of sample used for inoculation

      - culture media

      - loop size and sterility

      - temperature and time of incubation.

- References

Practice of Medical Microbiology – Cruickshank