SOP - 008

PROCEDURE FOR

AUTOMATED IDENTIFICATION & ANTIMICROBIAL SUSCEPTIBILITY TESTING

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	12

15.	Interfering Substances	12

16.	Result Calculation	12

17.	Biological Reference Interval	12

18.	Reportable Interval	13

19.	Procedure when Results are not in Range	13

20.	Critical Alert Values	13

21.	Laboratory Clinical Interpretation	13

22.	Potential Source of Variability	13

23.	References	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ATCC | American Type Culture Collection |
| 2 | AST | Antimicrobial Susceptibility Testing |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | IAMM | Indian Association of Medical Microbiologists |
| 6 | ID | Identification |
| 7 | MI | Microbiology |
| 8 | MIC | Minimum Inhibitory Concentration |
| 9 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 10 | rpm | Revolutions per minute |
| 11 | SOP | Standard Operating Procedure |
| 12 | TDPL | TRUSTlab Diagnostics Private Limited |
| 13 | V2C | VITEK 2 Compact |

- Objective

  - This SOP describes the procedure for the preparation and identification and antimicrobial susceptibility testing using the VITEK 2 Compact instrument.

- Scope

This procedure applies to all staff of the Microbiology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Identification of a pathogen and antimicrobial susceptibility testing are crucial steps in treating infectious diseases. Identification determines the type of microbe, while susceptibility testing determines which antibiotics are effective against it, guiding appropriate treatment.

- Test Principle

The VITEK 2 Compact system uses a fluorogenic methodology for organism identification and a turbidimetric method for susceptibility testing using a 64 well card that is barcoded with information on card type, expiration date, lot number and unique card identification number.

- Specific Performance Characteristics

- Antibiotic sensitivity is based on broth micro dilution method as the instrument can read turbidity after bacteria are exposed to predefined dilutions of antibiotics at a defined time period. The MIC of antibiotics is based on break point studies and are defined as per available guidelines on the subject. The instrument is capable of identifying bacteria (aerobic and/or facultative anaerobes) based on fluorogenic detection and biochemical utilization up to species level. Specialized software is used for analyzing the data and reporting.

- Type of Sample

- Pure isolates of organisms to be tested may be taken from chocolate agar, MacConkey agar and sheep blood agar.

- Patient Preparation

- N/A.

- Type of Container & Additives

- N/A.

- Equipment & Reagents

Sterile inoculating loops or sterile swabs

VITEK 2 Compact Identification and AST cards

75 mm × 12 mm clear polystyrene tubes (single use only)

DensiCHEK Plus Meter with McFarland Standards for calibration (0.0, 0.5, 3.0, and 5.0 McF)

Sterile suspension solution

Bar-coded 10 well cassette card holders

VITEK® 2 compact system for identification and antibiotic sensitivity test

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

Suitable infection prevention precautions to be followed.

Biohazardous spills can occur inside the V2C instrument. All organism suspensions, cards, cassettes, test tubes, sample transfer tubes, waste bin and the user interface panel should be considered as potentially infectious. Use disinfectant to clean any spill that occurs in the V2C instrument

Use non-powdered gloves when handling the cassette with live organisms.

Discard used reagents and samples as per standard disposal procedure.

- Calibration Procedure

- Biosafety Cabinet and incubator calibrations are done on an annual basis by an NABL accredited calibration agency. Necessary maintenance measures are undertaken for all the equipment.

- The V2C instrument will be serviced annually under a preventative maintenance agreement.

- DensiCHEK Plus is calibrated using the calibration standards on daily basis when in use.

- Test Procedure

Precautions

- When handling specimens or cultures, aseptic technique is important to avoid their contamination and to protect the worker from infection from the sample.

- The following points should be observed when culturing specimens or performing subcultures:

      - Caps and lids from culture media should not be placed on the workbench, but retained in the hand or placed in such a way while the sample is being processed, taking care not to contaminate the hand or cap. Caps and lids should be replaced as soon as possible.

      - If the work is being carried out on the open bench, a Bunsen burner/loop sterilizer should be in close proximity to sterilize loops or wires.

      - Keep samples away from the face when opening culture containers.

      - Aerosol production should be minimized by:

        - Opening caps slowly as the contents of containers are sometimes under pressure.

        - Avoiding vigorous swirling or shaking of the sample prior to opening.

        - Cooling loops that have been heated before use.

        - Avoiding expelling the last drop from a pipette.

        - If forceps or scissors are used when handling specimens, they should be heated in a Bunsen flame/loop sterilizer and allowed to cool before use.

        - Loops and wires should be heated in the Bunsen flame/loop sterilizer with the loop end downwards until red-hot, then removed and allowed to cool. Care should be taken that the loop does not contain fluid or large particles of matter that may ‘splatter’ when placed in the flame, and that the entire loop\wire up to the loop holder is heated to red-hot state.

Procedure

Bench Work

      - Take 2 polystyrene tubes for one isolate and place it on a cassette.

      - Take 3 ml suspension solution in each tube from the dispensette.

      - Mark Tube-1 as ID tube (for identification) and mark Tube-2 as AST tube (for sensitivity).

      - Pick similar isolated colonies from quadrant 4 from the plate and make a suspension in Tube-1.

      - Fix the turbidity of Tube-1 up to Bacteria (Gram negative and Gram positive): 0.50-0.63.

    - To use the DensiCHEK Plus Meter to read samples

      - Ensure the instrument is ON and set to the PLASTIC tube.

      - Blank the DensiCHEK Plus by filling a test tube with sterile saline and inserting the tube into the instrument. Press the “0” key and slowly rotate the test tube.

      - Ensure one full rotation is completed before the reading is displayed.

      - The instrument will display a series of dashes followed by “0.00”.

      - To measure a sample, place a well-mixed organism suspension into the instrument and slowly rotate the test tube.

      - Ensure one full rotation has completed before the reading is displayed. The instrument will display a series of dashes followed by a reading.

      - Remove the test tube after completion of a reading.

      - The instrument will automatically shut off when test tubes are not  inserted after one minute.

      - Transfer 145 µl for Gram negative bacilli and 280 µl for Gram positive cocci from Tube-1 to Tube-2.

      - Place the ID card in Tube-1 and the AST card in Tube-2.

    - Machine Work

      - Place cassette in filler section.

      - Press “Start fill” button.

      - The V2C instrument will beep when the filling cycle is complete.

      - When the cards are finished filling, the Load Door is automatically unlocked.

      - Place the cassette in the Load Door.

      - When the cards are loaded, remove the cassette and dispose of the tubes and straws in a biohazard container.

      - When the cards are processed and results are obtained, cards will be automatically ejected into the waste collection bin.

    - Software Work

      - Click “Cassette view”.

      - Click on the “Red” Cassette entry on left.

      - Click on the accession of ID card and drag till accession of AST card.

      - Click “Define isolate” icon.

      - Enter the lab ID.

      - Click “OK”.

      - Enter Bench name = Sample type.

      - Click “Save cassette information” icon.

      - Click “Patient view”.

      - Click “Add patient”.

      - Enter: Patient ID = Lab ID, Name = Patient Name, Location = Patient Location, Physician = Dr. Name, Lab ID = Lab ID, Type = Bench Name, Source = Sample type.

      - Click “OK”.

- Quality Control Procedure

Each new lot number of ID and AST cards is tested with stock culture organisms to ensure proper performance.

Susceptibility cards are tested once against stock culture organisms to ensure proper performance

QC organisms to be used

      - ID-GP cards - Enterococcus casseliflavus ATCC 700327 / Staphylococcus saprophyticus ATCC BAA-750

      - ID-GN cards - Enterobacter hormaechei ATCC 700323

      - AST N235/N405/ N406 - Escherichia coli ATCC 25922 / Pseudomonas aeruginosa ATCC 27853

      - AST GP628 - Enterococcus faecalis ATCC 29212 / Staphylococcus aureus ATCC 29213.

    - EQAS - IAMM EQAS three times a year.

- Interfering Substances

- Adequate aseptic precautions are to be taken to avoid contamination during the process. Very mucoid organisms may not provide acceptable results. Alternative methods should be used for these organisms. Improper sub culturing and filling of VITEK cards may result in inconsistent or erroneous bio patterns. Unsterile normal saline can cause erroneous results.

- Result Calculation

Results are concurrently printed and the data sent to the Results View folder on the left side of the screen also called the Navigation Tree where the information is archived. A red cassette in the Navigation Tree is indicative of an error. If an error occurs during processing, refer to the Software User Manual (see section 15).

Review results printout and file appropriately.

- Biological Reference Interval

- N/A.

- Reportable Interval

- 24 Hours

- Procedure when Results are not in Range

- When the results are found inappropriate results are rechecked.

- Critical Alert Values

- All significant growth observed in samples from sterile body sites is informed as critical alert.

- Laboratory Clinical Interpretation

- Specific clinical interpretation is given as per sample type and pathogen isolated and its antimicrobial susceptibility testing.

- Potential Source of Variability

- The following factors play a crucial role in the result variations and appropriate care should be taken to eliminate them:

pure colony culture

reagent storage and continual performance

equipment maintenance

- References

- Equipment and reagent usage instructions