SOP - 009

PROCEDURE FOR QUALITY ASSURANCE IN MICROBIOLOGY & SEROLOGY

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Input	8

6.	Output	8

7.	Monitoring & Measurement	8

8.	Quality Control Procedure	8

9.	References	20

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ATCC | American Type Culture Collection |
| 2 | BAL | Bronchoalveolar Lavage |
| 3 | ET | Endotracheal |
| 4 | MI | Microbiology |
| 5 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 6 | SOP | Standard Operating Procedure |
| 7 | TRUSTlab | TRUSTlab Diagnostics Private Limited |
| 8 | TSI | Triple Sugar Ion |

- Objective

- The purpose of this SOP is to explain various quality control processes carried out by the technicians of the Microbiology and Serology departments (Daily & Periodically) to monitor and evaluate testing processes to ensure accurate test results and maintenance of records/registers pertaining to quality control activities.

- Scope

This procedure applies to all staff of the Microbiology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure:

      - QM ensures that all the technicians of the departments are knowledgeable above the contents of department manual & procedures including changes relevant to the scope of their testing activities.

      - QM ensures that only the latest approved procedure manuals of the department are available to technicians at all times

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definitions

Quality Assurance: It has been defined by WHO as, the total process where the quality of laboratory reports can be guaranteed. It has been summarized as the “Right result, at the right time, on the right specimen, from the right patient, with the result interpretation based on relevant reference data, and at a right price”.

Quality Control:	Quality control covers that part of Quality Assurance, which primarily concerns the control of errors in the performance of tests and verification of the test results. Quality control covers all the aspects of every procedure done within the department. Quality Control must be practical, achievable and affordable

- Input

      - Samples for Internal quality control

      - Samples for Inter Laboratory Comparison

      - Samples for proficiency testing (PT)

      - QC Samples for EQAS

      - Test equipment, Kits, Standard quality control material

      - ATCC Strains of bacteria

      - Culture Media

      - Equipment, consumables, Bunsen flame, Bio Safety Cabinet, Inoculating Loop, Microscope, Laminar Airflow

- Output

      - Results of IQC Samples

      - Results of ILC Samples

      - Results of EQAS & PT Samples

- Monitoring & Measurement

      - Forms and Records

      - Standard deviation, coefficient of variation (CV), on Internal QC results.

      - Deviations of IQC, I.L.C, P.T, EQAS results and their corrective actions.

- Quality Control Procedure

General

    - TRUSTlab has designed internal quality control (IQC) systems and procedures that verify the attainment of the intended quality of result.

    - TRUSTlab follows the criteria for ensuring quality of examination procedures mentioned in clause 5.6 of the NABL manual 112: “Specific criteria for accreditation of medical laboratories”.

    - The IQC procedure performed in individual departments include:

      - Running of IQC samples (Controls Daily)

      - Running of Calibrators (Periodically & as and when required)

      - IQC Samples are run daily at the beginning of the day work before processing the patient samples.

      - The results of the IQC samples are approved by the Quality Manager first and then the patient samples are run.

      - The control ATCC strains are processed periodically (weekly).

      - Media that is commercially procured as well as in-house prepared media are checked for sterility.

      - The stains (Gram’s Strain and Ziehl Neelsen Stain) are checked for their sterility and quality control on a daily basis with ATCC controlled strains.

    - Technical specifications for glassware, chemicals, kits, reagents, media and antibiotic discs used for testing are determined and followed while placing purchase order to procure these items.

    - These specifications are used for conducting inspections of items delivered by the suppliers.

    - The chemical, reagents & kits, media and antibiotic discs delivered by the supplier are stored at the temperature and humidity conditions specified in the product insert or as specified by the manufacturer.

    - The expiry date of the reagents, kits, dehydrated media are checked.

    - Only approved and standard glassware of ‘A’ quality, analytical grade chemicals, approved kits and reagents, commercially available standardized media and antibiotic discs are used.

    - When a lyophilized chemical/reagent/ATCC culture ampoule/EQAS culture ampoule is reconstituted, the date of reconstitution is marked on the vial.

    - For reagents/media prepared in house, the name & concentration of the reagent, date of preparation, validated period are indicated on the bottle.

    - For commercially available reagents/kits/media the date of opening is clearly indicated on the pack.

    - Antibiotic discs are verified before use. Expired reagents, kits, media and antibiotic discs are discarded.

    - Only deionized water supplied along with kit or reagent or deionized water from the micropore water plant at the laboratory is used for re-constitution of reagents, controls, ATCC culture ampoules, EQAS sample and media preparation.

    - Equipment, Instruments (Major & Minor), Micropipettes, Temperature Measuring Devices are calibrated through external agency at intervals specified in the equipment calibration plan and the details are recorded in the equipment calibration register maintained in the department.

Serology

    - Positive and Negative controls of assayed QC samples are run once a day for all the serological parameters (HIV, HBsAg, Anti HCV, TORCH, DENGUE IgG, IgM, NS1Ag, H.pylori, IgG, IgM, etc. by automation/ according to the method performed.

    - Positive and Negative controls are run for parameters like WIDAL, VDRL, before performing the test.

    - For ELISA, Positive & Negative controls are run along with the patient samples as and when required.

    - The results of the IQC assay is documented in the IQC Register verified and approved by the HOD/Consultant Microbiologist

    - The IQC results of the parameters like are recorded in the work register along with the patient sample results.

    - External Quality Assurance (EQAS):

      - TRUSTlab participates in EQAS programmes provided by an NABL accredited PT service provider.

      - IAMM Micro EQAS from CMC Vellore (every quarter the samples are received and processed and documented).

      - RML Vellore (every quarter the samples are received and processed and documented).

Microbiology

    - IQC procedure is done for media, biochemical test, antibiotic discs, stains and test like Catalase, Coagulase, Bacitracin Discs, Optochin Discs and Oxidase Discs.

    - All commercially available and in house prepared media are examined visually for colour change, precipitate formation, lysis of blood and contamination.

    - Each Petri Plate is checked for the depth of the media poured. Any atypical observation is brought to the notice of the Microbiologist.

    - IQC is done for:

      - Sheep blood agar

      - Mac Conkey agar

      - Mueller Hinton agar

      - T.S.I Slant (Triple Sugar Iron agar slant)

      - Christensen’s Urea slant

      - Simmon’s Citrate agar slant

      - Tryptone Soy agar

      - Peptone water

      - Nutrient broth

    - On receiving these media, a plate/slant/test tube with the media from each lot is set aside for performing:

      - Sterility Check: A single piece of media from each lot of the above-mentioned uninoculated media is placed for overnight incubation (18-24 Hours) and checked for any growth. Absence of any growth in the uninoculated media specifies its sterility. The result is recorded in the sterility check register.

      - Performance Testing: This test is done to check if the media permits the growth of standard controlled strains. Standard ATCC (American Type Control Center) strains are procured. An inoculum of a standard strain is prepared in comparison to 0.5 McFarland’s turbidity standard. The inoculum is inoculated onto a set of un-inoculated media plates, slants and liquid media. Incubated overnight and checked for growth the following day.

| S No. | Media | Control Organism | Expected reaction |
| --- | --- | --- | --- |
| 1. | Blood agar | Staphylococcus aureus ATCC 25923 | Good growth, B-haemolytic colonies |
|  |  | Streptococcus pneumonia ATCC49619 | Good growth, A-haemolytic colonies |
| 2. | Chocolate agar | Haemophilus influenza ATCC19418 | Good growth |
|  |  | Neisseria gonorrhea | Good growth |
| 3. | Mac Conkey agar | Positive QC Escherichia coli CODE: 25922 | Pink Colonies (Lactose Positive) |
|  |  | NEGATIVE QC Pseudomonas Aeruginosa CODE 27853 | Colourless colonies, no spreading |
| 4 | C.L.E.D agar | Positive QC Escherichia coli ATCC: 25922 | Yellow colonies (lactose Positive) |
|  |  | Negative QC Proteus mirabilis ATCC 25933 | Bluish Green colour colonies, no spreading |
| 5 | Xylose-Lysine Dextrose | Salmonella species | Red colonies (Positive Lysine) |
|  |  | Escherichia coli ATCC: 25922 | Yellow colonies (Positive Sugars) |
|  |  | Shigella species | Transparent colonies (Negative) |
| 6 | Mueller Hinton agar | Escherichia coli ATCC: 25922 | Good growth |
|  |  | Staphylococcus aureus ATCC 25923 | Good growth |
|  |  | Pseudomonas Aeruginosa ATCC 27853 | Good growth |

QC – VITEK 2 COMPACT

    - The Vitek 2 Compact system uses a fluorogenic methodology for organism identification and a turbidimetric method for susceptibility testing using a 64 well card that is barcoded with information on card type, expiration date, lot number and unique card identification number:

- Antibiotic sensitivity is based on broth micro dilution method whereas the instrument can read turbidity after predefined concentration of bacteria are exposed to predefined dilutions of antibiotics. The MIC of antibiotics is based on break point studies and are defined as per available guidelines on the subject. The instrument is capable of identifying bacteria (aerobic and/or facultative anaerobes) based on fluorogenic detection and biochemical utilization up to species level.  Specialized software is used for analyzing the data and reporting. The current version of software in use is 8.1 as per manufacturer’s manual.

- The instrument is capable of reporting MICs in terms of mg/ml and interpreting it as Resistant, Sensitive or Intermediate.

    - IQC

      - Each new lot number of ID cards is tested with stock culture organisms on a weekly basis to    ensure proper performance.

      - Susceptibility cards are tested weekly against stock culture organisms to ensure proper performance.

      - All shipments of new lots for ID and sensitivity cards must be entered in the VITEK 2 C prior to use.

    - QC Organisms to be Used

      - Enterococcus faecalis ATCC 29212

      - Staph aureus ATCC 29213

      - E. coli ATCC 25922

      - PSA ATCC 27853

      - E. coli ATCC 35218

    - With New Lots

      - ID-GP cards (Streamlined QC when applicable)

      - Enterococcus casseliflavus ATCC 700327

      - Staph saprophyticus ATCC BAA-750

      - ID-GN cards (Streamlined QC when applicable)

      - Enterobacter hormaechei (E.cloacae) ATCC 700323

      - Stenotrophomonasmaltophilia ATCC 17666

IQC for Stains

      - IQC for stains is performed daily and whenever a new lot of stain kit is used.

      - The stain kit bottles are checked for signs of deterioration, such as abnormal turbidity, formation of a precipitate, decolorization.

      - In case of formation of a precipitate, the stain is filtered and then used.

      - In case of   abnormal turbidity/ decolorization the stain kit is discarded and intimated to the microbiologist.

      - IQC for GRAM’S STAIN: Controls smear for Gram’s stain is prepared from a mixed culture of Staphylococcus aureus (ATCC 25923) and Escherechia coli (ATCC 25922). These control smears are run along with the day to day patients samples.

      - IQC for Gram’s stain is run daily once using standard ATCC strains.

      - IQC for Ziehl Neelsen Stain for Acid fast bacilli: 2 Sputum smears of known High & Low AFB Positivity are prepared from Mycobacterium tuberculosis strain. These smears are stained daily with routine patient smears to check for quality control.

IQC of various Biochemical’s Procedures

      - IQC for biochemical reactions is done every week. With each new batch of the house prepared media tubes, certain tubes from each lot of media are inoculated with standard strains and tested for the reactions.

| S No. | Procedure Test | Control Organism | Expected Result | Expected Reaction |
| --- | --- | --- | --- | --- |
| 1. | Catalase | Staphylococcus aureus (ATCC strain- 25923) Enterococcus. faecalis (ATCC strain -29212) | + - | Bubbling reaction No bubbling |
| 2. | Coagulase | Staphylococcus aureus (ATCC strain- 25923) Staphylococcus epidemidis | + - | Clot formation in 4 hours No Clot |
| 3. | Indole | Escherichia coli (ATCC 25922) Klebsiella pneumonia (ATCC700603) | + - | Red ring at surface Yellow ring at surface |
| 4. | Methylred | Escherichia coli (ATCC 25922) Klebsiella pneumonia (ATCC700603) | + - | Instant red colour No colour change |
| 5. | Bacitracin disc | Streptococcus Group A Enterobacter faecalis | + - | Zone of inhibition>12mm No zone of inhibition |
| 6. | Optochin disc | Streptococcus pneumonia Streptococcus viridians | + - | Zone of inhibition>-12mm No zone of inhibition |
| 7. | Oxidase disc | Pseudomonas aeruginosa (ATCC 27853) Esherichia coli (ATCC25922) | + - | Purple colour in 10 seconds. No Change in colour |
| 8 | Citrate utilization test | Escherichia coli (ATCC 25922) Klebsiella pneumonia (ATCC700603) | - + | No colour change Blue colour formed |
| 9 | Urease test | Escherichia coli (ATCC 25922) Klebsiella pneumonia (ATCC700603) | - + | No colour change Bright Pink colour formed |
| 10 | T.S.I(Tripple Sugar Iron Test) | Escherichia coli (ATCC25922) Pseudomonas aeruginosa (ATCC 27853) Proteus mirabilis (ATCC25933 ) |  | Acid Slant / Acid butt Alkaline slant / Alkaline butt Alkaline slant and acid butt with H2S between the two portions of medium. |

      - The results of the above mentioned biochemical reactions are documentated in the IQC register for Biochemical reactions.

IQC for Bacitracin, Optochin and Oxidase Discs

      - Blood agar plates are streaked with the inoculum of standard strains and Optochin and Bacitracin discs are placed in the center of the plates respectively.

      - IQC is checked for a new batch of Bacitracin   and Optochin discs.

      - The results are documented in IQC register for ATCC strains.

      - Commercially available oxidase discs are checked for IQC on weekly basis.

      - The discs are checked for colour change. If any change in colour is observed the discs are discarded.

| S No. | Procedure Test | Control Organism | Expected Result | Expected Reaction |
| --- | --- | --- | --- | --- |
| 1. | Bacitracin disc | Streptococcus Group A Enterobacter faecalis | + - | Zone of inhibition>12mm No zone of inhibition |
| 2. | Optochin disc | Streptococcus pneumonia Streptococcus viridians | + - | Zone of inhibition>-12mm No zone of inhibition |
| 3. | Oxidase disc | Pseudomonas aeruginosa Escherichia coli | + - | Purple colour in 7-10 seconds No Change in colour |

IQC of Antibiotic Discs

      - Quality Assurance is essential to ensure the quality of AST (Anti Biotic Sensitivity Testing) by Kirby Bauer discs diffusion method.

      - Routine IQC listing with a range of commercially procured ATCC (American type control cultures) strains is a major part of the Quality Assurance Process as it facilitates monitoring of the performance of test and also check for the various resistance patterns exhibited by the organisms thus helping to monitor the antibiotic formula prescribed to the patient.

      - Most standardized methods include tables of acceptance of zone size ranges for control stains and in addition to checking zone diameters, are written within the publicized ranges. Rules or statistical approaches are applied to indicate deviation from acceptable performance.

      - IQC for Antibiotic Discs is done weekly once and also when a new batch of antibiotic is received.

      - Kirby Bauer discs diffusion method is done to test the antibiotic sensitivity test.

      - Control Strains are inoculated on Mueller Hinton agar, the plate is allowed to dry for a few minutes, and the antibiotic discs are placed on the plate. Around 6 antibiotics are placed in a plate of 90 mm and 9 antibiotics are placed in plate of 110mm.at a distance of 20 mm between each disc.

      - The plates are incubated over night with plates inverted at 37’C.

      - After the incubation of 16-18 Hours, the zone diameters of the antibiotic discs are measured with the help of the calipers or zone measuring scale and the zone diameters are interpreted as sensitive, intermediate and resistant.

      - When control tests indicate unacceptable performance, the source of the error is investigated viz: - Depth of the Media poured, number of the antimicrobial discs applied to each plate, Concentration of the inoculum.

      - A single control organism antimicrobial combination out of range will not necessarily require action unless the result is far out of range or there is other indication of a problem (E.g.: small control zone together with a frequency of resistant isolates than normal).

      - 2 successive results either above or below the control limits (indicate a likely systemic error) resulting in a problem in the test.

      - More than 2 results outside the limits in any direction, 2 2S in a series of 20 tests may indicate error.

      - If zone diameters are consistently above or below the Mean of the range, a systematic error may be indicated particularly if this is true for several agents. In this case the methodology should be checked including the depth of the agar time of inoculation, time between inoculation and application of discs, time between application of the discs & incubation, inoculums concentration and incubation conditions.

Quality Control Material

      - TRUSTlab adopts mutual consent standards or methods which are clearly established in peer reviewed journals and textbooks, specified, characterized and mutually agreed upon by all parties concerned.

      - In TRUSTlab Quality Control Materials are periodically examined with a frequency that is based on the stability of the product and the risk of harm to the patient from an erroneous result.

      - TRUSTlab documents statements regarding reagents, procedures or the examination system where traceability is provided by the suppliers or manufacturer like the product insert, operating manuals etc.

Quality Control Data

      - Internal Quality Control Procedure performed in the respective Departments is documented in the department manuals.

      - QM/HOD/Senior Technician ensures that the IQC & EQAS samples are introduced along with the patient samples in the routine day-to-day laboratory workload and processed in the same manner as patient specimens by the technicians who routinely perform patient sample testing.

      - The QC values are checked for the acceptable range specified in the product insert.

      - If the QC is not within the acceptable range specified, the reagents are checked and changed if necessary/IQC sample is also checked. The instrument is recalibrated, and QC is rerun to obtain a value that is close to the target values and within the specified range.

      - QM/HOD/Senior Technician analyses the results of any undue deviation and takes necessary corrective and preventive action if any and records the same in corrective action record.

      - The corrective actions taken are evaluated and validated by the QM and these corrective actions are documented in the Corrective Action record.

      - All data of the daily QC values are compiled, and a monthly QC report is prepared. The coefficient of variation (CV), Standard Deviation (SD), Mean of the values are calculated from the monthly QC data.

      - QC values are accepted if they are within the target range, the standard deviations, coefficient of variation obey the QC rules mentioned below.

      - Westgard rules are followed for the acceptance/non-acceptance of the QC and SD values.

        - QC is rejected if it is outside 3SD (1 3S rule).

        - Two consecutive values obtained are outside 2SD on the same side of the mean but within 3SD (2 2S rule).

        - Ten consecutive values are above or below the mean, but within 2SD (10x rule).

      - Rules followed when 2 levels of QC materials are used:

        - QC is rejected if either QC values is outside 3SD (1 3S),

        - Both QC values are outside 2SD on the same side of the mean but within 3SD (2 2S).

        - The difference between the QC values is >4SD i.e., one level QC is >2SD and the other level QC is < 2SD (R4S).

        - Ten consecutive of the same level QC   are >/< the mean, but within 2SD (10X) rule.

        - Five consecutive values of one level QC and five consecutive values of other level QC are >/< the mean but within 2SD (10X).

      - In serology I 3S rule, 2 2S rules and 4S rule hold good to the QC values.

      - QC is rejected when the QC values are not within range as mentioned in the Kit Insult.

      - In Microbiology controlled ATCC Strains are used, the results are correlated with the values of the controlled strains as per CLSI (clinical laboratory standard institute) guidelines. Any deviation in the values, the QC is re-done till the value is close to the target within the specified range.

QC OUTLIERS AND ROOT CAUSE ANALYSIS CHECKLIST

      - When QC is out of range, a root cause analysis is done according to the checklist given below

        - Checking the expiry date of the reagents and the quality control material.

        - Changing to a new reagent from the same lot.

        - Changing to a new lot of reagents.

        - Changing quality control material.

        - Checking the calibration status of the analyte.

        - Recalibration for the analyte in the instrument using the calibrator. The details of out-of-range QC values and the corrective and preventive actions taken are approved by the QM/HOD. Patient samples are processed only after the approval of the QC ranges and L.J. charts.

      - Whenever internal controls and calibrations cannot be performed for any analyte or parameter or test procedure, TRUSTlab ensures accuracy and reliability of test results in the following ways:

        - Inter lab comparison

        - Split sample testing

        - Following validated test procedures

        - Following procedures mentioned in peer reviewed journals, reputed and standard textbooks.

Inter Laboratory Comparisons

      - TRUSTlab participates in an Inter lab comparison programme(s) as per the NABL112, “Specific criteria for accreditation of medical laboratories” appropriate to the examination and interpretation of the examination results.

      - TRUSTlab participates in inter-lab comparison programme with laboratories which are NABL accredited, for parameters in the scope of testing which are not covered under EQAS (External quality assurance scheme)

      - TRUSTlab, chooses its referral labs with the same methodology of testing samples.

      - TRUSTlab, follows a specific procedure for choosing, monitoring and evaluating its referral laboratories. A criterion is set for choosing, monitoring and evaluating referral laboratory, which is included in the Evaluating Referral Lab and consultant evaluation form.

      - A memorandum of understanding (MoU) is maintained for each referral lab.

      - SCOPE of the referral lab is also included in the ILC record.

- References

NABL112, “Specific criteria for accreditation of medical laboratories” appropriate to the examination and interpretation of the examination results.

CLSI Guidelines-2021