SOP - 006

PROCEDURE FOR

RAPID PLASMA REAGIN (RPR) TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	10

15.	Interfering Substances	10

16.	Result Calculation	11

17.	Biological Reference Interval	12

18.	Reportable Interval	12

19.	Procedure when Results are not in Range	12

20.	Critical Alert Values	12

21.	Laboratory Clinical Interpretation	12

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BFP | Biological False Positive |
| 2 | CDC | Centre for Disease Control |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | FTA-ABS | Fluorescent treponemal antibody absorption |
| 6 | IgG | Immunoglobulin G |
| 7 | IgM | Immunoglobulin M |
| 8 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 9 | NS1 | Non-Structural protein 1 |
|  | rpm | Revolutions per minute |
| 10 | RPR | Rapid Plasma Reagin |
| 11 | SE | Serology |
| 12 | SOP | Standard Operating Procedure |
| 13 | TDPL | TRUSTlab Diagnostics Private Limited |
| 14 | TPHA | Treponema pallidum hemagglutination assay |
| 15 | VDRL | Venereal Disease Research Laboratory |

- Objective

  - This procedure describes the step-by-step tasks involved in the process of the Rapid Plasma Reagin (RPR) test.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

The Rapid Plasma Reagin (RPR) test is an in vitro qualitative and semi quantitative test to diagnose Syphilis using human serum/plasma for detection of non-specific antibodies to Treponema pallidum.

- Test Principle

The RPR reagent contains modified VDRL Antigen coated with microparticulate             carbon particles which flocculate when mixed with sample containing Reagin.

A reactive specimen is indicated macroscopically by visible black clumps against the white background, in contrast a non-reactive specimen appears to have a smooth uniform light gray colour.

This test is a screening test and if positive, further testing with specific confirmatory tests is required.

- Specific Performance Characteristics

N/A.

- Type of Sample

- Serum/Plasma or Whole Blood must be used for testing. Should a delay in testing occur, store the sample at 2-8°C. Samples can be stored for up to a week. Do not use hemolysed, icteric, lipemic serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

RPR Kit – Reagent, Control serum, Disposable plastic cards, Applicator sticks

Plastic droppers

Centrifuge

Normal Saline

- Safety

The use of disposable gloves and proper biohazardous clothing is strongly recommended while running the test.

In case there is a cut or wound in hand, do not perform the test.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in-vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

Protein-A Conjugate and Buffer Solution contain Sodium Azide as a preservative. If these materials are to be disposed of in a sink or other common plumbing systems, ensure to flush with generous amounts of water to prevent accumulation of potentially explosive compounds. In addition, consult the manual guideline "Safety Management No. CDC-22", Decontamination of Laboratory Sink Drains to remove Azide salts" (Centre for Disease Control, Atlanta, Georgia, April 30, 1976.)

- Calibration Procedure

- Centrifuge calibration is done on an annual basis by an NABL accredited calibration agency.

- Test Procedure

Precautions

      - For every new kit, check the intactness of the reagent bottle, do not use the reagents if they are opened or leaked.

      - Do not use the reagents after expiry date.

      - Do not use the reagents from different lots of kits.

      - Do not use the reagents, if they appear turbid or discoloured.

      - Bring all the reagents and samples to room temperature before starting the procedure. If the temperature of reagents is less than 23°C test reactivity decreases and at temperatures greater than 29°C the test reactivity increases.

      - Avoid the contact of test circles as the oily deposits formed thereof may invalidate the test results. Avoid touching the dropper tip to the reagent.

      - If the moistened humidifying cover is not used while rotating, proper humidity will not be maintained, and the test component may dry on the card giving rise to false reactive results.

      - Too fast or too slow speeds of the mechanical rotator may produce unpredictable test results due to improper antigen antibody interaction.

      - Rotation of the reaction mixture for longer than recommended time increases the test reactivity whereas too short a rotation period decreases the test reactivity.

      - Tilt the card gently by hand after removal from the rotator. Vigorous movements can hamper flocculation.

      - Read the test results immediately (after card rotation) under a high intensity lamp or strong daylight. While interpreting the results, if the light produces glare on the card may obscure the reaction.

Procedure

    - Qualitative Test

      - Place one drop (50 µl) of positive, negative control serum and test samples on the card in separate well. Add a drop of reagin reagent and place the card on the VDRL rotator and rotate for 8 minutes at 180 rpm. Read the result immediately after the rotation. Positive control serum reaction should show black needle shape floccules in the reaction mixture, while the negative control serum should show no clumps.

    - Semi-Quantitative Test

      - Place one drop (50 µl) of 0.9 % saline in 2nd, 3rd, 4th and 5th circles of the card using a micropipette. Do not spread the saline solution.

      - Using a micro-pipette, add 0.05 ml (50 µl) of the sample to the 1st and 2nd circle.

      - Using the same micro-pipette, mix the sample with saline by aspirating back and forth several times. Aspirate (50 µl) from the 2nd circle and transfer it to the 3rd circle. Repeat the same operation successively up to 5th circle. Aspirate (50 µl) from 5th circle and discard. This creates a titre series of 1:2, 1:4, 1:8, 1:16, 1:32 starting from the 1st circle to the 5th circle.

      - Add a drop of reagin reagent and place the card on the VDRL rotator and rotate for 8 minutes at 180 rpm. Read the result immediately after the rotation.

      - The end point is the highest dilution showing any visible black clumps.

- Quality Control Procedure

Each testing device has an inbuilt Control Dot to confirm the validity of the device. The appearance of the control dot after completing the procedural test indicates a valid test.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – IAMM VIROLOGY EQAS, MHL EQAS

- Interfering Substances

Frozen Samples

      - The Test is best when used with fresh samples that have not been frozen and thawed. However, most frozen samples will perform well if the following suggested procedure is followed. Allow the sample to thaw in a vertical position in the rack. Do not shake the sample. This allows particles to settle to the bottom. If a centrifuge is available, the sample can be centrifuged at 3500 rpm for 15 min.

      - Insert the dropper just below the top surface of the sample and withdraw one drop of sample.

      - If the above procedure still yields a high background, dilute 1 drop of sample with 2 drops of normal saline. Use 1 drop of this diluted sample in the test.

Thick or Viscous Samples

      - Whenever possible, clear specimens should be used. However, viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3500 rpm for 15 minutes and retested on a fresh device to avoid inconsistent results.

- Result Calculation

Qualitative Test

      - REACTIVE

- Characteristic clumping ranging from marked and intense (Reactive) to slight but definite (minimally to moderate reactive) clumps against white background, indicates that the sample is reactive for regain.

      - NON-REACTIVE

- Absence of clumps or slight roughness indicates that the sample is nonreactive for reagin.

Semi-Quantitative Test

| Undiluted 1:1 | Serum dilution 1:2 | Serum dilution 1:4 | Serum dilution 1:8 | Serum dilution 1:16 | Report |
| --- | --- | --- | --- | --- | --- |
| R | N | N | N | N | Reactive Undiluted (1:1)/ R1 |
| R | R | N | N | N | Reactive 1:2 dilution |
| R | R | R | N | N | Reactive 1:4 dilution |
| R | R | R | R | N | Reactive 1:8 dilution |

- R = Reactive; N = Non-reactive

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 4 Hours

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- Rapid Plasma Reagin Test is a non-specific flocculation test used to determine infection with Treponema pallidum. The test described here is primarily valid as a screening procedure. It is advisable to confirm positive samples with other methods employing specific treponemal antigen, e.g. TPHA, FTA-ABS etc. A serum titre of >= 1:8 is considered significant. The test has prognostic value during treatment.

- Potential Source of Variability

This method may give biological false positive (BFP) results in the presence of diseases such as leprosy, malaria, infectious mononucleosis, hepatitis, mumps, rheumatoid arthritis and collagen disease. When BFPs do occur, they may be of low titre (<1:8).

Presence of dust, high heat, poor light and the drying of the reaction mixture before being read also affects results, particularly when the test is performed in field conditions.

- References

- RPR Test – RECKON kit insert