SOP – 002

PROCEDURE FOR OPERATING THE CENTRIFUGE

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Procedure	7

6.	Calibration Procedure	8

7.	Reference	11

Flow Chart 1: Remi R-8C Centrifuge Start-Up Procedure	12

Flow Chart 2: Remi R-8C Centrifuge Shut-Down Procedure	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BC | Biochemistry |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | LCD | Liquid Crystal Display |
| 5 | QC | Quality Control |
| 6 | QM | Quality Manager |
| 7 | rpm | Revolutions per minute |
| 8 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- This document describes the procedure for use, calibration and maintenance of the Centrifuges at TRUSTlab.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing tests on this instrument.

- Responsibilities

All technical members of the Biochemistry department, who operate the centrifuge, are responsible for its Operation, Calibration and Maintenance.

The Head of Department shall ensure proper implementation of this SOP.

The Quality Manager shall monitor for compliance with this SOP.

- Definition

- Centrifugation: Centrifugation is the process of separating two or more liquids in a mixture by rotation in a container so that the lighter density liquid rises to the top. It is occasionally necessary to centrifuge samples in refrigerated conditions to ensure sample integrity.

- Procedure

Requirements

- Calibrated digital Tachometer

Procedure of Usage

      - Before Centrifugation

        - Always wear gloves when handling tubes or rotors while centrifuging infectious materials or human samples.

        - Keep the instrument on stable platform.

        - Check tubes, bottles, and rotors for cracks and deformities before each use.

        - Explosive or highly flammable chemicals must not be used.

        - Check the lid for cleanliness & any damage before starting the rotor.

        - Damaged rotors will not be used.

        - Ensure that the rotor & lids (if any) are correctly attached & in place before starting the centrifuge.

        - Centrifuge tubes shall never be filled above the maximum as recommended by the manufacturer.

        - The rotor should be loaded in a balanced manner, within the limits specified. For e.g. tubes containing materials of similar densities should be used to balance the rotor on the opposing side.

        - The centrifuge should not be operated with the lid open. Centrifuge tube material should be compatible with solvents being used.

        - Only use tubes that fit correctly in the slots provided.

        - Set the desired time and speed (rpm).

        - Ensure proper labeling of the tubes.

        - To ensure balanced loading, if the number of tubes is odd, a tube filled with water, in the same quantity as the other tubes, should be kept for balancing. Close the lid and switch on the centrifuge.

      - During Centrifugation

        - Keep the lid always closed during the entire run-cycle of the centrifuge. Never open a centrifuge until the rotor has stopped.

        - Do not exceed safe rotor speed.

        - The operator must not leave the centrifuge until full operating speed is attained, and the centrifuge appears to be running safely without vibration.

        - Stop the centrifuge immediately if an unusual condition (noise or vibration) begins and check for load balance.

      - After Centrifugation

        - Allow the centrifuge to come to a complete stop before opening. Never attempt to lift the lid or slow the rotor by hand while rotor is in motion.

        - Wear a new pair of outer gloves to remove rotor and samples (using forceps).

        - Check inside the centrifuge for possible spills and leaks, clean centrifuge and rotor thoroughly if necessary.

        - Wash hands after removing gloves.

- Calibration Procedure

Use a tachometer for calibration of rpm. The device used for calibration shall be calibrated with traceability to National/International standards.

Calibration of Speed (rpm) using a Tachometer

      - Press the button provided on the right-hand side of the tachometer unit to check for the LCD display and also the light emitted from the unit. The display reads ‘0’, on pressing the button.

      - Ensure that the tachometer is focused on the shaft or flywheels directly.

      - Stick the reflective tape (provided with the tachometer) of approximately 2cm x 2cm area on the surface of the shaft (flywheel), whose rpm is to be measured.

      - Switch on the centrifuge.

      - Set any minimum three speed levels covering the working range of the centrifuge and perform the following steps:

        - Set the lower speed first.

        - Hold the tachometer in your palm and press the button on the right-hand side.

        - Beam the light focus on the reflective tape from a distance of approx. 6 to 8inches depending on the ambient light.

        - You will observe the LCD provided on the tachometer, coming ON and OFF to ensure that the reflection from the reflective tape is being sensed.

        - The tachometer reads the rpm of the rotating object and displays on  its LCD.

        - Repeat this experiment three times for each set speed value.

      - Set the remaining speed values in the operating range and follow the above steps from ‘a’ to ‘e’.

      - For best results, beam the light from the tachometer at a right angle to avoid unwanted reflections.

      - Document the details in centrifuge calibration log.

      - Calculate %difference = [(Set speed – measured speed)/Set speed] x100.

Acceptance Criteria

- The % difference shall be within the ± 5% for each set speed value.

Frequency of Calibration

- Calibrate the centrifuge on a yearly basis, on or before the due date.

Trouble Shooting

- For trouble shooting refer operator manual or call service support.

Maintenance

      - Use a clean, absorbent cloth or paper to remove excess water, to reduce potential effects of corrosion and staining of parts as and when required.

      - Leave the centrifuge cover (lid) open when not in use.

      - In case of spillage of body fluids (blood, plasma, serum and urine) – do the following:

          - Clean the area with a cotton plug dipped in 1% Sodium hypochlorite solution (preferably use a pair of tongs).

          - Wash with a mild detergent.

          - Scrub cavities with a brush.

          - Wash with tap water.

          - Wash with Milli-Q or equivalent grade water.

          - Dry the parts by wiping with a clean absorbent cloth.

      - If a tube breaks, the centrifuge should be turned off and allowed to stand undisturbed for an appropriate amount of time before opening, to allow aerosols to settle. Thirty (30) minutes is a commonly recommended waiting period. Clean and disinfect the rotor.

      - If infectious material was placed in the centrifuge, implement protocols for safe decontamination and cleanup of equipment.

      - Cleaning and disinfection of tubes, rotors, and other components requires considerable care.

      - No single method is suitable for all items, and the various manufacturers’ recommendations must be followed to avoid rotor fatigue, distortion, and corrosion.

      - After use, tubes, rotors, and centrifuge interiors shall be cleaned and/or disinfected.

      - Perform regular checks to ensure that the side and rear grills of the centrifuge are not obstructed. Obstructions restrict airflow and in turn will reduce the life of motor and refrigeration efficiency.

      - Check the carbon brushes in the rotor for wear & tear after 300-400 hours of running time approximately.

      - If the brushes are worn out, they must be replaced with new ones. Prolonged use of centrifuge even after the carbon brushes are worn out may result in burn- out of motor.

- Reference

- Equipment Manufacturer’s Operator Manual

| Name of the instrument | CENTRIFUGE | Remarks |
| --- | --- | --- |
| Date of installation | 12-03-2022 |  |
| Start Up and Shut down procedure | Refer flow chart-1 & 2 |  |
| Maintenance | Daily Maintenance |  |
| Calibration details | Once a year Calibration done date :12-03-2022 Calibration due date :11-03-2023 |  |
| Safety measures | Do not put hands on instrument. Always close the instrument top door. |  |
| Precautions if any |  | NA |
| Decontamination procedures |  | NA |
| Any other (taking out of service, breakdown and repairs |  | NA |

- Flow Chart 1: Remi R-8C Centrifuge Start-Up Procedure

Switch the centrifuge ON at the operation switch, located on the front panel

Lift the top cover of the centrifuge

Check the surfaces of the centrifuge are clean, ensure centrifuge Buckets and tubes are clean and dry

Use matched tubes, buckets and other tubes

Ensure the lid is clean

Do not overfill tubes or containers and make sure they are balanced

Inspect tubes and containers for cracks or flaws

Make sure the centrifuge is operating properly before leaving the area

Do not exceed Manufacturer's prescribe maximum run-speed

Do not open lid until the rotor has come to total stop

Seat the rotor on drive shaft properly

- Flow Chart 2: Remi R-8C Centrifuge Shut-Down Procedure

Switch off the Main power button of the electric switch board

Switch off the operation button of the centrifuge

Check the surfaces of the centrifuge are clean, check and ensure centrifuge buckets are empty and tubes are clean and dry

Ensure the lid is clean

Close the centrifuge lid

Remove samples from the centrifuge buckets

Ensure all the samples have completed centrifugation