SOP – 006

PROCEDURE FOR OPERATING TOSOH HLC-723GX HPLC ANALYZER

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Procedure	7

6.	Calibration Procedure	8

7.	Reference	11

Flow Chart 1: Remi R-8C Centrifuge Start-Up Procedure	12

Flow Chart 2: Remi R-8C Centrifuge Shut-Down Procedure	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BC | Biochemistry |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | NaF | Sodium Fluoride |
| 5 | QC | Quality Control |
| 6 | QM | Quality Manager |
| 7 | rpm | Revolutions per minute |
| 8 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- This document describes the procedure for use, calibration and maintenance of the TOSOH HLC-723G8 HPLC Analyzer at TRUSTlab, in order to operate it correctly and safely.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing tests on this instrument.

- Responsibilities

All technical members of the Biochemistry department, who operate the centrifuge, are responsible for its Operation, Calibration and Maintenance.

The Head of Department shall ensure proper implementation of this SOP.

The Quality Manager shall monitor for compliance with this SOP.

- Definition

- HPLC: High-performance liquid chromatography (HPLC) is a broad analytical chemistry technique used to separate compounds in a chemical mixture. These separations utilize the pressure-driven flow of a mobile phase through a column packed with a stationary phase.

- Procedure of Operation

START

STOP

HOME

E.RESET

Switching on the Instrument

Check for levels of Elution buffer 1, 2,3 and Wash Solution.

“SWITCH ON” the instrument, PC and GX reporting software.

“Press the “Power Button” on the sheet keys”

Wait for Instrument Initialization (Pump Clean)

System will come to “STAND-BY”

Check Column and Filter count. Replace if necessary

Instrument is ready for Routine operations.

Maintenance

- [Main Screen] –

- Press the 	key on the main screen to display the MAINTE screen

- Key Functions:

|  | Displays another screen to reset the column and filter counter or to replace the reagents |
| --- | --- |
|  | Used only by service person (nothing executes, even if pressed) |
|  | Used only by service person (nothing executes, even if pressed) |
|  | Used only by service person (nothing executes, even if pressed) |
|  | Continuously rotates the turn table. Stops the table when the key is pressed again |
|  | Initializes (washes) the sampling unit |
|  | Returns to the previous screen |

Trouble Shooting

- Pre-assay Checklist

      - The following table provides a checklist of procedures to be performed on a daily basis before starting assays (pressing START key):

| No. | Items to Check | Content | Refer to |
| --- | --- | --- | --- |
| 1 | Calibration settings | Check the CALIB key display | P.46 |
| 2 | Column | Check counter → replace | P.139 |
| 3 | Filter | Check counter → replace | P.143 |
| 4 | Elution buffers | Check volume → replace | P.128 |
| 5 | GX Hemolysis & Wash Solution | Check volume → replace | P.128 |
| 6 | USB stick | Check remaining volume → replace or initialize | P.97 |
| 7 | Printer paper | Check volume → replace | P.147 |
| 8 | Waste eluent bottle | Check waste volume → treat waste | P.45 |

      - Be sure to check the following items before starting an assay:

| No. | Items to check/replace | Maintenance schedule | Refer to |
| --- | --- | --- | --- |
| 1 | Column | After 1500 tests | P.139 |
| 2 | Filter | After 350 tests | P.143 |
| 3 | Suction filter | Every 1 year | P.149 |

      - The following items are checked by service personnel:

| No. | Items to check / replace | Service frequency (guide or target) |
| --- | --- | --- |
| 1 | Check the sample sensor | Every 2 years or After 15000 tests |
| 2 | Check needle descent position |  |
| 3 | Clean dilution port and wash assembly |  |
| 4 | Check the screws of sampling unit |  |
| 5 | Check the screw of the parts driving the valves |  |
| 6 | Check the column oven temperature |  |
| 7 | Check the solenoid valve’s action (3 locations) |  |
| 8 | Check vacuum pump’s action |  |
| 9 | Replace the rotor seal of the injection valve |  |
| 10 | Replace the rotor seal of the AS valve |  |
| 11 | Wash or replacement of the pump check valves |  |
| 12 | Replace the plunger seal |  |
| 13 | Check the waste pump’s action |  |
| 14 | Check the barcode reader |  |
| 15 | Replace the sample loop | Annually or every 20,000 sample injections |
| 16 | Replace the waste chamber body | When dirty |
| 17 | Replace the Teflon tip of syringe | When dirty or worn |
| 18 | Needle O-ring | When worn |
| 19 | Replace the drain valve packing | When worn |

Press “Calib” key on touch screen display and enter the concentration of Calibrators

Use sample cups for calibrators in Calibration Port

Press “Start” to initiate calibration

Review new calibration

After Calibration validate with control

Use sample cups for controls with Tosoh bar-coded adaptors in Turn Table

Validate Q.C data

Instrument is ready for routine operations

Use bar-coded EDTA primary tubes in sample racks and press “Start”

Review results In the Reporting software

Calibration Procedure

Acceptance Criteria

- Calibration results were in the acceptable range

Frequency of Calibration

- Calibrate the instrument on a yearly basis, on or before the due date.

- Reference

- TOSOH Equipment Manufacturer’s Operator Manual