SOP – 007

PROCEDURE FOR

THE OPERATION OF ERBA LAURA URINE ANALYZER

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Principle of Working	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration Procedure	9

13.	Procedure	9

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	11

16.	Result Calculation/Interpretation	11

17.	N/A.	11

18.	Biological Reference Interval	11

19.	Reportable Interval	11

20.	Procedure when Results are not in Range	11

21.	Critical Alert Values	11

22.	Laboratory Clinical Interpretation	11

23.	Potential Source of Variability	11

24.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | CP | Clinical Pathology |
| 3 | CV | Coefficient of Variation |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HBV | Hepatitis B Virus |
| 6 | HIV | Human Immunodeficiency Virus |
| 7 | HOD | Head of Department |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | pH | Potential of Hydrogen |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to provide guidelines for the safe and proper operation and maintenance of the ERBA LAURA Urine Analyzer used for Urine Analysis in the Clinical Pathology Department at TRUSTlab. This SOP also provides information about the procedural principles and other details.

- Scope

- This SOP applies to all laboratory personnel of the Clinical Pathology Department responsible for operating this Urine Analyzer.

- Responsibilities

It is the responsibility of the technicians of the Clinical Pathology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Urine analysis, also known as urinalysis, is a medical test that examines the physical, chemical, and microscopic properties of urine. It's a common diagnostic tool used to assess overall health, kidney function, and detect various conditions like infections or diabetes.

- Principle of Working

- The Urine Analyzer Erba Mannheim LAURA Version 2 is a semi-automated system designed to analyze urine samples for various parameters such as pH, specific gravity, protein, glucose, ketones, bilirubin, urobilinogen, nitrites, leukocytes, and blood.

- Specific Performance Characteristics

- N/A.

- Type of Sample

Urine is the primary specimen for use in this instrument.

Collect a minimum 15 ml of a random sample of urine.

If required, centrifuge the sample to remove any particulate matter before analysis.

Process the samples within 1 hour of collection. If delay is anticipated instruct patients to store the sample at 2-8°C in the refrigerator for a maximum of up to 4 hours until transport to the laboratory.

Do not accept if the sample volume is inadequate or if samples are received without an ID number.

- Type of Container & Additives

- Sterile screw-capped container.

- Patient Preparation

N/A.

- Equipment & Reagents

Equipment

- ERBA LAURA Urine Analyzer

Reagents

- DEKAPHAN Urine Strips

- Safety

Follow universal precautions like wearing suitable gloves and eye/face protection while handling all reagents, samples and operating the instrument.

Universal precautions imply that all body fluids are treated as if known to be infectious for Human Immunodeficiency Virus (HIV), Hepatitis B Virus (HBV) or other Blood borne Pathogens.

Dispose of all waste in accordance with applicable national and/or local regulations.

Some reagents contain sodium azide, which may react with copper or lead plumbing to form explosive metal azides. Use caution in disposing of these reagents. If disposing to drain, flush with large volumes of water to prevent azide buildup.

Waste material containing patient samples or biological products should be considered biohazardous when disposing or treating.

Clean up all spills immediately and thoroughly. Disinfect the area for any spills involving biohazardous materials. Dispose of all contaminated materials appropriately.

- Calibration Procedure

      - Perform daily calibration using the provided calibrators and controls, as per the manufacturer's instructions.

      - LAURA® automatically checks itself after each switch-on.

      - The calibration of LAURA® can be checked via the control mode from SW menu.

      - The special control strips are using for the checking’s of the calibration and the right function of the optical parts.

      - Record the calibration results and adjust the analyzer if necessary.

- Procedure

Preparation of the Instrument:

      - Ensure that the Urine Analyzer is placed on a clean, stable and level surface.

      - Power on the analyzer and allow it to warm up as per the manufacturer's instructions.

Procedure:

      - Take the required amount of urine sample as specified by the manufacturer and load it onto the sample tray of the analyzer.

      - Enter sample information such as patient ID, sample ID, and other relevant details into the system.

      - Start the analysis using the appropriate test parameters and reagents.

      - Data Interpretation:

          - Allow the Urine Analyzer to complete the analysis process.

          - Record the results displayed on the analyzer's screen or print them as required.

Maintenance:

      - The software of LAURA® contains a special mode for maintaining cleanliness and general hygiene during daily operations.

      - This mode allows easy access to all parts which are in contact with urine.

      - This mode eases the maintenance and the cleaning of the reader.

      - It is provided with single-use plastic parts that can be disposed of in the waste container, thus ensuring hygienic conditions of operation.

      - The Analyzer automatically recognizes if the waste container is to full and flashes a message on the display.

      - The lab personnel must keep the instrument clean and free from any spills or debris.

      - Regularly check and replace reagents and consumables when required.

Troubleshooting:

      - In case of any error messages or unusual results, consult the instrument manual for troubleshooting guidelines.

      - b. If the issue persists, inform the laboratory supervisor or maintenance personnel.

- Quality Control Procedure

- As defined in the procedure Urine IQC (TDPL/CP/SOP-002).

- Interferences & Cross Reactions

- N/A.

- Result Calculation/Interpretation

- N/A.

- Biological Reference Interval

- Negative.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

A microscopic examination of urinary sediment for pus cells, bacteria and casts is often used to correlate a positive test for urine protein.

This examination aids in laboratory diagnosis of renal and other disorders.

- Potential Source of Variability

- N/A.

- References

      - ERBA LAURA Urine Analyzer Operator’s Manual