SOP – 014

PROCEDURE FOR

DRY BATH INCUBATOR

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Principle	7

5.	Environment & Safety Control	7

6.	Calibration Procedure	8

7.	Equipment Record	8

8.	Procedure	9

9.	Maintenance	10

10.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | HOD | Head of Department |
| 3 | MSP | Management System Procedure |
| 4 | N/A | Not Applicable/Available |
| 5 | QC | Quality Control |
| 6 | QM | Quality Manager |
| 7 | SOP | Standard Operating Procedure |
| 8 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to provide guidelines for operating and maintaining the Dry Bath Incubator in the Molecular Biology Department, to ensure safe and proper usage.

- Scope

- This procedure applies to all staff of the Molecular Biology Department.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Principle

The dry bath incubator is a microprocessor-controlled thermostat device with a high precision of temperature control. It is used for sample preparation. It is used as an alternative to the traditional water bath devices. It is used in the cultivation, preservation, and reaction of various samples.

- Environment & Safety Control

| General Safety Instruction | DANGER – Misuse of electrical equipment can cause electrocution, burns, fire and other hazards Basic safety precautions always should be taken, including all those listed below. Close supervision is necessary when equipment is used by or near children, handicapped persons or invalid Persons. Check that the voltage setting matches the supply voltage. Plug the equipment into a supply outlet which has an earth connection. Switch off the MAINS supply, immediately after use. Do not place the equipment in liquid, nor put it where it could fall into liquid. If the equipment becomes wet, unplug it before touching it. Use the equipment only for the purpose described in the instructions for use / intended use. Do not use accessories which are not supplied or recommended by the manufacturer. Do not use the equipment if it is not working properly or if it has suffered any damage. Note Examples of typical defects include: a] Damage to the flexible supply cord or its plug. b] Damage caused by dropping the equipment. c] Damage caused by dropping the equipment into water or splashing water onto it. Send this device to service center without delay. Do not let the equipment or its flexible cord come into contact with surfaces which are too hot to touch. Do not block air openings nor place equipment on a soft surface which might block them, and keep air openings free from lint, hair, fluff, etc. Do not place anything on top of the equipment. Do not use the equipment out of doors |
| --- | --- |

- Calibration Procedure

Frequency: Once a year.

Procedure:

      - Switch ON the instrument.

      - Put the Distilled water in tube and place it in the five holes of the dry bath incubator, four corner & one centers.

      - Leave for 15minutes for stabilization.

      - Put the calibrated thermometer for each location one by one & record the Result after stabilization.

- Equipment Record

      - Equipment Name: Dry Bath Incubator

      - Manufacturer Name: LabLine

      - Serial No & equipment ID: 5508 &  TDPL/PR/CL-037

      - Contact number application engineer:

      - Location: Central Lab, Begumpet Hyderabad

- Procedure

Starting the instrument:

- Connect the device to power outlet.

- Turn power switch “ON’’.

- The machine will start for functional work.

- Display will show default Temperature and Timer setting (37°C and 30 Minutes respectively).

        - For Adjustment of Temperature (TEMP):

          - UP: It is a Programming switch, by pressing it, increase the Temp.

          - DOWN: It is a Programming switch, by pressing it, decrease the Temp.

        - For Adjustment of Time:

          - UP: It is a Programming switch, by pressing it, increase the Temp.

          - DOWN: It is a Programming switch, by pressing it, decrease the Temp.

        - NOTE: - While “UP” or “DOWN” keys are pressed “TIME” in display will blink.

- Once desired Temperature & Time are SET, press “STA” key to start warming process.

- NOTE: - Once “STA” key is pressed, “TEMP” or “TIME” will stop blinking in display. They will be steady, which shows machine is now under working condition.

- In center of display “HTR” will show condition of heater i.e.: - “ON” or “OFF”.

- If user needs to use machine at default setting (37°Cand 30 Minutes respectively), they need to turn power switch “ON” & press “STA” key. Machine will run at this setting.

- Once set time is over, the machine will give buzzer sound and stop working.

- After working turn power switch OFF

Heating Block Replacement Procedure:

        - Remove the supply power cord from device.

        - It is advisable to wait till heating block turns cool before replacing it

        - Open the two screws provided on top side of heating block.

        - Hold the top side knob and pull the block toward top side with little force.

        - Note: Never give power supply to device when the block is removed, it will damage device.

        - Put the required Heating block in place and fix the two screws from the top side.

        - Fix the power supply cord.

- Maintenance

| Procedure of cleaning and decontamination: | Remove power supply cable from socket. Do not Connect plug if device is wet. Use safety glows, to save your fingers from contamination. Dry Bath incubator can be cleaned with Isopropyl Alcohol or 70% ethanol. Use of any other disinfectant may damage the device and parts, so not permitted. Clean the device and parts by wiping with 70%ethanol. The Deeping in chemical for long time may damage the parts. |
| --- | --- |
| Instrument Maintain | Routine cleaning: It is recommended to clean device at every six months. (Caution: Always remove power cord from source plug before performing any type of service or cleaning to device.) Electrical components: Do not need any routine or preventive maintenance. If the device needs to be sent to the service center then it must be cleaned and decontaminated for safety. It is advised to get check earthing continuity after every 12 months. |

- References

- N/A