SOP – 018

PROCEDURE FOR

USAGE OF INOCULATING LOOP

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Procedural Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	8

13.	Procedure of Instrument Operation	9

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	10

16.	Biological Reference Interval	10

17.	Reportable Interval	11

18.	Procedure when Results are not in Range	11

19.	Laboratory Clinical Interpretation	11

20.	Potential Source of Variability	11

21.	References	11

Flowchart: Start Up & Shut Down Procedure for LAF	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | HOD | Head of Department |
| 3 | MSP | Management System Procedure |
| 4 | N/A | Not Applicable/Available |
| 5 | QC | Quality Control |
| 6 | QM | Quality Manager |
| 7 | SDA | Sabouraud Dextrose Agar |
| 8 | SOP | Standard Operating Procedure |
| 9 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose and scope of this SOP is to describe procedure for maintenance of the inoculating loop used for inoculating specimens onto various media used for isolation of pathogens in microbiology department.

- Scope

- This procedure applies to all staff of the Microbiology Department.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Definitions

An inoculation loop, also called a smear loop, inoculation wand or microstreaker, is a simple tool used mainly by microbiologists to retrieve an inoculum from a culture of microorganisms. The loop is used in the cultivation of microbes on plates by transferring inoculum for streaking.

- Principle

      - Nichrome loops calibrated to deliver 0.01 ml (10 ul) are used to inoculate culture media with liquid specimens when colony counts are required.

      - The loops are composed of 26 gauge twisted nichrome wire. Although applicable to nearly any liquid or specimen, this inoculation method is most frequently used when urine is submitted for bacterial culture.

- Specific Performance Characteristics

- N/A

- Type of Sample

- N/A.

- Type of Container & Additives

- N/A.

- Patient Preparation

- N/A.

- Equipment & Reagents

Equipment:

- Inoculation Loop

Reagents & Consumables

      - Calibrated pipette of 0.001 to 2 micro liter

      - Calibrated loop measuring 0.005

      - Filter paper

      - Methylene blue

- Safety

Handle all reagents with care and avoid contact with eye, mouth, and skin.

Handle all samples as potentially infectious.

Discard used reagents and sample as per disposal procedure.

- Calibration Procedure

Frequency:

- Calibration of the loop volume is based on the thickness of the wire and circumference of the loop.

Procedure:

      - Take 0.05 microliter of Methylene blue in a test tube and then with a calibrated loop blot on to the filter paper for 10 times. After blotting the same volume of methylene blue on the filter paper, no volume of methylene blue was retained in the test tube

Observations:

      - No Volume should remain in the test tube after calibration.

      - The diameter of the inoculating loop should be within regular margins.

- Procedure of Instrument Operation

Preparation

      - Nichrome wire loop is heated on a flame to red hot stage, before use.

      - Different loops have been labeled for urine, pus and sputum and for antibiotic susceptibility testing, as per their usage.

Maintenance

    - Daily Maintenance:

- The delivery volume of calibrated loops may change due to damage, corrosion, or build-up of incinerated material.

      - It should be visually inspected to confirm that the loop is still round, and the circle has not parted.

      - The continual heating and cooling of the Nichrome wire loop to sterilize between use means there is always a risk that the 10 μl calibrated loop can become distorted or there is an excessive buildup of burnt organic material on the surface of the loop.

      - Both conditions can significantly alter the liquid volume delivered by the loop and hence the accuracy.

      - The loop should be sterilized in the blue part of the Bunsen flame till red hot, between streaking of two samples.

      - After inoculation of the given sample, the loop is dipped in spirit solution for half an hour for disinfection and cleaning of the loop.

    - Weekly Maintenance:

      - To verify that the loops have maintained their proper size, internal calibration of inoculating loop must be done by methylene blue stain.

      - The procedure is as described in Section 12 of this SOP.

Troubleshooting

- N/A.

- Quality Control Procedure

- N/A.

- Interferences & Cross Reactions

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

- Hi-Media Labs – Inoculation Loop Maintenance