MSP - 003

PROCEDURE FOR

EQUIPMENT MANAGEMENT

TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Procedure	7

5.1	Equipment Planning	7

5.2	Procurement of Biomedical Equipment	8

5.3	Installation Flow	8

5.4	Process for Calibration	9

5.5	Calibration Checklist	9

5.6	Process of Preventive Maintenance	9

5.7	Acquiring the AMC	10

5.8	Acquiring the CMC	11

5.9	Breakdown Management Procedure	11

5.10	Breakdown Record Maintenance	11

5.11	End-User Training	12

5.12	Condemnation of Equipment	12

6.	Quality Control	13

7.	Related Documents	13

8.	Reference Regulations & Guidelines	13

9.	Related Records	13

Annexure 1	15

APPROVAL AUTHORITY

| Prepared by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Issued by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Approved by: | Chairman & Managing Director | Mr. Venkata Cherukuri Name | Date: 14/11/2024 |
|  |  | Signature |  |

RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 04 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

DISTRIBUTION LIST

LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMC | Annual Maintenance Contract |
| 2 | BE | Biomedical Engineering |
| 3 | BME | Biomedical Engineering |
| 4 | CMC | Comprehensive Maintenance Contract |
| 5 | FOL | Folder |
| 6 | FOM | Form |
| 7 | GE | General |
| 8 | ISO | International Organization for Standardization |
| 9 | MAN | Manual |
| 10 | MSP | Management System Procedure |
| 11 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 12 | PM | Preventive Maintenance |
| 13 | QA | Quality Assurance |
| 14 | REG | Register |
| 15 | RPM | Revolutions per Minute |
| 16 | SOP | Standard Operating Procedure |
| 17 | TAMS | TRUSTlab Asset Management Software |
| 18 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this procedure is to provide specific instructions for equipment management; selection, installation, calibration, verification, breakdowns, etc. together with the reception, storage, acceptance testing and disposal of the equipment.

- Scope

- The scope covers all medical equipment and allied instruments, used in the testing of a patient’s biological specimens, that are procured by TRUSTlab either directly from local manufacturers and distributors, or by importing them from other countries.

- Responsibilities

It is the responsibility of the Biomedical Engineering department and Lab Manager to implement this policy.

It is the responsibility of the Biomedical Engineering department to identify, get quotations, procure, oversee installation, preventive maintenance & breakdown repairs, and dispose of the equipment.

It is the responsibility of the Biomedical Engineering department to remove the approved-for-disposal devices from the departments and store them before disposing of the equipment.

- Definitions

- N/A.

- Procedure

- Equipment Planning

    - One month in advance of the start of the financial year, user departments have to fill the New Equipment Requisition Form (TDPL/BE/FOM-004) and submit to the Management justifying the same. The requests are certified by the Management and the same is handed to the Purchase committee for further actions.

- Procurement of Biomedical Equipment

    - The Biomedical Engineering (BME) Manager is also a part of the Purchase Committee.

    - Any new equipment requirement is intimated to BME, and an assessment for the same is done i.e. to check whether the requirement is appropriate or not.

    - If found appropriate, then the suppliers are contacted to get the details regarding the same equipment.

    - Technical and commercial comparisons are prepared and verified.

    - The BME department co-ordinates with the Purchase Committee in the price negotiations.

    - Once the details like technical specifications, price etc. are finalized by the Purchase Committee, these details are forwarded to the Purchase department. The Purchase Order is generated & forwarded to the vendor.

    - The Purchase department then takes further steps to procure the same.

- Installation Flow

  - The new equipment is received at the location as is mentioned in the Purchase Order.

  - The BME department is responsible for checking equipment details and spare parts details.

  - Before installation, the pre-installation details are acquired from the manufacturer/dealer.

  - The corresponding department is informed regarding the pre-installation requirements. E.g. the Maintenance Department will be informed of any electrical and plumbing work to be done on site before installation.

  - After the arrival of the consignment, it is moved to a storage room/installation site.

  - Once the site is ready, the instrument engineer is informed to visit the site and to start installation if all the criteria are met.

  - Consignment is opened in the presence of either a BME team member and/or an assigned employee who is a qualified clinical operations member.

  - Once the installation is complete a demo is given to the BME department, and the application training given to the end-user.

- Process for Calibration

    - A list of medical equipment/devices requiring calibration is prepared and maintained in TAMS.

    - Each calibration is done through external sources, and the status is updated continuously in TAMS.

    - Instruments that require to be calibrated at an outside location are collected and sent to the identified calibration agency.

    - Calibration records & reports are maintained within the BME department.

    - The equipment is well maintained to preserve its accuracy and fitness for use.

- Calibration Checklist

- Whenever a Calibration is done, the following is confirmed:

    - Physical condition of instrument /test equipment.

    - The Calibration Checklist provided by the equipment manufacturer is followed to check whether the equipment is running as per the norms.

    - Calibration certificate with traceability to be obtained from calibration agency with verification marked as O.K. /Not O.K. The original hardcopy of the calibration certificate is maintained with the BME department in Equipment Calibration Certificates Folder (TDPL/BE/FOL-003).  A soft copy is uploaded into the respective equipment files in TAMS software and a copy is provided to the user department too.

    - Calibration sticker is displayed on the machine which shows the last calibration date and the next calibration due date.

- Process of Preventive Maintenance

    - The Preventive Maintenance of an instrument having an AMC/CMC contract is done by communicating with the BME Manager and company engineer.

    - The availability of necessary spares, consumables, tools and necessary materials are ensured through standardization and/or advance planning, through Stores and guidance by BME department Manager.

    - Carrying out PM as per PM schedule –

      - Preventive maintenance schedule is prepared periodically by BME department.

      - The BME department identifies the equipment for PM for the applicable month from the PM schedule.

      - BME co-ordinates with the user department for available date/time of equipment for PM within the month.

    - For PM to be carried out In-house –

      - PM for equipment are carried on scheduled date by BME.

      - After PM, the PM report is duly signed by BME Manager and End user.

      - PM reports are maintained with the BME department in the Equipment Preventive Maintenance Service Folder (TDPL/BE/FOL-005) and also maintained in TAMS software.

      - The BME department places a sticker on the equipment with details of the PM done and the next due date.

    - For PM to be carried out by vendor –

      - BME department coordinates with the vendor to fix PM on scheduled date/time.

      - Carry out PM on scheduled date/time in presence of BME Manager.

      - Instrument Engineer prepares PM report which is counter-signed by the end user.

      - Instrument Engineer submits PM report with checklist to BME department.

      - PM reports are verified by the BME department and signed by BME Manager.

      - PM reports are filed with BME department in the Equipment Preventive Maintenance Service Folder (TDPL/BE/FOL-005) and also maintained in TAMS software.

      - The BME department places a sticker on the equipment with details of the PM done and the next due date.

- Acquiring the AMC

    - The details regarding the equipment are collected one month prior to the warranty expiry date. The BME Team receives notification of the same via TAMS.

    - Depending on service history, the necessity of an AMC is evaluated.

    - The AMC Proposal is requested from the respective company.

    - The price is negotiated and the number of breakdowns and preventive maintenance calls per annum is discussed.

    - The quotation is then approved by the Chairman & Managing Director.

    - The same is then forwarded to the Purchase department.

    - Then the Purchase Order is released, and documentation of the contract is done.

    - The hardcopies of the AMCs are maintained with the BE Department in the Equipment Warranties/AMC/CMC Folder (TDPL/BE/FOL-004).

- Acquiring the CMC

    - The details regarding the equipment are collected one month prior to the warranty expiry date.

    - Depending on service history, the necessity of a CMC is evaluated.

    - The quotation regarding CMC is asked for from the respective equipment supplier/manufacturer.

    - The price is negotiated. The accessories included in the contract and the number of maintenance calls per annum are discussed.

    - The quotation is then approved by the Chairman & Managing Director.

    - The same is then forwarded to the Purchase department.

    - Then the Purchase Order is released, and documentation of the contract is done.

    - The hardcopies of the CMCs are maintained with the BE Department in the Equipment Warranties/AMC/CMC Folder (TDPL/BE/FOL-004).

- Breakdown Management Procedure

    - The user department raises a ticket in the TAMS software to register a complaint about the breakdown and informs the BME department.

    - Where applicable, the BME Manager allots the complaint to the staff in the department.

    - The staff then tries to rectify the breakdown that has occurred.

    - If it cannot be repaired by in-house staff, the company supplying the equipment is informed about the complaint.

    - The work of the company representative is supervised.

    - If the equipment is irreparable then it is considered for condemnation.

- Breakdown Record Maintenance

    - The breakdown of an equipment or device is entered into the TAMS software through the ticketing process and then reported to BME department.

    - Instruments/devices which are under AMC (Annual Maintenance Contract) /CMC (Comprehensive Maintenance Contract) are given to the respective Company for maintenance.

    - If instruments are out of AMC/CMC, the BME department assigns the breakdown to an approved supplier.

    - A report of failure/breakdown is taken from the said Company for monitoring purposes and is filed in the Equipment Breakdown Maintenance Folder (TDPL/BE/FOL-006). The details of the breakdown are also recorded in TAMS in the Equipment Breakdown Service Register (TDPL/BE/REG-001).

    - In-house servicing of equipment is sometimes done for breakdowns, like for errors which arise due to improper application as well as other minor errors.

    - Once the equipment is thus repaired, a Work Order Completion of Equipment Repairs & Maintenance Form (TDPL/BE/FOM-003) is prepared, and the end-user’s signature is taken on the same. The report is filed in the Equipment Breakdown Maintenance Folder (TDPL/BE/FOL-006). The details of the breakdown are also recorded in TAMS in the Equipment Breakdown Service Register (TDPL/BE/REG-001).

- End-User Training

    - As and when new equipment is installed or whenever requested by the user departments, a training session is conducted related to the respective equipment.

    - Training certificates are issued to the attendees by the Company providing the equipment. These certificates are maintained in the Personnel Folder (TDPL/HR/FOL-004) and details are maintained in the Employee Training/Re-training Details & Attendance Form (TDPL/HR/FOM-032).

- Condemnation of Equipment

    - On finding out that a piece of equipment is beyond economical repair, an official letter is given to the Chairman & Managing Director describing its status and cost of repair.

    - Once the request for condemnation is approved by the Chairman & Managing Director the equipment is brought back to BME department and an Equipment Disposal Form (TDPL/BE/FOM-005) is filled out by the BME Manager.

    - The equipment then undergoes a process called cannibalization where as far as possible the spare parts are used up.

    - On receiving intimation from Central Stores that e-waste disposal is going to take place in the institution, a list of all the equipment to be disposed of is made.

    - These pieces of equipment are then sent to Central Stores from where it is given out for disposal to the vendors that are authorized by the Pollution Control Board.

    - A record of the condemned equipment is maintained in TAMS

- Quality Control

- The Quality Manager or designee will do random Quality Control checks to ensure that this SOP is implemented as written. If Quality control checks reveal non-compliance with this SOP, the SOP deviations and violations will be investigated, corrective and preventive action initiated, implemented, reviewed and closed out. All corrective and preventive action will be documented in the Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021).

- Related Documents

- N/A

- Reference Regulations & Guidelines

  - Specific Criteria for Accreditation of Medical Laboratories (NABL 112 A; Issue 01)

  - Guidance document: Medical Laboratories (NABL 112B; Issue 01)

  - ISO 15189:2022 Medical Laboratories – Requirements for Quality and Competence

- Related Records

  - New Equipment Requisition Form (TDPL/BE/FOM-004)

  - Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021)

  - Equipment Preventive Maintenance Register (TDPL/BE/REG-003)

  - Equipment Breakdown Service Register (TDPL/BE/REG-001)

  - Equipment Calibration Certificates Folder (TDPL/BE/FOL-003)

  - Equipment Preventive Maintenance Service Folder (TDPL/BE/FOL-005)

  - Equipment Warranties/AMC/CMC Folder (TDPL/BE/FOL-004)

  - Equipment Breakdown Maintenance Folder (TDPL/BE/FOL-006)

  - Work Order Completion of Equipment Repairs & Maintenance Form (TDPL/BE/FOM-003)

  - Personnel Folder (TDPL/HR/FOL-004)

  - Employee Training/Re-Training Details & Attendance Form (TDPL/HR/FOM-032)

  - Equipment Disposal Form (TDPL/BE/FOM-005)

- Annexure 1

| Equipment | Maintenance Step | Frequency | Maintenance done by |
| --- | --- | --- | --- |
| Autoclave | Record temperature and pressure | Each use | Internal Staff member |
|  | Cleaning and check functioning with an indicator | Monthly | Housekeeper supervised by Lab Manager |
|  | Clean and adjust | Annually | Internal staff member |
|  | Unscheduled service | As required | Internal staff member |
| Incubator | Check door seal | Quarterly | Internal Staff member |
|  | Clean | Quarterly | Housekeeper supervised by Lab Manager |
|  | Recalibrate thermometer | Annually | Approved Third party Vendor |
|  | Unscheduled service | As required | Internal staff member |
| Hot air oven | Record temperature and time | Each use | Housekeeper supervised by Lab Manager |
|  | Cleaning | Monthly | Housekeeper supervised by Lab Manager |
|  | Unscheduled service | As required | Internal staff member |
| Biosafety Cabinets | Disinfect work surface & Check air pressure | Daily | First user of the day |
|  | Disinfect entire hood | Bi-Monthly | First user of the 1st and 15th day of the month |
|  | Decontamination, Testing, Certification | Annual | Approved Third Party Vendor |
|  | Unscheduled Service | As Required | Internal staff member |
| Centrifuge | Clean | Weekly | Internal staff member |
|  | RPM Check by calibrated Tachometer | Annual | Approved Third Party Vendor |
|  | Unscheduled Service | As Required | Internal staff member |
| Microscope | Cleaning, turn off light, cover | Each use | User |
|  | Unscheduled service | As Required | Internal staff member |
| Freezer | Measure Temperature | Daily | Internal staff member |
|  | Check door seal | Quarterly | Internal staff member |
|  | Clean and Defrost if required | Monthly | Housekeeper supervised by Lab Manager |
|  | Recalibrate thermometer | Annual | Approved Third Party Vendor |
|  | Unscheduled service | As Required | Internal staff member |
| Pipettes | Inspect and Calibrate | Annual | Approved Third Party Vendor |
|  | Unscheduled service | As Required | Internal staff member |
| Refrigerators | Measure Temperature | Daily | Internal staff member |
|  | Check door seal | Quarterly | Internal staff member |
|  | Cleaning | Monthly | Housekeeper supervised by Lab Manager |
|  | Recalibrate thermometer | Annual | Approved Third Party Vendor |
|  | Unscheduled service | As Required | Internal staff member |
| Water baths | Record temperature, check water level, Replace water | Each use | Internal staff member |
|  | Cleaning | Monthly | Housekeeper supervised by Lab Manager |
|  | Recalibrate thermometer | Annual | Approved Third Party Vendor |
|  | Unscheduled service | As Required | Internal staff member |
| Vortex | Cleaning | Monthly | Internal staff member |
| Shaker | Cleaning | Monthly | Internal staff member |