MSP - 012

PROCEDURE FOR MEASUREMENT OF UNCERTAINTY

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Procedure	8

6.	Quality Control	10

7.	Related Documents	10

8.	Reference Regulations & Guidelines	10

9.	Related Records	11

- APPROVAL AUTHORITY

| Prepared by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Issued by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Approved by: | Chairman & Managing Director | Mr. Venkata Cherukuri Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | CI | Confidence Interval |
| 2 | CLIA | Clinical Laboratory Improvement Amendments |
| 3 | CV | Coefficient of Variation |
| 4 | EQA | External Quality Assurance |
| 5 | ISO | International Organization for Standardization |
| 6 | MAN | Manual |
| 7 | MSP | Management System Procedure |
| 8 | MU | Measurement Uncertainty |
| 9 | PT | Proficiency Testing |
| 10 | QA | Quality Assurance |
| 11 | SD | Standard Deviation |
| 12 | SOP | Standard Operating Procedure |
| 13 | TDPL | TRUSTlab Diagnostics Private Limited |

Objective

- To determine Measurement Uncertainty (MU) for each measurement procedure in the scope of tests.  TRUSTlab estimates measurement of uncertainty because a measurement only gives an approximation of the true value of the quantity to be measured. TRUSTlab also uses MU to monitor the performance of the quality management system.

Scope

  - The SOP covers the estimation, application and communication of measurement of uncertainty in TRUSTlab. This SOP applies to all TRUSTlab personnel involved in performing routine procedures in the laboratory.

Responsibilities

  - Quality Manager

  - Lab Manager

  - Lab Head

Definitions

- Analyte: The substance or constituent of interest that is the subject of measurement. However, a substance can have a number of properties, some or all of which can be utilized to quantify the substance in an appropriate measuring system.

- Uncertainty: A parameter associated with the result of a measurement that characterizes the dispersion of the values that could reasonably be attributed to the measurand.

- Traceability: Property of the result of a measurement or the value of a standard, whereby it can be related to stated references, usually national or international standards, through an unbroken chain of comparisons all having stated uncertainties.

- Repeatability: A measure of variability derived under specified repeatability conditions, i.e. independent test results are obtained with the same method on identical test items in the same laboratory by the same analyst using the same equipment, batch of culture media and diluents, and tested within short intervals of time.

- Reproducibility: A measure of precision derived under reproducibility conditions i.e. test results are obtained with the same method on identical test items in different laboratories with different operators using different equipment. A valid statement of reproducibility requires specification of the conditions used.

- Measurand: A measurand is defined as the particular quantity subject to measurement (or the particular quantifiable property of the analyte used in the measuring system), where the quantity is the attribute of a substance that can be distinguished and determined quantitatively. It is essential to define as fully as possible the quantity that is measured (i.e. the measurand) by a given procedure. There are four aspects of a measurand that should be described (see Table 1):

- Quantity intended to be measured

- System

- Kind-of-quantity and measurement unit

- Method

- Table 1: Examples of Measurand Definition

| Quantity intended to be measured | System | Kind-of-quantity | Measurement unit | Method |
| --- | --- | --- | --- | --- |
| ALT | Venous serum | Amount of substance concentration | mmol/L | Flame photometry |
| Creatine kinase MB | Serum | Activity concentration | mIU/L at 37ºC | Immuno-inhibition |
| Haemoglobin | Venous whole blood | Mass concentration | g/dl | Spectrophotometry |

Procedure

Determine MU on measured quantity values (Quantitative Tests)

- The uncertainty of measurement is only determined for all quantitative tests within the laboratory. One of the components of MU is precision/imprecision. This is obtaining from processing stable controls. TDPL processes controls for each analyte where measurements are in metric values with frequency described under each discipline. The SD and % CV are derived from the laboratory mean and not from the control’s target value assigned by the manufacturer. Actual % CV up to the first decimal place for each parameter is used for calculation. Imprecision data obtained from the routine application of internal quality control is recommended as the quantitative estimate of MU. The imprecision is derived is derived from long term IQC data, calculated as SD or %CV. For the purpose of recording estimates of uncertainty of measurement, imprecision is documented as K x % CV. K is the coverage factor and at 95% confidence interval, it equals to ± 1.96. So, the MU is set as % CV x 1.96. Three months’ % CV is used to calculate MU. This quarterly MU is also used to monitor the laboratory performance (as a quality indicator).

Mean

- The Mean is defined as the arithmetic average of a set of data points. It is expressed as:

- Mean= Σxin

- Where,

- xi = Individual observations / data / results in the set

- n = the total number of data /observations / results in the set

- Σ = Symbol for Summation

Standard Deviation

- The standard deviation (SD) is a measure of the dispersion of observations/data/data points above and below the Mean and is used to set acceptable limits for values obtained with internal QC samples.

- SD= (xi-x)x2(n-1)

- Where,

- Σ(x2) = the sum of the squares of each value of x

- (Σx)2 = the square of the sum of all data points

- n = the total number of data points in the set

Coefficient of Variation

- The coefficient of variation (CV) is a measure of the variability of an assay and is expressed as a percentage:

- CV%= SDMean×100

Measurement Uncertainty

- MU=CV%2

Acceptance Criteria

- Performance Requirements for the measurement uncertainty (e.g. Standard Deviation; Clinical decision points)

- The MU is accepted if it’s less than the total allowable error as per CLIA limits.

- In case the MU is greater than the total allowable error, verify the steps

- If verification confirms that the MU is greater than the total allowable error, perform a root cause analysis and corrective action (handle as a nonconformity).

Application of MU

- Every measurand whose result is generated by the laboratory is compared to the respective uncertainty within which the true value lies.

Reporting MU

- An MU report is written in a format shown in appendix 1 and is made available to the clinicians up on request.

- Quality Control

- The Quality Manager or designee will do random Quality Control checks to ensure that this SOP is implemented as written. If Quality control checks reveal non-compliance with this SOP, the SOP deviations and violations will be investigated, corrective and preventive action initiated, implemented, reviewed and closed out. All corrective and preventive action should be documented in the Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021).

Related Documents

TRUSTlab’s Quality Manual (TDPL/QA/MAN-001)

Reference Regulations & Guidelines

Specific Criteria for Accreditation of Medical Laboratories (NABL 112 A; Issue 01)

Guidance document: Medical Laboratories (NABL 112B; Issue 01)

ISO 15189:2022 Medical Laboratories – Requirements for Quality and Competence

ISO 13528:2005(E) Statistical methods for use in proficiency testing by interlaboratory comparisons

ISO 15189: 2012 Medical laboratories – Particular requirements for Quality and competence (clause 5.5.1.4)

EURACHEM/CITAC Guide CH 4, Quantifying Uncertainty in Analytical Measurement, 2nd Edition, 2000.

Checklist E-186: “Measurement Uncertainty in Chemical Analysis”

Order from www.labcompliance.com/solutions/examples

Monica-Cheesbrough-District-Laboratory-Practice-in-Tropical-Countries-Part-1

Related Records

Non-Conformity & Corrective Action Form (TDPL/GE/FOM-021)