SOP – 012

PROCEDURE FOR

TOTAL BILIRUBIN TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	12

16.	Result Calculation	13

17.	Biological Reference Interval	13

18.	Reportable Interval	13

19.	Procedure when Results are not in Range	13

20.	Critical Alert Values	13

21.	Laboratory Clinical Interpretation	13

22.	Potential Source of Variability	14

23.	References	14

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CV | Coefficient of Variation |
| 5 | CRR | Clinical Reportable Range |
| 6 | EDTA | Ethylenediaminetetraacetic Acid |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | NIST | National Institute of Standards and Technology |
| 15 | pH | Potential of Hydrogen |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | SD | Standard Deviation |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TBIL | Total Bilirubin |
| 23 | TRF | Test Requisition Form |
| 24 | TDPL | TRUSTlab Diagnostics Private Limited |
| 25 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Total Bilirubin in serum and plasma samples using Beckman Coulter automated biochemistry analyser. This SOP also provides information on Test principles, clinical significance etc.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Total bilirubin is a blood test that measures the overall amount of bilirubin, a yellow pigment produced when red blood cells break down. It includes both direct (conjugated) and indirect (unconjugated) bilirubin.

- Test Principle

- A stabilized diazonium salt, 3, 5-dichlorophenyl diazonium tetrafluoroborate (DPD) reacts with conjugated bilirubin directly and with unconjugated bilirubin in the presence of an accelerator to form azobilirubin. The absorbance at 540 nm is proportional to the total bilirubin concentration.  A separate sample blank is performed to reduce endogenous serum interference.

Caffeine

Surfactant

Reaction Principle:

- Bilirubin + DPD	                                                  Azobilirubin

- Specific Performance Characteristics

Limit of Quantitation:

      - LLOQ: 0 mg/dl

      - ULOQ: 30 mg/dl

Analytical measurement range (AMR): 0-30 mg/dl

Clinical Reportable Range (CRR): 0-30 mg/dl

- Type of Sample

A minimum of 2ml of Serum/Plasma

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF/Grossly incomplete TRF/Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample acceptance/processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No other specific preparation required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxC 700 AU Chemistry Analyzer

Reagents

      - Concentrations

| - Reactive Ingredient | - Final Concentration |
| --- | --- |
| Caffeine | - 2.1 mmol/l |
| 3,5-dichlorophenyldiazonium tetrafluoroborate | - 0.31 mmol/l |
| Surfactant | - N/A |
| - Preservative | - N/A |

- The concentrations of the reactive components of the reagents shown on the kit label are the actual concentrations in the individual R1/R2 vials. The reagent composition which is shown in the Instructions for Use is the final concentration of these components in the reaction cuvette after addition of R1, Sample and R2.

      - Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2 - 8°C.

          - Once open, reagents stored on board the instrument are stable for 90 days.

          - Protect TBILC from light.

- Safety

For in vitro diagnostic use only. Do not pipette by mouth. Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

- Calibration Procedure

Calibrator Required:

      - System Calibrator

      - The calibrator value is traceable to the National Institute of Standards and Technology (NIST) Standard Reference Material (SRM) 937.

Calibration schedule:

      - Recalibrate the assay every 30 days.

      - Calibrate when there is a change in reagent lot number or significant shift in control values.

      - Calibrate when major preventative maintenance has been performed on the analyser or a critical part was replaced.

- Test Procedure

Refer to the appropriate Beckman Coulter AU Analyzer User Guide.

Sample Processing Using Racks:

      - Place Each sample in a rack (if volume is not sufficient transfer sample into a validated low volume container).

      - Open the top and/or the front cover of the Rack Input area.

      - Place the rack in the input area with the Rack ID barcode facing the back of the analyzer.

      - Do not leave any empty space between the racks.

      - Close the Rack Input area cover.

      - Select Start button.

To monitor sample status on the Analyzer, select - Result.

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1 and Level 2, Two levels of Quality Control

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation –- CMC EQAS

- Interferences & Cross Reactions

- Results of studies conducted to evaluate the susceptibility of the method to interference were as follows:

The criterion for no significant interference is recovery within 10% of the initial value.

Hemolysis: No significant interference up to 450 mg/dL of hemolysate.

Lipemia: No significant interference up to 1000 mg/dL of intralipid.

- Result Calculation

TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

| Age | Male (mg/dl) | Female (mg/dl) |
| --- | --- | --- |
| Adults | 0.3 - 1.2 | 0.3 - 1.2 |
| 3 – 5 Days | 1.5 - 12.0 | 1.5 - 12.0 |
| 1 – 2 Days | 3.4 - 11.5 | 3.4 - 11.5 |
| 0 – 1 Day | 1.4 - 8.7 | 1.4 - 8.7 |

- Reportable Interval

Analytical Measurement Range (AMR): 0-30 mg/dl

Clinical Reportable Range (CRR): ): 0-30 mg/dl

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- Newborn babies: >15 mg/dl

- Laboratory Clinical Interpretation

Increased Serum/Plasma levels were seen in the following conditions:

- Hemolytic jaundice

- Hepatocellular destructions (Viral, toxic, alcoholic)

- Biliary tract obstructions

- Physiological hyperbilirubinemia

- Gilberts disease

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Presence of fibrin in the sample may lead to erroneous results and hence centrifuge the sample only after complete clot formation.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition.

      - Product insert of Total Bilirubin, Beckman Coulter AU680 / AU 480 / DXC700.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).