SOP – 017

PROCEDURE FOR CHLORIDE TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	8

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	11

16.	Result Calculation	11

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BC | Biochemistry |
| 2 | °C | Degree Centigrade |
| 3 | GI | Gastrointestinal |
| 4 | ISE | Ion Selective Electrode |
| 5 | EDTA | Ethylenediaminetetraacetic Acid |
| 6 | EQAS | External Quality Assurance Scheme |
| 7 | QC | Quality Control |
| 8 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - This SOP describes the step-by-step tasks involved in performing the Chloride test.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- A chloride test measures the amount of chloride, an electrolyte, in the blood or urine. This test helps doctors diagnose and monitor conditions related to fluid balance, acid-base balance, and various medical conditions like kidney disease and heart failure.

- Test Principle

- Sensa Core’s ST-200 aQua Electrolyte Analyzer is completely automated, microprocessor-controlled electrolyte system that uses current ISE technology to make electrolyte measurements. The ST-200 aQua product measures various combinations of Chloride serum and urine.

- Method: Ion Selective Electrode (ISE) Direct method

- Specific Performance Characteristics

Analytical Sensitivity      :  The minimum detection limit by this kit is 25.0 mmol/l

Detection Limit                 :   25.0 – 200.0 mmol/l

- Type of Sample

- Only the specimens listed below were tested and found acceptable.

- Serum: Use serum free of hemolysis and gross lipemia, collected by standard venipuncture technique.

- Plasma: Use only lithium heparin.

- Urine: Collect 24 hour urine without additives.

- Patient Preparation

- Specimen should be collected using recommended procedures for collection of diagnostic blood specimen by venipuncture.

- Type of Container & Additives

- Plain red top vacutainer tube and Green Top (Lithium Heparin).

- Sterile container for Urine.

- Storage and stability for serum and plasma:

| - Temperature | - 15- 25°C | - 2-8°C | - -20°C |
| --- | --- | --- | --- |
| - Stability | - 2 weeks | - 2 weeks | - stable |

- Equipment & Reagents

Instruments

      - Sensa Core ST-200 aQua

Material

      - Ready to use

Reagent Storage & Stability

- Temperature: 5°C to 45°C.

- Safety

Process the sample on the same day within 3 hrs of collection.

Safety Precautions:

      - Handle all samples as potentially infectious.

      - Wear suitable gloves before handling any sample

      - Handle all reagents with care and avoid contact with eyes, mouth and skin.

      - Do not perform mouth pipetting.

      - Discard used reagents and samples as per disposal procedure.

- Calibration Procedure

Calibration frequency:

- Perform a full calibration

      - Every 24 hours

      - After 25 samples

      - After daily cleaning and maintenance

      - After changing the reagent bottles

      - After replacing any electrode

      - As required following quality control procedures.

- Test Procedure

    - Refer to Sensa Core’s ST-200 aQua Manual for ISE mode Operation and, Sensa Core Reference manual for estimation of electrolytes.

- Quality Control Procedure

Internal Quality Control

    - Quality Control Material Selection: Biorad control levels 1, 2 & 3 are used.

    - To ensure adequate quality control, normal and abnormal control with assayed values should be run as unknown samples in the following scenarios:

      - When a new reagent fluid pack is used.

      - After preventative maintenance is performed or a critical component is replaced.

    - Control results falling above the upper limit or below the lower limit of the established ranges indicate the assay may be out of control. The following corrective actions are recommended in such situations:

      - Repeat the same controls.

      - If the results are outside the limits, take fresh control serum and repeat the test.

      - If the results are still out of range, recalibrate and then repeat the test.

      - If the results are still out of range, perform a calibration with fresh reagent, then repeat the test.

      - If the results are still out of range, contact Technical Services or your local distributor.

External Quality Control

- EQAS participation – CMC EQAS on a monthly basis.

- Any outlier is noted, and Corrective and Preventive action is taken and is documented.

- Interferences & Cross Reactions

Haemolysis alters the values.

Hemoglobin does not interfere in the tested concentration range up to 1000 mg/dl (621 μmol/l) hemoglobin (approximate H index 1000).

- Result Calculation

N/A.

- Biological Reference Interval

- Serum/Plasma	-    94 - 110 mmol/l

- Urine (24 hr)	-    100 - 250 mmol/24 h

- Reportable Interval

- 25.0 – 200.0 mmol/l

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

      - < 80 mmol/l

      - > 120 mmol/l

- Laboratory Clinical Interpretation

      - Chloride increased in metabolic associated with prolonged diarrhea with loss of sodium bicarbonate.

      - Renal tubular with decreased excretion of hydrogen ions and decreased reabsorption of bicarbonate. Chloride may be decreased in respiratory alkalosis, drugs, some cases of hyperparathyroidism.

- Potential Source of Variability

- Haemolysis, improper calibration.

- References

- Sensa Core’s ST-200 aQua Potassium Kit Insert Catalog No. R22112-W9