SOP – 029

PROCEDURE FOR ESTRADIOL TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	10

8.	Type of Container & Additives	10

9.	Patient Preparation	10

10.	Equipment & Reagents	10

11.	Safety	11

12.	Calibration Procedure	11

13.	Test Procedure	12

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	15

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	16

22.	Potential Source of Variability	16

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 |  |  |
| 2 | AMR | Analytical Measurement Range |
| 3 | BC | Biochemistry |
| 4 | BSA | Bovine Serum Albumin |
| 5 | °C | Degree Centigrade |
| 6 | CA | Carbohydrate Antigen |
| 7 | CEA | Carcinoembryonic Antigen |
| 8 | CLIA | Chemiluminescence Immunoassay |
| 9 | CV | Coefficient of Variation |
| 10 | CRR | Clinical Reportable Range |
| 11 | EDTA | Ethylenediaminetetraacetic Acid |
| 12 | EQAS | External Quality Assurance Scheme |
| 13 | HOD | Head of Department |
| 14 | IQC | Internal Quality Control |
| 15 | LLOQ | Lower Limit of Quantitation |
| 16 | MAb | Monoclonal Antibodies |
| 17 | QC | Quality Control |
| 18 | QM | Quality Manager |
| 19 | SD | Standard Deviation |
| 20 | SST | Serum Separator Tube |
| 21 | TDPL | TRUSTlab Diagnostics Private Limited |
| 22 | ULOQ | Upper Limit of Quantitation |

- Objective

  - The purpose of this SOP is to describe the procedure for quantitative estimation of Estradiol in serum samples using Beckman Coulter DxI analyzer.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- An estradiol test measures the level of estradiol, a form of estrogen, in the blood. It's primarily used to assess reproductive health, including ovarian function, menstrual cycles, and fertility. The test can also be used to monitor menopause, evaluate hormone replacement therapy, and diagnose conditions like polycystic ovary syndrome (PCOS) or ovarian tumors.

- Test Principle

The Access Estradiol assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of estradiol levels in human serum and plasma using the Access Immunoassay Systems.

Estradiol (17β-estradiol, 1,3,5 (10)-Estratrien-3,17β-diol) is a natural estrogen with a molecular mass of 272.3 daltons. Most circulating estradiol is bound to protein. It is estimated that only 1-3% of estradiol is free (unbound). In non-pregnant women, estradiol is secreted by the ovary and the corpus luteum. The adrenals and testes (in men) are also believed to secrete minute amounts of estradiol. Estradiol levels are lowest at menses and into the early follicular phase and rise in the late follicular phase to a peak just prior to the hLH (human Luteinizing Hormone) surge, initiating ovulation. As the hLH peaks, the levels of estradiol decrease before rising again in the luteal phase. Endometrial growth is stimulated by estradiol and progesterone (secreted by the corpus luteum) in preparation for implantation of a fertilized egg. If conception does not occur, the secretion of estradiol and progesterone by the corpus luteum decreases, initiating menses.

- Levels of estradiol are used to monitor ovulatory status. Because estradiol levels reflect follicular maturation, the measurement of estradiol as cited in scientific literature has been used as a valuable tool in the assessment of sexual development, etiology of amenorrhea, causes of infertility and menopause. Abnormally high levels in males are indicative of feminizing syndromes such as gynecomastia.

- The Access Estradiol assay is a competitive binding immunoenzymatic assay. A sample is added to a reaction vessel with paramagnetic particles coated with goat anti-rabbit: rabbit anti-estradiol and a TRIS-buffered protein solution. After 20 minutes, estradiol alkaline phosphatase conjugate is added. Estradiol in the sample competes with the estradiol-alkaline phosphatase conjugate for binding sites on a limited amount of specific anti-estradiol antibody. Resulting antigen: antibody complexes are bound to the capture antibody on the solid-phase. After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos* 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is inversely proportional to the concentration of estradiol in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 20 pg/ml

      - ULOQ: 4800 pg/ml

Analytical measurement range (AMR): 20-4800 pg/ml

Sensitivity: The lowest detectable level of estradiol distinguishable from zero (Access Estradiol Calibrator S0) with 95% confidence is 20 pg/ml (73 pmol/l). This value is determined by processing a complete six-point calibration curve, controls, and 10 replicates of the zero calibrator in multiple assays. The analytical sensitivity value is interpolated from the curve at the point that is two standard deviations from the fitted zero calibrator signal.

Specificity: Serum samples containing up to 10 mg/dl (171 μmol/l) Bilirubin, hemolyzed samples up to 1 g/dl (10 g/l) hemoglobin and lipemic samples containing the equivalent of 1800 mg/dl (20.32 mmol/l) triglycerides do not affect the concentration of estradiol assayed utilizing a serum sample containing approximately 1089 pg/ml estradiol.

- Type of Sample

Serum is the recommended sample

      - Ambient (+15 to 25°C): 24 hours

      - Refrigerated (+2 to +8°C): 4 days

      - Frozen (-20°C): long term storage.

- Type of Container & Additives

- Gel/ Clot Activator tube.

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No other specific preparation required.

- Equipment & Reagents

Equipment

- Beckman DxI 800 CLIA analyzer

Reagents

- Access Estradiol Reagent Pack

- Cat. No. 33540: 100 determinations, 2 packs, 50 tests/pack

- Provided ready to use.

- Store upright and refrigerate packs at 2 to 10°C.

- Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

- Stable until the expiration date stated on the label when stored at 2 to 10°C.

- After initial use, the pack is stable at 2 to 10°C up to 28 days.

- Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

- If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

| R1a | Paramagnetic particles coated with goat anti-rabbit IgG: rabbit anti-estradiol in TRIS buffered saline, with bovine serum albumin (BSA), < 0.1% sodium azide. |
| --- | --- |
| R1b | TRIS, sodium chloride, protein (bovine, goat) and < 0.1% sodium azide. |
| R1c | Estradiol-alkaline phosphatase conjugate (bovine), protein (BSA, rabbit) < 0.1% sodium azide. |

- Safety

Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the AMH Reagent Kit on the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

Handling precautions are described in detail in the UniCel DxI System Operations Manual.

- Calibration Procedure

- An active calibration curve is required for all tests. For the Access CEA assay, calibration is required every 28 days. Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

- Test Procedure

- Refer to the appropriate system manuals and/or Help system for a specific description of installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

Procedure

      - Mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument. Do not invert open (punctured) packs, instead, mix reagents by swirling gently

      - Use ten (10) µL of sample for each determination in addition to the sample container and system dead volume. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

      - The system default unit of measure for sample results is ng/ml.

Handling of Repeat Test Results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever higher).

      - If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

    - Biorad: Level 1, Level 2 and Level 3, any two levels of Quality Control sera i.e. low, normal and high are run once daily

    - Preparation of Controls:

      - Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

      - Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

      - Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

      - Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

      - Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided).

      - Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”. Store the aliquots at -20°C.

    - Storage and Stability of Controls:

      - Unopened: Stable until the expiry date when stored at 2 to 8°C.

      - Reconstituted: Stable for 7 days when stored at 2 to 8°C; stable up to 30 days when stored at -20°C. Note: Aliquots are stored at the above temperatures as required.

    - Procedure

      - Run Quality Control according to TRUS lab-defined Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20°C, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8°C, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion

      - Place a portion of the material in a sample cup (approximately 350 µL).

      - Place the cup in the tray positions indicated by the screen Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levy Jennings chart and calculate the cumulative precision i.e. %CV.

    - Handling of QC failure:

      - Verify the following points and if required, consult with HOD/QM:

          - Review the environmental conditions and recent events and look for any deviation.

          - Check on-board reagents (for stability, deposits, fungal growth) reference fluid and controls.

          - Check the quality of the sterile water for injections used for the preparation of controls and calibrators, contaminants, i.e. bacterial growth may contribute to inaccurate results.

          - Check reaction temperature.

          - Check the expiry date of kit and contents.

          - Check instrument settings and light source.

          - If there are any control failures, take appropriate corrective and preventive actions as per TRUSTlab’s Quality Process MSP.

          - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

          - Take appropriate corrective actions and re-run the QC.

          - Contact application support/service support from vendor if required corrective action fails to yield acceptable results.

    - Repeat and Recheck criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation – Metropolis EQAS

- Interferences & Cross Reactions

Hemoglobin, triglycerides, bilirubin, human serum albumin and rheumatoid factor, tested up to the following concentrations, respectively, 500 mg/dl, 1,800 mg/dl, 30 mg/dl, 5 g/dl and 500 IU/ml, do not interfere with the Access CEA assay.

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

Male : <20 – 47 pg/ml

Female: Post Menopausal - <20 – 40 pg/ml

Non-Pregnant : Mid-follicular phase – 27 -122 pg/ml

Mid-luteal Phase – 49 – 291 pg/ml

- Peri-ovulatry phase – 95 – 433 pg/ml

- Reportable Interval

- AMR: 20 to 4800 ng/ml

- CRR > 10000 ng/ml

- Dilution Procedure:

- If a sample contains more than the stated value of the highest Access CEA Calibrator (S5), report the result as greater than that value (i.e. > 1000 ng/ml). Alternatively, dilute one volume of the sample with nine volumes of Access CEA Calibrator S0 (zero) or Access CEA Diluent. After assaying the diluted sample, multiply the obtained value by the dilution factor, ten (10). If a sample contains less than the lower limit of detection for the assay, report the result as less than the lower limit of detection (< 0.1 ng/ml). Refer to the appropriate system manuals and/or Help system for instructions on how to enter a sample dilution in a test request.

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

      - N/A.

- Laboratory Clinical Interpretation

The measurement of serum CEA has shown substantial benefit in the prognosis and management of patients with malignant diseases, especially colorectal cancer. Serial measurements can be used to monitor patients for progression, regression or recurrence of cancer following treatment. A persistent elevation of CEA following therapeutic or surgical intervention signals residual disease or recurrence, whereas decreasing levels to within the normal range is indicative of successful intervention.

CEA is also elevated in the serum of patients with non-malignant diseases and in heavy smokers, therefore CEA should not be used in the diagnosis of cancer or for screening asymptomatic patients.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE as it indicates contamination of the reagent.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Access DXI CEA Kit insert

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)