SOP – 030

PROCEDURE FOR

FERRITIN TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	11

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	14

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CV | Coefficient of Variation |
| 5 | CRR | Clinical Reportable Range |
| 6 | EDTA | Ethylenediaminetetraacetic Acid |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | pH | Potential of Hydrogen |
| 15 | QC | Quality Control |
| 16 | QM | Quality Manager |
| 17 | QNS | Quantity Not Sufficient |
| 18 | SD | Standard Deviation |
| 19 | SOP | Standard Operating Procedure |
| 20 | SST | Serum Separator Tube |
| 21 | TRF | Test Requisition Form |
| 22 | TDPL | TRUSTlab Diagnostics Private Limited |
| 23 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Ferritin in serum using Beckman Coulter DxI 800. This SOP also provides information about the test principle, clinical significance and other test specific details.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Ferritin is a protein that stores iron within the body. It acts as a "pantry" for iron, releasing it when the body needs it to make red blood cells and other vital functions. A ferritin test measures the amount of ferritin in the blood, which can indirectly indicate how much iron is stored in the body.

- Test Principle

The Access Ferritin assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of ferritin levels in human serum and plasma (heparin) using the Access Immunoassay Systems.

- Specific Performance Characteristics

Accuracy: 95%

Limit of Quantitation:

      - LLOQ: 8.0 ng/ml

      - ULOQ: 450 ng/ml

Analytical measurement range (AMR): 8 - 450 ng/ml

Clinical Reportable Range (CRR): 8 - 450 ng/ml

- Type of Sample

A minimum of 2 ml of Serum (preferred).

Collect blood sample by venipuncture

For serum, allow samples to clot adequately before centrifugation.

Keep tubes stoppered at all times.

Physically separate serum or plasma from contact with cells as soon as possible.

Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - Follow blood collection tube manufacturer’s recommendations for centrifugation

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot Activator – Red top Vacutainer

Heparin tube- Green top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800

Reagents

    - Access Ferritin Reagent Pack Cat. No. 33020: 100 determinations, 2 packs, 50 tests/pack.

    - Provided ready to use.

    - Store upright and refrigerate at 2 to 10°C.

    - Refrigerate at 2 to 10°C for a minimum of two hours before use on the instrument.

    - Stable until the expiration date stated on the label when stored at 2 to 10°C.

    - Stable at 2 to 10°C for 28 days after initial use.

    - Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

    - If the reagent pack is damaged (i.e., broken elastomer), discard the pack.

    - All antisera are polyclonal unless otherwise indicated.

| R1a: | Paramagnetic particles coated with goat anti-mouse IgG: mouse monoclonal anti-ferritin complexes suspended in TRIS buffered saline, with a surfactant, bovine serum albumin (BSA), < 0.1% sodium azide, and 0.1% ProClin 300. |
| --- | --- |
| R1b: | Goat anti-ferritin-alkaline phosphatase (bovine) conjugate in TRIS buffered saline, with a surfactant BSA, protein (goat, mouse), < 0.1% sodium azide, and 0.1% ProClin 300. |

    - CAUTION:

- Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

    - Do not use reagents beyond the expiration date

- Safety

For in vitro diagnostic use only.

Use this reagent only for the DXI 800.

Do not pipette by mouth.

Do not ingest. Harmful if swallowed

Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

Used cuvettes contain human body fluids, it should be handled with appropriate care to avoid skin contact.

Reagent packs: Sealed cartridges are stable until the Expiry date.

- Calibration Procedure

Calibrator Required:

- Manufacturer supplied 6 Point Calibrator.

Calibration Schedule:

- An active calibration curve is required for all tests. Calibration is required every 28 days. Calibration is also needed when the following occur:

      - When there is a change in reagent lot number or significant shift in control values.

      - When major preventative maintenance has been performed on the analyzer or a critical part was replaced.

      - When quality control results are found to be consistently outside acceptable range.

- Test Procedure

Refer to the appropriate Beckman Coulter DxI 800 User Guide.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use ten (10) μL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) μL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate manual system and/or Help system for the minimum sample volume required.

The system default unit of measure for sample results is ng/ml. To change sample reporting units to the International System of Units (SI units) µg/l, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply ng/ml by multiplication factor

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Analyzer generated results are reviewed by the operator & only clinically valid results are saved for approval by personnel authorized to release the reports.

Invalid results are reviewed for repeat & recheck criteria.

The results are then evaluated & reported by the personnel authorized to release the reports.

- Quality Control Procedure

Internal Quality Control

- Biorad: Level 1, 2 and 3. Any two levels of Quality Control sera i.e. low, normal and high are run as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity   (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS Participation – Metropolis EQAS

- Interferences & Cross Reactions

Samples containing up to 5 mg/dl (86 μmol/l) bilirubin, lipemic samples containing the equivalent of 900 mg/dl (10.16 mmol/l) triglycerides, and hemolyzed samples containing up to 300 mg/dl (3 g/l) hemoglobin do not affect the concentration of ferritin assayed. Grossly hemolyzed samples should not be used in that ferritin can be released from lysed red cells.

Samples containing 5-9 g/dl (50-90 g/l) albumin do not affect the concentration of ferritin assayed.

The antibodies used in this kit were raised against liver ferritin. Reactivity to spleen ferritin, as demonstrated by spiking recovery in a serum sample, is equivalent to that of liver ferritin.

The following table describes the cross-reactivity of the assay with common chemotherapeutic agents and other potential interferents

- Result Calculation

Ferritin concentrations were measured in 113 serum samples from apparently healthy male and female subjects using the Access Ferritin assay. The results were as follows.

| Subject Category | n |  | Geometric Mean (ng/ml, μg/l) | 95% Range (ng/ml, μg/l) |
| --- | --- | --- | --- | --- |
| Males |  | 49 | 105.6 | 23.9-336.2 |
| Females |  | 64 | 51.4 | 11.0-306.8 |

Interpret the results of this test in conjunction with the patient’s clinical presentation

- Biological Reference Interval

| Age | Range |
| --- | --- |
| - | 0.29–1000 ng/ml |

- Reportable Interval

Analytical Measurement Range (AMR): 8.0 - 450 ng/ml

Clinical Reportable Range (CRR): 8.0 - 450 ng/ml

- Procedure when Results are not in Range

The Ferritin assay exhibits total imprecision of less than 10% across the assay range. One study, using commercially available human serum-based control material generating 20 assays, two replicates per assay, over 20 days provided the following data, analyzed via analysis of variance (ANOVA):

| Sample | Grand Mean (n=40) (ng/ml) | Within Run (%CV) | Total Imprecision (%CV) |
| --- | --- | --- | --- |
| Low | 37.2 | 2.6 | 4.1 |
| Medium | 118.9 | 3.6 | 4.3 |
| High | 311.8 | 3.9 | 6.3 |

- The Dil-Fer assay exhibits total imprecision of less than 20% across the range of the Dil-Fer assay.

- Critical Alert Values

> 1000 ng/ml

- Laboratory Clinical Interpretation

Ferritin plays a significant role in the absorption storage and release of Iron.

Stored Iron represents about 25 % of total Iron in the body, and most of this Iron is stored as Ferritin. Ferritin is found in many body cells, especially in the Liver, Spleen Bone Marrow and in Reticuloendothelial cells.

Ferritin is a compound composed of Iron molecules bound to Apoferritin, a protein shell.

Ferritin predicts and monitors Iron deficiency; differentiate Iron deficiency from chronic disease as cause of Anaemia.

Increased levels of Ferritin are found in acute and chronic Liver disease, alcoholism, malignancies, Hyperthyroidism and Iron load.

Decreased levels of Ferritin are found in Iron deficiency and Haemodialysis.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of Ferritin, Beckman Coulter DxI 800

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)