SOP - 035

PROCEDURE FOR HBA1C TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	10

7.	Type of Sample	10

8.	Patient Preparation	10

9.	Type of Container & Additives	10

10.	Equipment & Reagents	10

11.	Safety	12

12.	Calibration Procedure	12

13.	Test Procedure	13

14.	Quality Control Procedure	17

15.	Interfering Substances	19

16.	Result Calculation	20

17.	Biological Reference Interval	20

18.	Reportable Interval	20

19.	Procedure when Results are not in Range	20

20.	Critical Alert Values	21

21.	Laboratory Clinical Interpretation	21

22.	Potential Source of Variability	22

23.	References	22

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BC | Biochemistry |
| 2 | DCCT | Diabetes Control and Complications Trial |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HbA1c | Glycated Haemoglobin |
| 6 | HPLC | High Performance Liquid Chromatography |
| 7 | LLOQ | Lower Limit of Quantitation |
| 8 | NGSP | National Glycohemoglobin Standardization Program |
| 9 | QC | Quality Control |
| 10 | rpm | Rotations per minute |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |
| 12 | ULN | Upper Limit of Normal |
| 13 | ULOQ | Upper Limit of Quantitation |
| 14 | WHO | World Health Organisation |

- Objective

  - This SOP describes the step-by-step tasks involved in the HbA1C testing process.

- Scope

This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Hemoglobin A1c measurements are used in the clinical management of diabetes to assess the long-term efficacy of diabetic control. Although a fasting blood glucose measurement gives the clinician information about the patient’s status over the last twelve hours, the stable HbA1c offers a more accurate indication of the patient’s long-term diabetic control over the last two to three months.

- Test Principle

- Glycohemoglobin is a general term for hemoglobin-glucose complexes in which glucose is bound to the alpha and beta chains of hemoglobin. The most quantitatively prevalent complex is called HbA1c, in which glucose binds to the N-terminus of the beta chain of HbA. HbA1c is nonenzymatically synthesized in two steps: The glucose aldehyde group and the free amino group on the valine in the N terminus of the hemoglobin beta chain react to form the Schiff base, aldimine (also known as labile HbA1c or LA1c). A stable ketoamine form of the hemoglobin complex (SA1c) is then produced by a reaction known as Amadori rearrangement. The level of LA1c changes rapidly in response to changes in blood glucose concentration. However, the level of the SA1c does not fluctuate significantly in response to physiological factors. Consequently, the SA1c measurement provides a better indication of the average glucose level over the previous two to three months (the average red blood cell life span). The Tosoh Automated Glycohemoglobin Analyzer HLC-723G8 uses non-porous ion exchange high performance liquid chromatography (HPLC) for rapid, accurate and precise separation of the stable form of HbA1c from other hemoglobin fractions. Analysis is carried out without off-line specimen pretreatment or interference from Schiff base. The analyzer dilutes the whole blood specimen (3 μL) with Hemolysis & Wash Solution and then injects a small volume of this specimen onto the TSKgel G8 Variant HSi column. Separation is achieved by utilizing differences in ionic interactions between the cation exchange group on the column resin surface and the hemoglobin components. The hemoglobin fractions (designated as A1a, A1b, F, LA1c+, SA1c, A0, and H-V0, H-V1, HV2) are subsequently removed from the column by performing a step-wise elution using the varied salt concentrations in the Variant Elution Buffers HSi 1, 2, and 3. The time from injection of the sample to the time the specific peak elutes off the column is called retention time. The Tosoh Automated Glycohemoglobin Analyzer HLC-723G8 software has been written so that each of the expected fractions has a window of acceptable retention times. If the designated peak falls within the expected window, the chromatogram peaks will be properly identified. When a peak elutes at a retention time not within a specified window, an Glycohemoglobin in Whole Blood NHANES 2011-2012 4 unknown peak (P00) results. If more than one peak elutes at times not specified by the software windows, each is given a sequential P0x title. In order to keep the peaks within their appropriate windows, it may be necessary to change how fast or slow the buffers are moving through the system by changing the pump flow rate. The separated hemoglobin components pass through the LED photometer flow cell where the analyzer measures changes in absorbance at 415 nm. The analyzer integrates and reduces the raw data and then calculates the relative percentages of each hemoglobin fraction. The Total Area of the SA1c is divided by the sum of the total areas of all peaks up to and including the A0 to obtain a raw SA1c percentage. This uncorrected result is substituted as the “x” value in the linear regression formula determined during calibration. The analyzer prints the final numerical results and plots a chromatogram showing changes in absorbance versus retention time for each peak fraction (chap 1). The Tosoh G8 Automated HPLC Analyzer – HbA1c Variant Analysis Mode is certified by the National Glycohemoglobin Standardization Program (NGSP). The final reportable result is traceable to the Diabetes Control and Complications Trial (DCCT).

- Specific Performance Characteristics

Accuracy: 90%

Limit of Quantitation:

      - LLOQ: < 2.7%

      - ULOQ: > 19.7%

Limit of detection: <3.0% to >19 .0 %

Analytical Measurement Range (AMR): 2.7% to 19.7%

Clinical Reportable Range (CRR):  2.7% to 19.7%

Sensitivity: 90%

Specificity: 90%

- Type of Sample

- Whole Blood

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 3 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

Instruments

      - G8 Automated HPLC Analyzer

      - Labquake Rotator

      - Auto Dilutor

Material

      - TSKgel Glyco HSi Variant Column

      - Filter element, 5/pkg

      - Thermal paper for G7, 10 roll/box

      - DIAMAT HbA1c Sample Preparation Kit

      - Sample Preparation Vials, 100/kit

      - Capillaries, one glass dispenser vial containing 100 sodiumheparinized capillary tubes (5 μl)

      - Capillary tube holder, one holder for manipulating 5 μl capillary tubes. Reorder box of 20 holders

      - Labels, 4 sheets of 25 blank labels each

Reagent Preparation

      - Tosoh Hemoglobin A1c Calibrator Set

      - Calibrator 1 (approximately 5.5%) 5 x 4 ml

      - Calibrator 2 (approximately 10.5%) 5 x 4 ml

      - Buffered human red blood cells, 2 mg/ml

      - Human hemoglobin

      - Tosoh Hemoglobin A1c Control

          - Normal 4 x 0.25 ml

          - Abnormal 4 x 0.25 ml

- Unopened and stored at 2-8°C, the Tosoh Hemoglobin A1c calibrator and control sets are stable until the expiration date printed on the label. After reconstitution, calibrators are stable for one week when stored at 2-8°C. Refer to the control package insert for stability data.

- TSKgel G8 Variant HSi column should be stored at 4-15°C in a cool location away from direct sunlight. The column is stable until the expiration date printed on the label. Columns are warranted for 2500 injections. ARDL has shown that column may last up to 3500 injections or more.

- 021956 G8 Variant Elution Buffer HSi No.1 (S) 1 x 800 ml

- 021957 G8 Variant Elution Buffer HSi No.2 (S) 1 x 800 ml

- 021958 G8 Variant Elution Buffer HSi No.3 (S) 1 x 800 ml

- Unopened G8 Variant Elution Buffers HSi (S) 1, 2, and 3 are stable until the expiration date printed on the label. After opening, Elution Buffers are stable for three months. Store at 4-30°C

- 018431US HSi Hemolysis & Wash Solution (L) 1 x 2000 ml Unopened Hemolysis & Wash Solution is stable until the expiration date printed on the label. After opening, Hemolysis & Wash Solution is stable for three months. Store at 4-30°C

- 021600 Filter Element 5/pkg 018581

- Sample Cups 1000/pkg

- 019563 Thermal Paper 10 rolls/pkg

- 018723Supply Line Filters for Buffer Lines 1/pkg

- 019500 Sampling Needle Assembly 1 each

- Safety

For in vitro diagnostic use.

This kit should only be used by adequately trained personnel.

All equipment required for this kit must be serviced by qualified service technicians.

This kit contains reagents manufactured from sheep blood components. The sheep are from a closed herd located in the USA. All recommended precautions for the handling of blood derivatives should be observed. Please refer to the U.S. Department of Health and Human Services publication "Biosafety in Microbiological and Biomedical Laboratories" or any other local or national regulation.

Handle all patient specimens as potentially infectious.

All the reagents may be irritating.

Disposal of all waste should be in accordance with local regulations.

Sodium azide can react with copper and lead plumbing to form explosive metal azides. At time of disposal, flush reagents with a large volume of water to prevent the buildup of azides.

- Calibration Procedure

HbA1c Calibrator Set: Calibrator 1 (5 x 4 ml) and Calibrator 2 (5 x 4 ml). Part # 018767, Tosoh Medics, Inc. Buffered human red blood cells, 2 mg/ml human hemoglobin, and 0.5 mM EDTA as preservative. Unreconstituted calibrator set is stable stored at 4-8°C until expiration date printed on label.

Reconstitute Calibrators 1 and 2 by adding 4 ml Milli-Q water to each vial then mix gently by inversion. Record dates of reconstitution and expiration on vial labels, then promptly store upright at 4-8°C. Always return calibrators promptly to refrigerator--do not leave vials at room temperature for an extended period.

Calibrator Lot Validation:

    - Each new lot of calibrators must be evaluated against the current lot prior to being put into use. (Evaluation against whole blood calibrators may be performed as needed – see Note 1). Analyze each level in duplicate within the same run over a period of two to three days (include both instruments) to verify that manufacturer-assigned values are valid. First, calibrate the run using the current lot of calibrators and analyze the controls. Analyze the new lot calibrators as unknowns immediately after the controls, running each level in duplicate. Record the values obtained including analyses from both instruments. When the tally is complete, calculate a Mean to confirm the assigned bottle value or to determine new assigned values. Prior to analyzing patient specimens, verify that analysis of current lot of controls against the new lot calibrators produces results that fall within established control ranges. New lots of the calibrator may be evaluated as necessary against whole blood calibrators obtained from the NGSP CPRL at the University of Missouri (UMO calibrators). Perform this procedure when validation of new lot calibrators against current lot calibration does not confirm manufacturer-assigned values.

      - Obtain aliquots of UMO whole blood calibrators and UMO controls. First calibrate both instruments with UMO calibrators using their assigned values. Analyze the UMO controls, the current lot of inhouse controls and both levels of new lot calibrator in duplicate as unknowns. Verify that the controls fall within their respective QC limits (preferably within 1 SD). Evaluate the results of the new lot calibrators against their manufacturer-assigned values.

      - Next, calibrate both instruments with the new lot of Tosoh calibrators using the manufacturer-assigned values. Analyze the UMO controls and in-house controls and verify that they fall within their respective QC limits (preferably within 1 SD). If control results are acceptable, the assigned values may be used. If controls do not fall within established ranges, repeat analysis of new lot calibrators against UMO calibrators to establish new assigned values.

      - Additionally, the most recent set of NGSP Monthly Monitoring samples may be thawed and analyzed on both instruments and the results compared with results obtained using the current lot calibrators.

- Test Procedure

Daily checks before starting the test:

      - Check waste containers. If necessary, empty and add 5% bleach.

      - Both analyzers are currently pre-programmed to warm-up at 7:00 AM Monday through Friday. If necessary to manually initiate warmup, press the POWER key located on the operation panel to switch on the analyzer. (See information in the Equipment manual).

      - Check for leaks. During warm-up, check for leaks, especially at filter and column connections. If a leak is found, tighten the connections.

      - Verify pressure is +/- 4 MPa of value indicated on the column inspection report and record. Press the Down‟ arrow to display the second Main screen. Verify that pressure stabilizes column value +/-. Record pressure in maintenance log. If pressure is greater than 4 MPa of assigned value, replace filter element. If pressure is less than 4MPa of assigned value, check for leaks and tighten fittings. Note: Low pressure during pumping may indicate air in the system. Check the lines connecting the buffers to the analyzer. If air is observed in the line, see procedure for removing air from buffer lines in operator’s manual. If the pressure does not rise during pumping, air may be present on the outlet side of the pump (see procedure for removing air from the pump outlet in operator’s manual).

      - Record the instrument temperature. The temperature should be 25.0°C ± 1.0°C.

      - Check F/C (filter and column) counts. Replace filter element at or before 400 injections (see procedure for replacing the filter element. Replace column as indicated by appearance of chromatograms. Column is guaranteed for 2500 counts; replace as necessary (ARDL routinely utilizes column to 3000- 3500 if chromatograms and retention times are acceptable).

      - Check Smart Media card space. Format a card, if necessary, before starting. (See operator’s manual for formatting instructions.) Press MENU on the MAIN screen then CARD to check remaining disk storage space. The CARD screen shows how full the CARD is in the upper left corner. A single CARD holds up to 12,000 sets of assay results (including chromatograms). Return to Main screen, which should display “AUTO SAVE: YES”.

      - Check buffers and hemolysis-wash solution for sufficient volume. (See procedure for installing buffers.)

      - Check printer paper. A single roll is enough for ~350 results. (See procedure for installing printer paper in operator’s manual.)

      - Print the parameters. From the MAIN screen select MENU. Press UTILITY and PARAM PRINT.

Reagent and Column Preparation and Installation

    - Installing G8 Variant Elution Buffer His

      - Press the STOP key to put system in STAND-BY status.

      - Remove buffer containers to be replaced.

      - Break the seal on the storage caps of the new buffer containers, leaving the caps in place.

      - Carefully place the capped buffer bags on the rack on the instrument. Verify that the buffer bag is supported on the rack by the octagonal shaped lip at the base of the threads.

      - For each reagent, remove the cap, carefully squeeze the buffer bag by hand to minimize air pockets, and then place the appropriate color-coded tubing into the corresponding bag. Ensure that the end of the tubing is touching the bottom of the container.

      - Securely tighten reagent caps.

      - From the MAIN screen, select MAINTE, then REAGENT CHANGE. Once the buffers have been correctly installed, select the appropriate buffer, and then press CHANGE.

      - Record reagent lot and date in use on daily protocol page. If the lot number changes, record on reagent/column lot change sheet in GX Documentation binder.

    - Installing Hemolysis & Wash Solution

      - Press the STOP key to put the system in STAND-BY status.

      - Remove the empty container.

      - Break the seal on the storage cap of the new container, leaving the cap in place.

      - Place the Hemolysis & Wash Solution bottle on the bench at the left side of the instrument, beside the H/W Solution port. Glycohemoglobin in Whole Blood NHANES 2011-2012 13

      - Remove the cap, then place the Hemolysis & Wash Solution tubing into the container. Ensure that the end of the tubing is touching the bottom of the container.

      - Securely tighten the reagent cap.

      - From the MAIN screen, select MAINTE, then REAGENT CHANGE. Once the Hemolysis & Wash Solution has been correctly installed, press H/W, then press CHANGE.

      - Record reagent lot and date in use on daily protocol page. If lot number changes, record on reagent/column lot change sheet in GX Documentation binder.

    - Installing TSK gel G8 Column

      - Press the STOP key to put the system in STAND-BY status.

      - Unscrew column fittings and remove used column.

      - Remove protective plugs from new column. Do not discard the plugs, as they are needed for storage.

      - Verify that the column master lot matches the Elution buffer lot.

      - Check flow direction - the arrow on the column should point to the left as it is placed on the instrument.

      - Slide the inlet tubing until it extends ¼ inch past each end fitting.

      - Connect the tubing to the inlet (right) side of the column. Take special care that buffer coming from the tube does not spill onto the analyzer unit by holding absorbent paper or gauze at the outlet end of the column when priming.

      - From the second page of the MAIN screen, press PUMP FLOW to start pumping the buffer. When buffer solution flows from the outlet (left) side of the column, press PUMP FLOW again to stop the pump. 9. Connect the tubing to the outlet side of the column.

      - Press PUMP FLOW to start pumping and check for leaks. The pressure should rise to the pressure level that is indicated on the column inspection report + 4 MPa. If leaks occur, tighten the fittings.

      - When the pressure reaches a steady state, press PUMP FLOW again to stop pumping.

      - From the MAIN screen, select MAINTE, then REAGENT CHANGE. Press COLUMN RESET to set column count back to zero.

      - Before calibrating the newly installed column, analyze at least three whole blood samples to prime the column. Calibrate the system and 5 duplicate specimens and controls.

      - Record column lot and date in use on daily protocol page. If lot number changes, record on reagent/column lot change sheet in GX Documentation binder. Also, record column change on the daily maintenance chart.

Procedure for Processing Patient Samples

    - Whole Blood: Ensure that the stopper on the tube is properly seated and that the barcode label is vertically aligned; re-affix barcode label vertically on tube if necessary. Mix each patient specimen several times by gentle inversion (or place briefly on a rotator). Then place the specimen tube in the rack in order from left to right according to its rack and position number as recorded on the protocol page. Align its barcoded label so that it faces the barcode reader (i.e., facing away from you in the rack as it’s loaded on the instrument).

- Note: Blood cells will begin to settle out as the tubes sit on the instrument waiting to be measured. This cell sedimentation over a period of approximately 5 hours does not affect the HbA1c result.

    - Low-volume samples (less than 1.0 ml whole blood in tube): Using the autodilutor, prepare hemolysates. Wipe the probe tip after drawing up sample and again after dispensing into labeled sample vial. Place the sample vial in the rack. Minimum dilution volume dispensed or pipetted into a sample vial is 150 μl.

    - HbA1c Sample Preparation Vials. Remove caps prior to sampling! Place the Prep vial in the rack using an adapter tube.

    - Place racks in ascending order into the sample loader. Place an empty rack after the last rack to be processed. Press START key to begin analysis. The racks will be moved automatically along the sample loader. The analyzer will prime the fluid lines with buffer, then analyze samples at 1.6-minute intervals. The analyzer aspirates 4 μl of whole blood. The whole blood sample is automatically diluted at approximately 1:200 and introduced into the assay line. When the assay begins, the sample rack is transferred, and continuous sampling starts and continues until an empty rack on the loader is detected.

    - IMPORTANT: Keep tubes in the sample rack until the whole rack is processed and printed reports are available and have been reviewed.

    - Processing automatically stops when the analyzer detects an empty rack (or 10 sequential empty spaces). When measurement ends, the analyzer will enter the WASH mode in which it washes the column by pumping buffer for 3 minutes. Once it has started, always allow the WASH sequence to go to completion! If washing is insufficient, column lifespan will be reduced and the result for the next sample could be affected.

    - If there is no further input from the operation panel while the analyzer is in STAND-BY mode, after 2 hours (Off Time setting), the analyzer shuts itself off automatically.

- Quality Control Procedure

Internal Quality Control

    - Two levels of glycated hemoglobin control (Normal and High) are analyzed in duplicate (or more) with each batch. Controls must also be analyzed at the start and end of every day, after calibration, maintenance, and as needed for troubleshooting. Controls are prepared from whole blood drawn from a normal (Normal) and a diabetic (High) individual (See paragraph below for preparation instructions). Stable indefinitely stored at –70°C. Controls should be diluted with Hemolysis & Wash Solution to obtain a Total Area on the chromatogram in the range of 500- 4000. If the value of one or more control specimens is out of the acceptable range, recalibrate the system and reanalyze the controls before testing patient sample.

      - Run Quality Control according to the protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion

      - Place a portion of material in a sample cup (approximately 500 µL) .

      - Place the cup in the positions indicated by the screen Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levy Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure

- Verify the following points and if required, consult with   Chief of Lab /Quality Manager:

      - Review the environmental conditions and recent events and look for any deviation.

      - Check Onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for the injection used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check expiry date of kit and contents.

      - Check instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per TRUSTlab’s Quality Process MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

- Contact application support/service support from vendor if required corrective action fails to yield acceptable results.

External Quality Control

- EQAS participation – CMC EQAS

- Interfering Substances

Variant hemoglobin interference studies were also conducted on diabetic and nondiabetic samples spiked with a high human specimen for HbF, and commercially available control material for HbAE, HbAD, HbAS, and HbAC. HbAE is not distinguished from the other peaks, and if present will interfere with sA1c measurement.

No interference was observed for HbF up to a concentration of 10%.

No interference was observed for HbAD, HbAS(approximate retention time 1.2) and HbAC (approximate retention time 1.3) up to a concentration of 45% since these variants elute after the A0 peak.

Interference studies were conducted with diabetic and non-diabetic samples. They were spiked with increasing amounts of glucose, bilirubin, triglycerides, sodium cyanate, acetylsalicylic acid, acetaldehyde and EDTA. Interference was determined as a variance greater than the (assigned value x 1.00 +/- 5%) Labile A1c, as indicated by glucose concentrations up to 1000 mg/dl does not interfere with the assay.

Icterus, as indicated by free and conjugated bilirubin concentrations up to 20 mg/dl does not interfere with the assay.

Lipoma, as indicated by triglyceride concentrations up to 1000 mg/dl do not interfere with the assay. Concentrations of up to 25 mg/dl sodium cyanate (carbamylated Hb) do not interfere with the assay. Concentrations of up to 25 mg/dl acetaldehyde (alcohol) do not interfere with the assay.

Concentrations of up to 50 mg/dl of acetylsalicylic acid (aspirin) which forms acetylated Hb, do not interfere with the assay.

Total Area

      - Dilution studies demonstrate that the assay is linear from a Total Area of 500 to 4000.

      - Optimum Total Area is 700 to 3000. Do not report an A1C result if the total area is less than 500 or greater than 4000.

      - For areas greater than 4000, a dilution can be made using a smaller sample size or more diluent to bring the area within acceptable range, and reanalyzed. For areas less than 500, a dilution can be made using a larger sample size or concentration by centrifugation and making a ‘normal’ dilution, and reanalyzed.

Abnormal Red Cell Survival

- The life span of red blood cells is shortened in patients with hemolytic anemias, and the actual life span depends upon the severity of the anemia. As a consequence, specimens from such patients may exhibit decreased glycohemoglobin levels compared to patients with normal red cell life span. The life span of red blood cells is lengthened in polycythemia or post-splenectomy patients. Specimens from such patients may exhibit increased glycohemoglobin levels.

- Result Calculation

N/A.

- Biological Reference Interval

- 4.3 – 6.0 % (DCCT/EDIC normal range

- Reportable Interval

Analytical Measurement Range (AMR): 2.7% to 19.7%

Clinical Reportable Range (CRR):  2.7% to 19.7%

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find route cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status is fine. Next consult with Clinician for clinical significance.

- If required by the doctor, the sample will be rerun with appropriate dilution methods.

- Critical Alert Values

- > 6.5%

- Laboratory Clinical Interpretation

- Reference Range as per American Diabetes Association (ADA):

| REFERENCE GROUP | HBA1C IN % |
| --- | --- |
| Non-Diabetic Adults >18 Years | < 5.7 |
| At Risk (Prediabetes) | 5.7 – 6.4 |
| Diagnosing Diabetes | ≥ 6.5 |
| Diabetics |  |
| · Excellent Control | 6 – 7 |
| · Fair To Good Control | 7 – 8 |
| · Unsatisfactory Control | 8 - 10 |
| · Poor Control | > 10 |

- Note: Dietary preparation or fasting is not required

      - A1C test should be performed at least two times a year in patients who are meeting treatment goals (and who have stable glycemic control).

      - Lowering A1C to below or around 7% has been shown to reduce microvascular and neuropathic complications of type 1 and type 2 diabetes. When mean annual HbA1c is < 1.1 times ULN (upper limit of normal), renal and retinal complications are rare, but complications occur in > 70% of cases when HbA1c is > 1.7 times ULN.

      - Falsely low HbA1c (below 4%) may be observed in patients with clinical conditions that shorten erythrocyte life span or decrease mean erythrocyte age. HbA1c may not accurately reflect glycemic control when clinical conditions that affect erythrocyte survival are present. Fructosamine may be used as an alternate measurement of glycemic control.

- Potential Source of Variability

- N/A.

- References

      - TOSOH Kit Insert

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)