SOP – 043

PROCEDURE FOR

LH TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	13

16.	Result Calculation	14

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LH | Luteinizing Hormone |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MSP | Management System Procedure |
| 14 | N/A | Not Applicable/Available |
| 15 | pH | Potential of Hydrogen |
| 16 | QC | Quality Control |
| 17 | QM | Quality Manager |
| 18 | QNS | Quantity Not Sufficient |
| 19 | SD | Standard Deviation |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TRF | Test Requisition Form |
| 23 | TDPL | TRUSTlab Diagnostics Private Limited |
| 24 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Luteinizing Hormone (LH) in serum and plasma samples using Beckman Coulter CLIA analyser. This SOP also provides information about the test principle, clinical significance etc.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Luteinizing hormone (LH) is a protein hormone produced by the pituitary gland. It plays a crucial role in both male and female reproductive systems. In men, LH stimulates testosterone production by the testes. In women, LH triggers ovulation, steroid hormone production in the ovaries, and progesterone release after ovulation.

- Test Principle

The DXI hLH assay is a sequential two-step immunoenzymatic (“sandwich”) assay. Sample is added to a reaction vessel, along with paramagnetic particles coated with goat anti-mouse: mouse anti-hLH complexes and TRIS buffered saline with protein. The hLH binds to the immobilized mouse anti-hLH on the solid phase. Materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Alkaline phosphatase conjugated goat anti-hLH is then added, which binds to the previously bound hLH on the particles. A second separation and wash step removes unbound conjugate. Then, the chemiluminescent substrate Lumi-Phos 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is directly proportional to the concentration of hLH in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.2 mIU/ml

      - ULOQ: 250 mIU/ml

Analytical measurement range (AMR): 0.2 - 250 mIU/ml

Clinical Reportable Range (CRR): ): 0.2 - 250 mIU/ml

Sensitivity: 95%

Specificity: N/A

- Note: Values greater than CRR may be reported, with further dilution, if requested by Clinician/doctor upon HOD’s approval.

- Type of Sample

A minimum of 2ml of non-hemolysed Serum/Plasma

Collect blood sample by venipuncture and separate plasma by centrifugation.

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

- No additives/preservatives in the tube

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

Heparin tube- Green top Vacutainer

No additives/preservatives in the tube

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

      - DXI hLH Reagent Packs

      - Calibrators

      - Substrate

      - DXI, Coulter, SYNCHRON LX®i: DXI Wash Buffer II, Cat. No. A16792

      - UniCel® DxI: UniCel DxI Wash Buffer II, Cat. No. A16793

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-8°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

- Do not use reagent kits beyond the expiration date.

Do not mix reagents from different reagent kits.

Prior to loading the Reagent Kit into the system for the first time, the microparticle bottle requires mixing to resuspend microparticles that have settled during shipment.

Septa MUST be used to prevent reagent evaporation and contamination and to ensure reagent integrity. Reliability of assay results cannot be guaranteed if septa are not used according to the instructions in this package insert.

To avoid contamination, wear clean gloves when placing a septum on an uncapped reagent bottle.

Once a septum has been placed on an open reagent bottle, do not invert the bottle as this will result in reagent leakage and may compromise assay results.

Over time, residual liquids may dry on the septum surface. These are typically dried salts which have no effect on assay efficacy.

For a detailed discussion of handling precautions during system operation, refer to the COULTER System Operations Manual.

- Calibration Procedure

Calibrator Required:

      - Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

Calibration schedule:

- Recalibrate the assay every 28 days.

- Test Procedure

Refer to the appropriate system manuals and/or Help system for a specific description of installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

Mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use forty-five (45) µL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) µL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

For detailed information on ordering dilutions, refer to the DXI System Operations Manual, Section 5.

Handling of repeat test results:

      - In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever higher).

      - If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

- Quality Control Procedure

Internal Quality Control

- Biorad, Level 1, Level 2 and Level 3 any two levels of Quality Control sera i.e. low, normal and high range as per NABL 112.

    - Preparation of Controls:

- Control material should be at room temperature before reconstitution. Keep vials for 20 - 30 minutes to reach the room temperature.

- Gently tap the lyophilized vial on the counter several times to dislodge any material adhering to the stopper.

- Using an auto pipette, reconstitute each vial with 5mL volume of sterile water for injection.

- Replace the stopper and allow the control to stand for 30 minutes at room temperature by swirling occasionally.

- Before sampling gently swirl the vial several times to ensure homogeneity (frothing to be avoided). Prepare aliquots of 500 µL each or as required, label with “Control name, Date of preparation, initial of the technician”.

- Store the aliquots at -20°C.

    - Procedure:

      - Run Quality Control according to TRUSTlab’s Quality Control protocol described in Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).

      - Remove the QC material stored at -20ºC, keep it for 30-45 minutes at room temperature. If the material is taken from the 2-8ºC refrigerator, keep it for 20-30 minutes at room temperature.

      - Mix the material thoroughly by gentle inversion.

      - Place a portion of the material in a sample cup (approximately 350 µl).

      - Place the cup in the tray positions indicated by the screen.

      - Perform Internal Quality Control once at the beginning of the day.

      - Check quality control values. It should be within the acceptable range i.e. + 2SD.

      - If it is within acceptable range, only then proceed further for patient’s samples.

      - In case of QC failure, remedial action is taken as per QC run programme and the details are entered in the QC failure log.

      - At the end of every month, print the Levey Jennings chart and calculate the cumulative precision i.e. %CV.

      - Review analyte performance via monthly LJs in terms of precision and accuracy as far as possible.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

External Quality Control

- EQAS participation –- Metropolis EQAS

Storage & Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2-8°C.

      - Reconstituted: Stable for 7 days when stored at 2-8°C. Stable up to 30 days when stored at -20°C.

      - Longer storage: -20°C

      - Aliquots are stored at the above temperatures as required.

- Interferences & Cross Reactions

Administration of gonadotropin-releasing hormone (GnRH), clomiphene citrate, and human menopausal gonadotropin (hMG) to females can result in elevated endogenous LH values, values. Administration of hMG to males can result in elevated endogenous LH values, while administration of testosterone decreases endogenous LH values.

Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays.

Patients routinely exposed to animals or to animal serum products can be prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

| Serum Specimen Status | Concentration of insignificant effect on assay |
| --- | --- |
| Hemolyzed | 500 mg/dl of hemoglobin |
| Lipemic | 3000 mg/dl of triglycerides |
| Icteric | 20 mg/dl of bilirubin |

Administration of estrogens or testosterone can result in decreased endogenous LH.

- Result Calculation

- TRUSTlab NRL follows estimation of CV%.

- Biological Reference Interval

| Males (mIU/ml) | Females (mIU/ml) |  |  |  |
| --- | --- | --- | --- | --- |
| 1.24 - 8.62 | Mid-Follicular Phase | Mid-Cycle Peak | Mid-Luteal Phase | Post menopausal |
|  | 2.12 - 10.89 | 19.18 - 103.03 | 1.20 - 12.86 | 10.87 - 58.64 |

- Reportable Interval

Analytical Measurement Range (AMR): 0.2–250 mIU/ml

Clinical Reportable Range (CRR): 0.2–250 mIU/ml

Dilution Procedure:

      - If a sample contains more than the stated value of the highest Access hLH Calibrator (S5), report the result as greater than that value (i.e., > 250 mIU/ml [IU/l]). Alternatively, dilute one volume of sample with one volume of Access hLH Calibrator S0 (zero) or Access Sample Diluent A. Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution.

      - The DxI system onboard dilution feature automates the dilution process, using one volume of sample with one volume of Access Sample Diluent A, allowing samples to be quantitated up to approximately 500 mIU/ml (IU/l). The system reports the results adjusted for the dilution.

- Procedure when Results are not in Range

Check the reagent lot and expiry.

Check the IQC status.

Find root cause for error.

Complete corrective action.

If methodology, IQC and Instrumental status are fine, then consult with Clinician for clinical significance.

If the doctor requests a rerun, the sample will be rerun using appropriate dilution methods.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

Luteinizing hormone (LH) is a glycoprotein hormone having two subunits. The alpha subunit is similar to those of follicle-stimulating hormone (FSH), human chorionic gonadotropin (hCG), and thyroid-stimulating hormone (TSH). The beta subunit is different from those of the other glycoprotein hormones and confers its biochemical specificity.

LH is secreted by the anterior pituitary in response to gonadotropin-releasing hormone (GnRH) secreted by the hypothalamus. In males, LH is also called interstitial cell-stimulating hormone (ICSH).1,2 In both males and females, LH secretion is regulated by a balance of positive and negative feedback mechanisms involving the hypothalamic-pituitary axis, the reproductive organs, and the pituitary and sex steroid hormones. LH and the other pituitary gonadotropin, FSH, play a critical role in maintaining the normal function of the male and female reproductive systems.

Abnormal LH levels with corresponding increased or decreased levels of FSH, estrogens, progesterone, and testosterone are associated with a number of pathological conditions.

- Increased LH levels are associated with menopause, primary ovarian hypofunction, and polycystic ovary disease in females and primary hypogonadism in males. Decreased LH levels are associated with primary ovarian hyperfunction in females and primary hypergonadism in males.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition.

      - Product insert of hLH, Beckman Coulter DxI 800.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).