SOP – 049

PROCEDURE FOR

TESTOSTERONE TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	10

8.	Type of Container & Additives	11

9.	Patient Preparation	11

10.	Equipment & Reagents	11

11.	Safety	12

12.	Calibration Procedure	12

13.	Test Procedure	13

14.	Quality Control Procedure	13

15.	Interferences & Cross Reactions	15

16.	Result Calculation	16

17.	Biological Reference Interval	16

18.	Reportable Interval	16

19.	Procedure when Results are not in Range	17

20.	Critical Alert Values	17

21.	Laboratory Clinical Interpretation	17

22.	Potential Source of Variability	18

23.	References	18

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | FSH | Follicle Stimulating Hormone |
| 9 | HOD | Head of Department |
| 10 | IQC | Internal Quality Control |
| 11 | LH | Luteinizing Hormone |
| 12 | LLOQ | Lower Limit of Quantitation |
| 13 | MSP | Management System Procedure |
| 14 | N/A | Not Applicable/Available |
| 15 | QC | Quality Control |
| 16 | QM | Quality Manager |
| 17 | QNS | Quantity Not Sufficient |
| 18 | SD | Standard Deviation |
| 19 | SHBG | Sex Hormone Binding Globulin |
| 20 | SOP | Standard Operating Procedure |
| 21 | SST | Serum Separator Tube |
| 22 | TRF | Test Requisition Form |
| 23 | TDPL | TRUSTlab Diagnostics Private Limited |
| 24 | ULOQ | Upper Limit of Quantitation |

- Objective

- The purpose of this SOP is to describe the procedure for the quantitative estimation of Testosterone in serum and plasma samples using Beckman Coulter CLIA analyzer DxI 800. This SOP also provides information about the test principle, clinical significance etc.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Biochemistry Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Testosterone is a hormone important for male sexual development and reproduction, but it's also present in women in smaller amounts. It plays a crucial role in male development, including muscle and bone mass, sperm production, and sexual function.

- Test Principle

The Access Testosterone assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of total testosterone levels in human serum and plasma using the Access Immunoassay Systems.

Testosterone in males is secreted by adult Leydig cells and is controlled principally by luteinizing hormone (LH). The majority of serum testosterone is bound to sex hormone binding globulin (SHBG), but it also exists loosely bound to albumin and in the free state. An abnormally low total testosterone level in males can be indicative of hypogonadism, hypopituitarism, hyperprolactinemia, renal failure, hepatic cirrhosis, or Kleinfelter’s syndrome. High total testosterone values in males can be caused by adrenal and testicular tumors, congenital adrenal hyperplasia or abnormalities of the hypothalamic-pituitary-testicular axis.

In females, testosterone is produced in the ovaries, adrenal gland, and peripheral fatty tissues and has a serum concentration that is approximately 10-fold less than in males. As with males, the majority of serum testosterone in females is bound to SHBG and albumin with a small amount in the free state. Increased female total testosterone levels may indicate polycystic ovary syndrome (PCOS), stromal hyperthecosis, ovarian and adrenal tumors, congenital adrenal hyperplasia and other disorders of the hypothalamic-pituitary-ovarian axis.

- Specific Performance Characteristics

Accuracy: > 90%

Limit of Quantitation:

      - LLOQ: 0.1 ng/ml

      - ULOQ: 16 ng/ml

Analytical measurement range (AMR): 0.1 – 16 ng/ml

Clinical Reportable Range (CRR): 0.1 – 16 ng/ml

Sensitivity: The lowest detectable level of testosterone distinguishable from zero (Access Testosterone Calibrator S0) with 95% confidence is 0.1 ng/ml (0.35 nmol/). This value is determined by processing a complete six-point calibration curve, controls, and 10 replicates of the zero calibrator in multiple assays. The analytical sensitivity value is calculated from the curve at the point that is two standard deviations from the fitted zero calibrator signal.

Specificity: Samples containing up to 10 mg/dl (171 μmol/l) bilirubin, 1000 mg/dl (10 g/l) hemoglobin, the equivalent of 1800 mg/dl (20.32 mmol/l) triglycerides (Triolein), or between 5.5-8.5 g/dl total protein (human serum albumin) do not significantly affect the concentration of total testosterone assayed.

The following table describes the cross-reactivity of the assay with substances that are similar in structure to testosterone. Potential cross-reactors were spiked into a testosterone sample of approximately 1.5 ng/ml.

| Substance | Analyte Added (ng/ml) | Cross-reactivity (%) |
| --- | --- | --- |
| Compounds Present in Human Serum |  |  |
| Testosterone-glucoronide | 100 | 0.4 |
| Testosterone-sulfate | 100 | 0.3 |
| 5-alpha-DHT | 100 | 2.0 |
| Androstanediol | 100 | 0.4 |
| Androstenediol | 100 | 0.6 |
| Androstenedione | 100 | 0.7 |
| DHEA | 1000 | 0.0 |
| DHEA-sulfate | 1000 | 0.0 |
| Androsterone | 100 | 0.2 |
| Corticosterone | 1000 | 0.0 |
| Cortisol | 1000 | 0.0 |
| Estradiol | 100 | 0.0 |
| Estradiol-sulfate | 100 | 0.0 |
| Estriol | 100 | 0.2 |
| Estrone | 100 | 0.4 |
| Estrone-glucuronide | 100 | 0.0 |
| Estrone-sulfate | 100 | 0.0 |
| Progesterone | 100 | 0.4 |
| 11-Deoxycortisol | 1000 | 0.0 |
| 17-alphaHydroxyprogesterone | 100 | 0.1 |
| 19-Hydroxytestosterone | 100 | 0.5 |
| 2-Hydroxyestradiol | 100 | 0.0 |
| Birth Control |  |  |
| Ethinylestradiol | 100 | 0.3 |
| Mestranol | 100 | 0.0 |
| Norethindrone | 100 | 0.05 |
| Norgestrel | 100 | 0.3 |

| Substance | Analyte Added (ng/ml) | Cross-reactivity (%) |
| --- | --- | --- |
| Drugs |  |  |
| Danazol | 100 | 0.3 |
| Mesterolone | 100 | 1.5 |
| Dexamethasone | 1000 | 0.0 |
| 19-Nortestosterone | 100 | 1.6 |
| Ethinyltestosterone | 100 | 0.07 |
| Structurally Related Compounds |  |  |
| 19-Norethisterone Acetate | 100 | 0.02 |
| 11B-Hydroxytestosterone | 100 | 4.1 |
| 11-Ketotestosterone | 100 | 6.7 |
| 17-alpha-Methyltestosterone | 100 | 0.2 |

- Type of Sample

Serum is the recommended sample.

Observe the following recommendations for handling, processing, and storing blood samples:

      - Collect all blood samples observing routine precautions for venipuncture.

      - For serum, allow samples to clot adequately before centrifugation.

      - Keep tubes stoppered at all times.

      - Physically separate serum or plasma from contact with cells as soon as possible.

      - Store samples, tightly stoppered, at room temperature (15 to 30°C) for no longer than eight hours.

      - If the assay is not expected to be completed within eight hours, refrigerate the samples at 2 to 8°C.

      - If the assay is not expected to be completed within 48 hours, or for shipment of samples, freeze at -20°C or colder.

      - Use the following guidelines when preparing specimens:

      - Ensure residual fibrin and cellular matter has been removed prior to analysis.

      - Follow blood collection tube manufacturer’s recommendations for centrifugation.

- Type of Container & Additives

SST - Yellow top Vacutainer

Clot activator - Red top Vacutainer

- Patient Preparation

The phlebotomists should be courteous and compassionate towards the patients and their attendants. Ask the patient to sit comfortably. Maintain calmness and silence during blood drawing to avoid mistakes.

Check the patient’s identification; ask the patient for his or her full name, address, identification number, and/or date of birth.

No specific preparation is required.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

      - Access Testosterone Reagent Pack. Cat. No. 33560: 100 determinations, 2 packs, 50 tests/pack

      - Provided ready to use.

      - Store upright and refrigerate at 2-10°C.

      - Refrigerate at 2-10°C for a minimum of 2 hours before use on the instrument.

      - Stable until the expiration date stated on the label when stored at 2-10°C.

      - Stable at 2-10°C for 28 days after initial use.

      - Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range.

      - If the reagent pack is damaged (i.e. broken elastomer), discard the pack.

      - All antisera are polyclonal unless otherwise indicated.

| R1a: | Paramagnetic particles coated with goat anti-mouse IgG; testosterone alkaline phosphatase conjugate with bovine serum albumin (BSA), < 0.1% sodium azide, and 0.1% ProClin 300. |
| --- | --- |
| R1b: | Sample Treatment Solution, < 0.1% sodium azide. |
| R1c: | Monoclonal anti-testosterone (mouse), protein (BSA, mouse, goat), < 0.1% sodium azide, 0.1% ProClin 300. |

Reagent Storage Condition Stability

          - Reagents are stable, unopened, up to the stated expiry date when stored at 2-8°C.

          - Once open, reagents stored on board the instrument are stable for 28 days.

- Safety

Use this reagent for DXI 800 only.

Do not pipette by mouth. Exercise the normal precautions required for handling laboratory reagents.

Exposed metal surfaces should be cleaned with 1% Sodium hypochlorite.

The reagents must be used only for the purpose intended by suitably qualified laboratory personnel, under appropriate laboratory conditions.

Used cuvettes contain human body fluids, it should be handled with appropriate care to avoid skin contact.

- Calibration Procedure

Calibrator Required:

      - Manufacturer-supplied 6 Point Calibrator.

Calibration schedule:

      - An active calibration curve is required for all tests

      - Recalibrate the assay every 28 days.

      - Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

- Test Procedure

CAUTION: Do not load more than one size of sample container in each rack. The rack indicator must be positioned at the correct setting for the size of sample container.

Prepare the sample container for each sample, and place barcode labels on the sample containers, as required.

Load each sample container into a rack, ensuring that the barcode labels are clearly visible through the slot in the rack.

Place the racks in the entry queue.

Refer to the appropriate system manuals and/or Help system for a specific description of the installation, start-up, principles of operation, system performance characteristics, operating instructions, calibration procedures, operational limitations and precautions, hazards, maintenance, and troubleshooting.

Mix contents of new (unpunctured) reagent packs by gently inverting the pack several times before loading on the instrument. Do not invert open (punctured) packs.

Use ten (10) μL of sample for each determination in addition to the sample container and system dead volumes. Use sixty-six (66) μL of sample in addition to the sample container and system dead volumes for each determination run with the DxI system onboard dilution feature. Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

The system default unit of measure for sample results is ng/ml. To change sample reporting units to the International System of Units (SI units), µg/l, refer to the appropriate system manuals and/or Help system. To manually convert concentrations to the International System, multiply ng/ml by multiplication factor.

CAUTION: Discard the primary reagent pack at the end of the Low Volume Usage onboard stability interval. When the Primary Reagent Stability option is set to USE BEYOND, the system does not display a flag when the reagent exceeds the onboard stability.

Do not use reagents beyond the expiration date.

- Quality Control Procedure

Internal Quality Control

- Quality control materials simulate the characteristics of patient samples and are essential for monitoring the system performance of immunochemical assays. Because samples can be processed at any time in a “random access” format rather than a “batch” format, quality control materials should be included in each 24-hour time period. More frequent use of controls or the use of additional controls is left to the discretion of the user based on good laboratory practices or laboratory accreditation requirements and applicable laws. Follow the manufacturer’s instructions for reconstitution and storage. Each laboratory should establish mean values and acceptable ranges to ensure proper performance. Quality control results that do not fall within acceptable ranges may indicate invalid test results. Examine all test results generated since obtaining the last acceptable quality control test point for this analyte. Refer to the appropriate

- System Manuals and/or Help system for information about reviewing quality control results.

    - Handling of QC failure:

- Verify the following points and, if required, consult with   HOD/QM:

      - Review the environmental conditions and review recent events and look for any deviation.

      - Check onboard reagents (for stability, deposits, fungal growth) reference fluid and controls.

      - Check the quality of the sterile water for injection used, and used for the preparation of controls and calibrators, contaminants i.e. bacterial growth may contribute to inaccurate results.

      - Check reaction temperature.

      - Check the expiry date of kit and contents.

      - Check the instrument settings and light source.

      - If there are any control failures take appropriate corrective and preventive actions as per IQC MSP.

      - Record in Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032).

      - Take appropriate corrective actions and re-run the QC.

      - Contact application support/service support from vendor if required the corrective action fails to yield acceptable results.

    - Repeat and Recheck Criteria:

      - If QC results in the analytical run are out of control limits and review requires a repeat.

      - If indicated by clinical validation.

      - In case of inadvertent equipment breakdown or power failure.

      - To handle Customer request/complaints upon approval by HOD

Storage & Stability of Controls

      - Unopened: Stable until the expiry date when stored at 2-8°C.

      - Reconstituted: Stable for 7 days when stored at 2-8°C. Stable up to 30 days when stored at -20°C.

      - Longer storage: -20°C

      - Aliquots are stored at the above temperatures as required.

Handling of Repeat Test Results

- In the case of repeat testing the value obtained first is reported, if the percentage difference between the values (original as reference) is less than 10 % or 1.5 times the parameter CV% (whichever higher).

- If the percentage difference between the values is greater than 10 % or 1.5 times the parameter CV% (whichever higher), root cause analysis will be done for appropriate corrective actions by the HOD for subsequent reporting & the same will be documented. If required a fresh sample may be requested (subject to feasibility).

External Quality Control

- EQAS participation –- Metropolis EQAS

- Interferences & Cross Reactions

Heterophile antibodies: These antibodies can interfere with the assay and lead to falsely elevated testosterone levels, particularly in women.

Androgenic compounds: Certain medications or substances can interfere with the assay, potentially leading to inaccurate results.

High levels of DHEA-S: Dehydroepiandrosterone sulfate (DHEA-S) can cross-react with the testosterone assay, especially at high concentrations, leading to an overestimation of testosterone levels.

Hemolysis: Hemolysis, or the breakdown of red blood cells, can interfere with the assay, potentially leading to decreased testosterone results.

Anti-animal Antibodies: Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce human anti-animal antibodies, e.g. HAMA (Human Anti-Mouse Antibody), that interfere with immunoassays.

Various other substances in the blood can interfere with the assay, including binding proteins like Sex Hormone Binding Globulin (SHBG).

Cross-Reactions: DHEA-S can cross-react with the testosterone assay, leading to an overestimation of testosterone levels.

Other medications: Certain medications, like Danazol and mifepristone, can interfere with the assay.

- Result Calculation

TRUSTlab NRL follows estimation of CV%.

Serum and plasma values should not be used interchangeably. Refer to the appropriate section of the table below:

| Sample Type | Reference Group | n | Median Age (years) | Age Range | Median Concentration (ng/ml) | 95% Reference Interval (ng/ml) |
| --- | --- | --- | --- | --- | --- | --- |
| Serum | Males | 240 | 41 | 18-66 | 3.84 | 1.75-7.81 |
|  | Females | 240 | 43 | 21-73 | 0.29 | < 0.1-0.75 |

| Sample Type | Reference Group | n | Median Age (years) | Age Range | Median Concentration (ng/ml) | 95% Reference Interval (ng/l) |
| --- | --- | --- | --- | --- | --- | --- |
| Plasma (heparinized) | Males | 240 | 41 | 18-66 | 3.86 | 1.68-7.58 |
|  | Females | 240 | 43 | 21-73 | 0.41 | < 0.1-0.90 |

- Biological Reference Interval

- 0.1 — 16 ng/ml

- Reportable Interval

Analytical Measurement Range (AMR): 0.1 - 16 ng/ml

Clinical Reportable Range (CRR): 0.1 - 16 ng/ml

- Procedure when Results are not in Range

Samples can be accurately measured within the analytic range of the lower limit of detection and the highest calibrator value (approximately 0.1-16 ng/ml [0.35-55.5 nmol/l]).

If a sample contains less than the lower limit of detection for the assay, report the result as less than that value (i.e., < 0.1 ng/ml [< 0.35 nmol/l]).

If a sample contains more than the stated value of the highest Access Testosterone Calibrator (S5), report the result as greater than that value (i.e. > 16 ng/ml [> 55.5 nmol/l]). Alternatively, dilute one volume of sample with one volume of Access Testosterone Calibrator S0 (zero). Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution.

For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce antibodies, e.g. HAMA, that interfere with immunoassays. Additionally, other heterophile antibodies such as human anti-goat antibodies may be present in patient samples.

Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

The Access Testosterone results should be interpreted in light of the total clinical presentation of the patient, including symptoms, clinical history, data from additional tests and other appropriate information.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

In men testosterone is synthesized almost exclusively by the Leydig cells of testis. Most of the circulating testosterone is bound to carrier proteins. In women, small quantities of testosterone are formed in the ovaries.

Determination of testosterone in women is helpful in diagnosis of Polycystic ovaries (Stein Leventhal syndrome), management of hirsutism & virilisation in females.

In men, reduced production may indicate:

      - Hypogonadism

      - Oestrogen therapy Chromosome aberrations (as in Klinefelter’s syndrome)

      - Liver cirrhosis

      - Delayed puberty

In men, increased production may indicate:

      - Precocious puberty

      - Congenital adrenal hyperplasia.

- Potential Source of Variability

      - Lysed serum specimens may give falsely elevated values.

      - If the reagents are turbid, DO NOT USE, as it indicates contamination of the reagent.

      - Strong lipemic and haemolytic sera should not be used.

      - Contaminated glassware is the greatest source of error. Disposable plastic tubes and clean glassware are recommended for the test.

- References

      - Interpretations on Diagnostics Tests, 7th Edition, Jacques Wallach.

      - Tietz. Textbook of Clinical Biochemistry and Molecular Diagnostics, 4th edition

      - Product insert of Testosterone, Beckman Coulter DxI 800.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).