SOP – 002

PROCEDURE FOR

URINE IQC

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Procedure	9

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	10

16.	Result Calculation	10

17.	Biological Reference Interval	10

18.	Reportable Interval	10

19.	Procedure when Results are not in Range	10

20.	Critical Alert Values	10

21.	Laboratory Clinical Interpretation	10

22.	Potential Source of Variability	10

23.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | CP | Clinical Pathology |
| 3 | CV | Coefficient of Variation |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HBV | Hepatitis B Virus |
| 6 | HIV | Human Immunodeficiency Virus |
| 7 | HOD | Head of Department |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | pH | Potential of Hydrogen |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to define a procedure for internal quality control in the department of clinical pathology. This SOP also provides information about the procedural principles, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Clinical Pathology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Clinical Pathology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Urine Internal Quality Control (IQC) involves using control materials to monitor the accuracy and precision of urine analysis tests. These controls are compared to expected values, ensuring the test results are within acceptable ranges. IQC helps to identify and address potential errors in the laboratory analysis process.

- Principle

- The use of quality control materials is indicative of an objective assessment of the precision of methods and techniques in use, and is an integral part of good laboratory practices.  The department requires daily estimation of assayed material for quality control of urine.

- Specific Performance Characteristics

- N/A.

- Type of Sample

- Control Level 1 & Control Level 2

- Type of Container & Additives

- N/A.

- Patient Preparation

N/A.

- Equipment & Reagents

Equipment

- ERBA LAURA

Reagents

- 10 Test Strips

- Safety

Follow universal precautions like wearing suitable gloves and eye/face protection while handling all reagents, samples and operating the instrument.

Universal precautions imply that all body fluids are treated as if known to be infectious for Human Immunodeficiency Virus (HIV), Hepatitis B Virus (HBV) or other Blood borne Pathogens.

Dispose of all waste in accordance with applicable national and/or local regulations.

Some reagents contain sodium azide, which may react with copper or lead plumbing to form explosive metal azides. Use caution in disposing of these reagents. If disposing to drain, flush with large volumes of water to prevent azide buildup.

Waste material containing patient samples or biological products should be considered biohazardous when disposing or treating.

Clean up all spills immediately and thoroughly. Disinfect the area for any spills involving biohazardous materials. Dispose of all contaminated materials appropriately.

- Calibration Procedure

- N/A.

- Procedure

      - The quality control material sourced from commercial sources is used for the regular monitoring of internal quality control at TRUSTlab’s CP Department.

      - The control samples are run in a manner similar to the regular samples, at the start of the day, before processing any of the test samples.

      - The control material is in liquid form. It is processed as per the guidelines of the manufacturer, prior to testing.

      - Every day, two levels of controls are performed for ERBA LAURA. The schedule of the frequency and the number of levels to be assayed and acceptable criteria are defined as follows:

| Equipment Name/Strip name | Test monitored | Levels of QC used (Singular level / Bi-level/ Two-Level) | Frequency of QC run / time of run | Acceptability Criteria |
| --- | --- | --- | --- | --- |
| ERBA LAURA | Protein, Glucose, Specific gravity, pH, Leucocytes, Nitrate, Ketone bodies, Urobilinogen, Erythrocytes, Haemoglobin | 2 (Normal & High) | Once in 24 hrs. | As per the kit insert |

      - Compare the results obtained with those mentioned in the manufacturer’s technical insert.

      - If the values are deviating, perform the quality control check with fresh strips and fresh quality control sample.

      - Once quality control values fall within acceptable limits, then processing of regular samples may proceed.

      - Results will be recorded in the Daily Urine QC Form (TDPL/CP/FOM-001).

      - All results are entered by the technician and verified by the HOD.

      - If results are out of range, check control validity and instrument settings.

      - If all checks are OK, rerun the controls again and see if they pass.

      - If the results are still out of range, then contact the instrument engineer.

      - Log all QC failures in the Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-036).

      - Record all the values and calculate monthly CV% for pH and specific gravity.

- Quality Control Procedure

- As described above.

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- Negative/Positive.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

      - Manufacturer’s kit inserts

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013)

      - NABL 112