SOP – 007

PROCEDURE FOR

ACTIVATED PARTIAL THROMBOPLASTIN TIME

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	10

15.	Interferences & Cross Reactions	10

16.	Result Calculation	10

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | APTT | Activated Partial Thromboplastin Time |
| 2 | °C | Degree Centigrade |
| 3 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HOD | Head of Department |
| 6 | HM | Haematology |
| 7 | INR | International Normalized Ratio |
| 8 | IQC | Internal Quality Control |
| 9 | ISI | International Sensitivity Index of the reagent |
| 10 | MSP | Management System Procedure |
| 11 | N/A | Not Applicable/Available |
| 12 | PPP | Platelet Poor Plasma |
| 13 | PT | Prothrombin Time |
| 14 | QC | Quality Control |
| 15 | QM | Quality Manager |
| 16 | QNS | Quantity Not Sufficient |
| 17 | rpm | Revolutions per minute |
| 18 | SOP | Standard Operating Procedure |
| 19 | TRF | Test Requisition Form |
|  | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the procedure for performing the Activated Partial Thromboplastin Time test by a conventional routine method using citrated, platelet-poor plasma. This SOP also provides information about the test principle, clinical significance and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- APTT stands for Activated Partial Thromboplastin Time. It's a blood test used to evaluate the efficiency of the coagulation cascade, which is the process of blood clotting. Essentially, it measures how long it takes for blood to clot in the presence of certain reagents.

- Test Principle

Cephaloplastin activates the coagulation factors of the intrinsic pathway of the coagulation mechanism in the presence of calcium ions.

APTT is prolonged by a deficiency of one or more of these clotting factors of the intrinsic pathway and in the presence of coagulation inhibitors like heparin.

Method: Opto-mechanical.

- Specific Performance Characteristics

Linearity: N/A

Measurement Range: 20 – 180 seconds

Sensitivity: 25.3 to 158.3 seconds

Clinical Reportable Range: N/A

- Type of Sample

Plasma - Trisodium citrate Sample (3.2%)

Draw blood without undue venous stasis and without frothing into a plastic syringe fitted with short needle of 19 to 20 SWG. The vein puncture must be a clean one. Transfer the blood to the tubes, after detaching the needle from the syringe.

Mix exactly nine parts of freshly collected blood with one part of Tri-sodium citrate.

Centrifuge immediately for 15 minutes at 3500 rpm and transfer the plasma into a clean test tube

Sample Stability: Plasma must be tested preferably immediately. The samples are stable only for:

      - 2 hours, at 22-24°C

      - 3 hours, at 2 -8°C

Sample Retention: Sample is retained for 24 hours at 2 – 8℃.

Sample Rejection Criteria:

      - Quantity not sufficient (QNS)

      - Sample received in wrong vacutainer

      - No TRF; Grossly incomplete TRF; Errors in labeling

      - Grossly haemolysed specimens

      - Highly lipemic specimens

      - Highly icteric specimens

      - Sample received beyond stability period

      - Incomplete TRF details or labelling errors are clarified with Customer/Doctor/ Patient before sample is accepted for processing. If critical information is still not obtained, the sample is rejected.

- Type of Container & Additives

- Sodium Citrate (3.2%) vacutainer.

- Patient Preparation

Patients should not have exercised heavily before blood collection.

Fasting or only non-fatty meals prior to blood collection is preferred as this results in procuring samples with a desirable lower opacity.

- Equipment & Reagents

Equipment

      - Centrifuge

      - Water bath

      - Stopwatch

      - Thermometer

      - Pipette (20 µl – 200 µl)

Reagents

      - Blood in 3.2% Sodium citrate

      - UNIPLASTIN

      - Plastic Test Tube (12x75mm)

- Safety

Handle all samples as potentially infectious.

Handle all reagents with care and avoid contact with eye, mouth and skin.

Do not perform mouth pipetting.

Discard used reagents and samples as per disposal procedure.

- Calibration Procedure

- N/A.

- Test Procedure

      - Before use, the reagent should be mixed well by gentle swirling. Do not shake.

      - Aspirate from the reagent vial enough reagent for immediate testing requirements in a thoroughly clean and dry test tube.

      - Bring the reagent vial to room temperature (20-30°C) before prewarming at 37°C for testing purposes.

      - Recap the reagent vial and replace the vial immediately into the 2-8°C refrigerator.

      - Separate test tubes containing Liquicelin-E Calcium Chloride Solution should be brought to 37°C. (Depending on volume, approximately 5 to 10 minutes required). Do not incubate the test plasma.

      - To a 12x75mm test tube, add 0.1 ml (100µl) of test plasma (PPP) and 0.1 ml (100µl) Liquicelin-E. Shake tube briefly to mix the reagent and plasma; place the tube at 37°C for 3 to 5 minutes.

      - Following incubation period, add forcibly add 0.1 ml (100µl) of prewarmed calcium chloride into the plasma Liquicelin-E; simultaneously start a stopwatch. Shake the tube briefly to mix contents, keep at 37°C for 15 seconds.

      - Following 15 seconds incubation, remove the tube; gently tilt the tube back and forth until a gel clot forms; stop the stopwatch; record time.

- Quality Control Procedure

Internal Quality Control

    - Procedure:

      - Previously tested known samples are aliquoted and preserved at -20°C and used each day of the test performed.

      - Platelet poor plasma-platelet count is checked every month. The count should be less than 10000/mm3.

External Quality Control

- EQAS participation –- CMC EQAS

- Interferences & Cross Reactions

- Turbid, icteric, lipemic or grossly hemolysed samples may generate erroneous PT results.

- Result Calculation

- The result may be reported directly in terms of the Mean APPT of the test plasma clotting time in seconds.

- Biological Reference Interval

22.0 -37.0 seconds

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

      - The aPTT test measures the length of time (in seconds) that it takes for clotting to occur when reagents are added to plasma (liquid portion of the blood) in a test tube.

      - The aPTT is used to evaluate the following coagulation factors:

        - XII, XI, IX, VIII, X, V, II (Prothrombin), I (Fibrinogen)

      - The aPTT test is used for:

        - Unexplained Bleeding or Bruising

        - Thromboembolism (Blood Clot).

        - An Acute condition such as Disseminated Intravascular Coagulation (DIC) that can cause both bleeding and clotting as coagulation factors are activated and used up at a rapid rate.

        - A Chronic Condition such as Liver Disease, as the liver is the source of most Coagulation factors

        - When a woman has had recurrent miscarriages.

        - As part of an evaluation for antiphospholipid syndrome or lupus anticoagulant.

      - It may be used as part of a pre-surgical evaluation for bleeding tendencies, especially if the surgery carries an increased risk of blood loss and/or if the patient has a clinical history of bleeding - such as nosebleeds and bruising that may indicate the presence of an inherited or acquired factor deficiency or of an acquired inhibitor.

      - When a patient is on Intravenous (IV) or injection heparin therapy, the APTT is ordered at regular intervals to monitor the degree of anticoagulation.

- Potential Source of Variability

- Incorrect mixture of blood and tri-sodium citrate, insufficient prewarming of plasma and reagent, contaminated reagents, glassware etc. are potential sources of errors.

- References

      - Practical Hematology, J. Dacie and M. Lewis, Tenth edition.

      - Uniplastin Kit insert - Tulip Diagnostics (P) Ltd.

      - Procedure for Managing Internal QC Program (TDPL/QA/MAN-001: MSP-013).