SOP – 002

PROCEDURE FOR

TISSUE PROCESSING & EMBEDDING

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Test Principle	8

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration	9

13.	Procedure	9

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	12

16.	Result Calculation	13

17.	Biological Reference Interval	13

18.	Reportable Interval	13

19.	Troubleshooting	13

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | °C | Degree Centigrade |
| 3 | CRR | Clinical Reportable Range |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HOD | Head of Department |
| 6 | IQC | Internal Quality Control |
| 7 | LOD | Limit of Detection |
| 8 | LOQ | Limit of Quantitation |
| 9 | MSP | Management System Procedure |
| 10 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 11 | N/A | Not Applicable/Available |
| 12 | pH | Potential of Hydrogen |
| 13 | QC | Quality Control |
| 14 | QM | Quality Manager |
| 15 | SOP | Standard Operating Procedure |
| 16 | TRF | Test Requisition Form |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

This document describes a detailed procedure on the processing of the fixed tissue sections submitted for histopathological or immunohistochemical examination.

Processed tissue sections are embedded in a solid medium (paraffin), firm enough to support the tissue and give it rigidity so that thin sections can be taken. These thin sections can then be subjected to histopathology staining and observed under microscope to give definitive diagnosis.

- Scope

- This procedure applies to all staff of the Histopathology Department of TRUSTlab, tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Histopathology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

Any other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definitions

Tissue processing: It involves treatment of tissue sections to allow paraffin wax impregnation by removing water from tissues and its gradual replacement with molten paraffin wax.

Dehydration: Process of removal of unbound water and aqueous fixatives from the tissues done in graded solutions slowly.

Clearing: Process of replacing the alcohol with a fluid which will dissolve the wax with which the tissue must be impregnated.

Impregnation: is the process of replacing the clearing agent and infiltration of the tissue with a medium (paraffin wax usually) so that the tissue structure is maintained. This will give a firm consistency to allow the cutting of thin sections without distortion and without alteration of the spatial relationships of the tissues and cellular elements.

Embedding: This involves the enclosing of properly processed correctly oriented specimens in a support medium that provides firmness to the tissue blocks thereby enabling cutting of thin sections.

- Test Principle

N/A.

- Specific Performance Characteristics

Uncertainty of measurement: NA

Accuracy: NA

Limit of quantitation (LOQ): NA

Limit of detection (LOD): NA

Analytical measurement range (AMR): NA

Clinical reportable range (CRR): NA

Sensitivity: NA

Specificity: NA

- Type of Sample

Specimens received at TRUSTlab.

Sample Retention: Sample is retained for 30 days.

Sample Rejection Criteria:

      - Wrongly labeled specimen

      - No tissue in the container

      - No fixative in the specimen container

- Type of Container & Additives

- Formalin-fixed specimen in sealed containers.

- Patient Preparation

N/A.

- Equipment & Reagents

Equipment

Automated Tissue Processor.

- •	Name:  Leica tissue processor.

- •	Environmental conditions: Stable ambient temperature.

Reagents

      - 10 % buffered formalin

      - Isopropyl Alcohol

      - Xylene

      - Paraffin wax

Reagent Storage Condition Stability

- All reagents mentioned above are stable at Room Temperature until the expiry date mentioned on the reagent package.

Additional Requirements

- Tissue Cassettes

- Safety

While handling the tissues and the various chemicals, the universal safety precautions as mentioned in biosafety manual.

Xylene is an irritant to skin and eyes. Use gloves while handling. If exposure occurs, wash with water and mild soap. Wash eyes with ample amount of water.

Xylene is explosive. Vapor accumulations may flash and/or explode if ignited. Keep ignition sources, open flames, etc. away from these fumes.

- Calibration

Calibration Schedule

- Calibration is done annually and also when there is a major equipment repair by the authorized service agency of the manufacturer.

- Procedure

Tissue Processing

      - Method: It is the process of subjecting tissue sections for fixation, dehydration, and infiltration with fixatives, alcohol, solvents and paraffin wax respectively.

      - Principle: Tissue processing is based on combination of the following principles:

        - Dehydration - It involves processing the tissues in gradually increasing strengths of alcohols to remove the water and fixative from the tissues.

        - Clearing - It is replacement of the dehydrating agent (alcohol) in the tissues by a clearing agent (xylene). This facilitates impregnation of the tissue by paraffin wax, because xylene is miscible with both the alcohol and paraffin wax.

        - Impregnation- It involves the replacement of the clearing agent in tissues by the embedding medium (paraffin wax)

      - Pre-Processing Requisites: Take fixed, sampled tissue sections placed in the plastic cassettes into a perforated stainless-steel basket of the tissue processor. Calcified tissue sections need to be decalcified before processing (Refer to SOP TLD/SOP/HP/01 for details of decalcification).

      - Sample Stability: Tissue sections kept in formalin are stable at Room Temperature

Procedure for Embedding

      - Description of the Equipment - Embedding Station

        - Major Components:

          - Paraffin wax reservoir with lid and dispenser: Capacity - 3 litres. There is a pivotable clip for dispensing the molten wax.

          - Cold Plate for cooling of blocks during and after embedding

        - Peripheral Material:

          - Plastic embedding rings

          - Plastic embedding cassettes

          - Forceps

          - Spatula, etc.

        - Calibration Frequency:

          - Calibration is done annually and also when there is major equipment       repair by the authorized service agency of the manufacturer.

      - Blocking with Embedding of tissues

        - The wax in the paraffin dispenser is kept molten at 20°C higher than the melting point of the wax.

        - For details of operation refer to the operator manual.

        - The final wax beaker on the automatic tissue processor is moved to the paraffin embedding station and the cassettes are transferred to the cassette warmer where the temperature is adjustable between 55 and 70°C.

        - After processing, the tissue is embedded in stainless steel embedding cassette along with plastic embedding cassettes.

        - The number of the specimen is written on one side of the embedding cassette with a marker pen and on the other side the year.

        - The label with the biopsy number is placed along with the embedding ring over the embedding cassette.

        - Place the processed tissue with the cutting surface facing downwards in the centre of the embedding cassette.

        - Molten wax from the paraffin dispenser is filled into the embedding cassette holding the tissue .

        - The mold is filled about one third with liquid paraffin, and while the paraffin is semi-liquid the sample can be oriented as required and the embedding cassette can be filled with the molten wax.

        - The assembly is then cooled.

        - For sectioning the metal embedding cassette is removed and the block is mounted by gripping the plastic embedding ring in the jaws of the microtome clamp.

      - In instances of instrument failure, manual blocking out is carried out using the same stainless steel embedding cassettes and plastic embedding rings and paraffin kept in molten state in a stainless steel container.

      - Maintenance of Embedding Station

        - Regularly remove dust from back of the instrument.

        - Xylene can be used to clean the work area.

        - Keep contaminants out of the paraffin reservoir.

        - Ensure residual amount of paraffin remains in the reservoir after draining to prevent solid contaminants from entering the dispenser. Absorb the paraffin with tissue paper. Do not remove the screen until the residual paraffin has been removed.

        - Remove and empty paraffin collection trays only when they are warm.

        - The paraffin in the collection trays must not be reused. Empty paraffin trays regularly so as to prevent overflowing in to the instrument which may damage the instrument.

- Quality Control Procedure

Internal Quality Control

      - Change the reagents (xylene and paraffin) weekly or when the xylene turns milky (whichever is earlier). Discard the alcohol of the first beaker and use the second beaker in its place, thus shift the third beaker in place of the second and so on, so that seventh beaker is in place of the sixth. Now fill the first beaker with fresh isopropyl alcohol and place it at the seventh position.

      - Enter the details about change of reagents in the Tissue Processor Reagent Change Register (TDPL/HP/FOM-005).

      - Temperature of the wax bath is recorded daily in the Daily Temperature Form for Tissue Processor (TDPL/HP/FOM-007).

      - QC failure: When faults due to poor processing as mentioned in trouble shooting occur, it is a quality control failure. The remedial measures are taken as mentioned in the Trouble shooting (faults due to poor processing).

      - The details are logged into the Daily IQC Outlier Non-Conformity & Preventive Action Form (TDPL/GE/FOM-032) in histopathology.

External Quality Control

- EQAS participation –- RML EQAS

- Interferences & Cross Reactions

Power failure - The Tissue Processor is connected to the UPS connection.

Proper fixation is important for good tissue processing, hence ensure that the tissue is properly fixed before processing (refer to TDPL/HP/SOP-001).

If more than 60 cassettes are placed in the basket, the sections kept at center of the basket may not be processed properly.

When the levels of the reagents are below the minimum level mark, tissue sections may become dry or may not be processed properly.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A

- Troubleshooting

When there is an instrument breakdown, contact the Service Engineer and inform the Pathologist. This is to be done even when there is a problem with the wax baths of the processor, and the wax is not in a molten state.

Enter the details about the problem in the Equipment Maintenance Form (Automatic Tissue Processor) (TDPL/BE/FOM-022).

When the thermostat of the wax bath has been set right, the temperature of the wax in the bath is to be measured and noted in the Daily Temperature Form for Tissue Processor (TDPL/HP/FOM-007), once in the presence of the service engineer and second time before starting the tissue processor for the day, before the instrument is declared fit for satisfactory operation.

After the instrument problem is corrected, complete the documentation in the Equipment Maintenance Form (Automatic Tissue Processor) (TDPL/BE/FOM-022) and get it signed by the engineer.

In the event of a breakdown of the tissue processor, processing is done manually.

Procedure for Manual Processing

      - Remove the glass beakers with reagents from the tissue processor and process the sections by manual method as follows:

      - Transfer the tissue basket from one reagent container to another container, manually, as mentioned below:

| Beaker No. | Reagent | Time |
| --- | --- | --- |
| 1 | 10% neutral buffered formalin | Overnight in formalin |
| For manual processing, if scheduled to start at 7 am |  |  |
| 2 | 70 % Isopropyl Alcohol | 1 hour |
| 3 | 80 % Isopropyl Alcohol | 1 hour |
| 4 | 90 % Isopropyl Alcohol | 1 hour |
| 5 | 95 % Isopropyl Alcohol | 1 hour |
| 6 | 100% Isopropyl Alcohol | 1 hour |
| 7 | 100% Isopropyl Alcohol | 1 hour |
| 8 | Xylene I | 1 hour |
| 9 | Xylene II | 1 hour |
| 10 | Paraffin wax I | 1 hour |
| 11 | Paraffin wax II | 1 hour |

Faults due to poor processing:

| Fault | Cause | Remedy |
| --- | --- | --- |
| Tissue has shrunken away from the wax | Insufficient dehydration | Reprocess the tissue |
| The tissue is quite soft when the block is trimmed | Insufficient fixation | Reprocess the tissue or take more sections |
| The tissue appears full of holes | Insufficient infiltration | Re-infiltrate and re-embed the tissue |
| On trimming, wax appears crystalline | Contaminated wax Block not cooled rapidly enough | Re-infiltrate, re-embed in fresh wax Re-embed and cool rapidly |
| On sectioning the tissue compresses but the wax does not | Insufficient infiltration in wax | Re-infiltrate and re-embed in wax |
| Specimen crumbles and drops out of wax leaving a rim of wax as a section | Insufficient infiltration in wax Overheated paraffin bath causing tissue hard and brittle | Re-infiltrate and re-embed in wax Embed in a hard wax |
| Sections crumble instead of flattening on water bath | Insufficient infiltration in wax | Re-infiltrate and re-embed in wax |
| Tissue is dried | Mechanical failure of the processing machine/power failure and tissue was out of solution and hung up. | Place the specimens in following rehydration solution for 18-24 hours: Sodium carbonate 1 gm Distilled water 70 ml Ethyl alcohol 30 ml |

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

      - LYNCH’S Medical laboratory technology, IIIrd Edition. Chapter 26. Page 890, 891.

      - Leica HistoCore PEARL Tissue Processor - Operator’s Manual

      - Procedure for Managing Internal QC Program (TDPL/AQ/MAN-001: MSP-013)