SOP - 001

PROCEDURE FOR

NUCLEIC ACID EXTRACTION

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	11

15.	Interfering Substances	11

16.	Result Calculation	11

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	11

23.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | DNA | Deoxyribonucleic Acid |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | GITC | Guanidinium thiocyanate |
| 4 | MI | Microbiology |
| 5 | MIC | Minimum Inhibitory Concentration |
| 6 | NA | Nucleic Acid |
| 7 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 8 | PCR | Polymerase Chain Reaction |
| 9 | PK | Proteinase K |
| 10 | PKRB |  |
| 11 | RNA | Ribonucleic Acid |
| 12 | rpm | Revolutions per minute |
| 13 | SOP | Standard Operating Procedure |
| 14 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - The purpose and scope of this document is to describe procedure for extraction of Nucleic Acid (DNA/RNA). This SOP is applicable to Nucleic Acid extraction from Nasopharyngeal or Throat swabs in whole blood, plasma, serum.

- Scope

This procedure applies to all staff of the Molecular Biology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Nucleic acid extraction involves isolating and purifying DNA or RNA from biological samples. This process typically includes breaking open cells, removing impurities, and then purifying the nucleic acid.

- Test Principle

RNA/DNA Viruses are lysed quickly by lysis buffer which contains high concentration of GITC. Lysis buffer and ethanol create appropriate conditions for binding nucleic acids to the silica membrane of the column. Poly A improves binding and recovery of low concentrated viral Nucleic Acid. Contamination like salts, metabolites and soluble macro molecular cellular components are removed with ethanolic buffers Wash Buffer 1 Wash Buffer 2. Nucleic acids can be eluted in low salt-based Elution Buffer.

- Specific Performance Characteristics

- N/A.

- Type of Sample

- Whole Blood EDTA / Plasma / Serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

Calibrated variable micropipettes

Sterile pipette filter tips (aerosol free)

1.5 ml/2ml Centrifuge tubes

1.5 ml/2ml Centrifuge tubes stand

NucleoDx Nucleic Acid Extraction Kit

Sterile powder free gloves

Lab coats, Facemask, Head cap

Cooling Centrifuge

Vortex mixer

Biosafety Cabinet

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Biosafety Cabinet and incubator calibrations are done on an annual basis by an NABL accredited calibration agency or by specific guidelines of equipment manufacturers.

- Test Procedure

Pre-Assay Instructions

    - Recommended sample volume for extraction and elution are as follows:

| Name of the Isolation Kit | Recommended Sample volume | Recommended Final Elution volume |
| --- | --- | --- |
| NucleoDx Nucleic Acid Extraction Kit | 200 µL | 60 µL |

    - Kit components for nucleic acid extraction are:

| Contents | 100 Rx | Storage |
| --- | --- | --- |
| Lysis Buffer | 30 ml | 20-25ᵒC |
| Binding Buffer | 30 ml | 20-25ᵒC |
| Wash Buffer-1 | 15 ml (Add 35ml 100% ethanol) | 20-25ᵒC |
| Wash Buffer-2 | 15 ml (Add 35ml 100% ethanol) | 20-25ᵒC |
| Precipitation Buffer | 45 ml | 20-25ᵒC |
| Proteinase K | 1 no. | -20ᵒC |
| Proteinase K reconstitution buffer | 2000 µl | 20-25ᵒC |
| Carrier RNA | 1 no. | -20ᵒC |
| Elution Buffer | 8 ml | 20-25ᵒC |
| Spin columns along with Collection tubes | 100 no.’s | 20-25ᵒC |

- NOTE: Use the kit for one year as per the Manufacture and Expiry dates mentioned on the kit.

Procedure

- Nucleic Acid can be isolated from Sample (whole blood, plasma, serum).

    - Preparation

      - Prepare wash buffer 1 and 2 by adding 35 ml of 100% ethanol to the respective buffer, to make a final volume of 50 ml each.

      - Add 2 ml of PKRB to the PK vial and store at -20ᵒC for up to 12 months.

      - Add 200 µl of elution buffer to Carrier RNA, mix well and store at -20ᵒC.

    - Lysis of Virus

      - Add 200 μl of sample to 1.5 ml centrifuge tube. Add 300 μl lysis buffer and add 20 μl of Proteinase K and 2 μl of Carrier RNA.

      - Incubate at room temperature (15–25°C) for 10 min with constant vertexing at every 3 minutes for 30 seconds.

    - RNA Precipitation

      - Add 400 μl of precipitation buffer to the lysate and invert the tube 5-6 times.

      - Loading on Column:

          - Transfer 700 μl of the lysate to the column and centrifuge the sample at 11,000 rpm for 1 minute. Discard the flow through.

    - Bind Viral RNA

      - Add 300 μl of binding buffer to the column and centrifuge for 30 seconds at 11,000 rpm and discard the flow through.

    - First Wash

      - Add 500 μl of Wash Buffer 1, close the cap, and centrifuge at 11000 rpm for 30 seconds. Discard the flow through.

      - Add 500 μl of Wash Buffer 2, close the cap, and centrifuge at 11000 rpm for 30 seconds. Discard the flow through.

      - Add 250 μl of Wash Buffer 2, close the cap, and centrifuge at 11000 rpm for 30 seconds. Discard the flow through.

    - Elution

      - Preheat the elution buffer at 50-60ᵒC before use. Place the columns in a fresh 1.5 ml centrifuge tube. Add 60 μl of preheated elution buffer and let it stand for 1-5 minutes at room temperature.

      - Centrifuge the Column at 11000 rpm for 1 minute.

      - Store the Elute containing Extracted NA at -200C till further processing.

- Quality Control Procedure

Add the Internal control in each sample during the extraction. If not added during the Extraction step, IC can be added in Master Mix preparation during PCR Setup.

- Interfering Substances

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

Lysed and contaminated samples can lead to erroneously high values.

Improper storage conditions and delayed transport of samples often give a low yield or no yield of NA and often leads to false negative results.

- References

- NucleoDx Nucleic Acid Extraction Kit Handbook