SOP - 002

PROCEDURE FOR

HIV, HBV, HCV QUALITATIVE BY PCR

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	9

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	12

15.	Interfering Substances	12

16.	Result Calculation	13

17.	Biological Reference Interval	13

18.	Reportable Interval	13

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	14

22.	Potential Source of Variability	14

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | Ct | Cycle Threshold |
| 2 | DNA | Deoxyribonucleic Acid |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HBV | Hepatitis B Virus |
| 6 | HCV | Hepatitis C Virus |
| 7 | HIV | Human Immunodeficiency Virus |
| 8 | IC | Internal Control |
| 9 | MI | Microbiology |
| 10 | MIC | Minimum Inhibitory Concentration |
| 11 | NA | Nucleic Acid |
| 12 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 13 | NTC | No-template Control |
| 14 | PCR | Polymerase Chain Reaction |
| 15 | QC | Quality Control |
| 16 | RNA | Ribonucleic Acid |
| 17 | rpm | Revolutions per minute |
| 18 | RT-PCR | Reverse Transcriptase PCR |
| 19 | SOP | Standard Operating Procedure |
| 20 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - The purpose and scope of this document is to describe the procedure for the detection of Human Immunodeficiency Virus - 1/2 RNA, Hepatitis B Virus DNA, and Hepatitis C RNA virus in plasma, serum or whole blood specimen by Real Time PCR.

- Scope

This procedure applies to all staff of the Molecular Biology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Qualitative PCR tests can detect the presence of HIV, HBV, and HCV in blood samples. These tests are used to identify the viruses, but they don't measure the quantity or viral load. Instead, they simply indicate if the virus is present or not.

- Test Principle

Pathogen detection by the polymerase chain reaction (PCR) is based on the amplification of specific regions of the pathogen genome. In real-time PCR the amplified product is detected via fluorescent dyes. These are usually linked to oligonucleotide probes that bind specifically to the amplified product. Monitoring the fluorescence intensities during the PCR run (i.e., in real-time) allows the detection and quantitation of the accumulating product without having to re-open the reaction tubes after the PCR run.

- Specific Performance Characteristics

- PathoPlex HIV-1/2, HBV, HCV Qualitative RT PCR kit – NeoDx Biotech Labs.

      - Linearity - Minimal Detection Limit:

          - HIV1/2 - 20 IU/ml

          - HBV – 32 IU/ ml

          - HCV – 22.7 IU/ml

      - Sensitivity: 99.8%

      - Specificity: 99.8%

- Type of Sample

Whole Blood EDTA / Plasma / Serum.

Collect 2.0 ml blood in Lavender top (K2- EDTA). Thoroughly mix the blood with the additives by gently inverting the tube four or five times. Whole blood should be separated into plasma and cellular components by centrifugation for 20 minutes at 3500 rpm within 4 hours.

- Note: Do not use heparin as an anticoagulant as it is an inhibitor of PCR.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

Nucleic Acid Extraction kit

PCR Amplification Kit

Biosafety Cabinet

Laminar Air flow

Calibrated variable micropipettes, Sterile pipette filter tips (aerosol free)

Vortex mixer

Mini centrifuge with rotor for 1.5 mL / 2 mL reaction tubes

Micro Centrifuge 0.2 mL /1.5 mL

Real Time PCR machine

PCR Tubes / strips (0.2 mL)

Cooling block (96 x 0.2 mL tubes)

1.5 mL / 2mL Centrifuge tubes

1.5 ml / 2ml Centrifuge tubes stand

Cooling block (1.5 ml tubes)

Sterile powder free gloves

Lab coats, Facemask, Head cap

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Equipment calibrations are done on an annual basis by an NABL accredited calibration agency or by specific guidelines of equipment manufacturers.

- Test Procedure

Pre-Assay Instructions

- Nucleic Acid Extraction (Isolation): Viral Nucleic Acid purification is carried out according to the instructions in NucleoDx Blood Nucleic Acid Extraction Kit.

    - Recommended sample volume for extraction and elution are as follows:

| Name of the Nucleic Acid Extraction Kit | Recommended Sample volume | Recommended Final Elution volume |
| --- | --- | --- |
| NucleoDx Blood Nucleic Acid Extraction Kit | 200 µL | 60 µL |

    - NeoDx HIV Real time RT-PCR Kit Contents (50 Reactions)

| Contents | 100 Rx | Storage |
| --- | --- | --- |
| Master Mix | 2*750 µL | -20ᵒC |
| Probe and Primer Mix | 150 µL | -20ᵒC |
| Positive Test Control | 100 µL | -20ᵒC |
| Nuclease Free Water | 1000 mL | -20ᵒC |

    - Fluorescent Probe Details

| Target | Reporter | Quencher |
| --- | --- | --- |
| HIV-1/2 | CY5 | BHQ2 |
| HBV | FAM | BHQ1 |
| HCV | HEX/VIC | BHQ2 |
| Human Beta Actin | TEXAS RED / ROX | BHQ2 |

    - Reaction Mixture Preparation

| Reagents | Volume per Reaction (35 µL) | Volume per 20 Reaction (17 µL) |
| --- | --- | --- |
| Master Mix | 15 µL | 300 µL |
| Primer and Probe Mix | 1.5 µL | 30 µL |
| Nuclease Free Water | 3.5 µL | 70 µL |
| Sample/ Control/ QS | 15 µL | 15 µL/ reaction |
| Volume | 35 µL | 20 µL Reaction Mix + 15 µL of sample |

    - Cycling Conditions

| Steps | No. of Cycles | Temperature (°C) | Time |
| --- | --- | --- | --- |
| Reverse Transcriptase | 1 | 45 | 20 min |
|  | 1 | 95 | 2 min |
| PCR cycling | 40 | 95 | 15 sec |
|  |  | 58 | 30 sec |

- NOTE: Fluorescence data should be collected during the 58°C/30sec extension step.

Preparation

      - Proper disinfection practices of the Biosafety Cabinet, Laminar Flow, Pipettes and other minor supplies used in the process.

      - Bring all the reagents and specimens to the required temperature before beginning the test. The reagents should not be subjected to frequent temperature fluctuations.

      - Make sure that the cooling block is pre-cooled to 2-8°C.

      - Do not use kit components beyond the expiration date.

Procedure

- The procedural sequence of reagent addition should be strictly adhered to avoid any discrepant results.

      - Before each use, all reagents need to be thawed completely, mixed (by repeated pipetting), and centrifuged briefly.

      - Calculate the number of reactions for each experiment including all controls with one excess reaction volume in the reaction cocktail to accommodate pipetting errors.

      - Assay should be run along with positive and negative controls and standards.

      - Prepare the reaction mix in a 1.5/2ml tube for the calculated number of samples referring to the Reaction Mixture table in Master Mix Preparation room.

      - Dispense 20 µl of reaction mix in each tube or strip.

      - Add 15 µl of nuclease free water to NTC well.

      - Carefully add 15 µl of nucleic acid extract to the wells in template addition room.

      - Add 15 µl of Quantitative Standard in the designated wells in Positive Control addition area in the designated wells.

      - Add 2 µl of Positive Test Control and make up the volume by adding 13 µl of nuclease free water.

      - Seal the plate carefully and spin down and set the program in RT PCR instrument and run the PCR reaction in the PCR Equipment – Quant Studio 5.

- Analysis should be performed independently using threshold settings after analysis of the Standard Curve for number of IU/ µl. Each sample is analysed along with an internal control and NTC to verify the sample run. After the completion of the run the samples are analysed and the values are compared to standard/ positive and negative control for qualitative result. The threshold should be set manually just above the background signal of the negative controls and negative samples. The threshold value will usually be in the range of 0.05–0.25.

- Quality Control Procedure

Internal Quality Control Procedure

      - A negative control/NTC and a positive template control is included in each run. This ensures that the laboratory is capable of detecting the entire range of values of the analyte in the clinical samples. QC must be acceptable (in control) before the testing and reporting of results. Any test results obtained when QC is unacceptable or not performed are invalid and must be repeated. The person performing testing must evaluate QC results and accordingly determine validity of the test run.

      - All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – MHL EQAS

- Interfering Substances

- Several factors may interfere and affect the Real-Time PCR amplification:

Elevated levels of bilirubin (≤15 mg/dl) and lipids (≤800 mg/dl) and hemolytic samples do not influence the system.

Heparin (≤10 IU/ml) affects the PCR. Samples that have been collected in tubes containing heparin as an anticoagulant should not to be used. Also, samples of heparinized patients must not be used.

Samples from delayed transport conditions.

Contamination of reagents – can show amplification of non-template control (NTC) which is usually molecular grade water added in place of sample.

- Result Calculation

- After the completion of the run, go to the analysis window. The run profile shows positive and negative samples based on the Ct values for HIV, HBV, HCV target Flourophore and IC (ROX) Flourophore. Make sure that IC for each sample is picked almost at the same Ct which indicates high fidelity of real time run.

Interpretation of Qualitative Results

| Fluorophore | Result | Interpretation |
| --- | --- | --- |
| TEXAS RED/ROX - IC | + | Valid Run |
|  | - | Invalid Run (Sample to be retested from first step of extraction and recheck sample acceptance criteria) |
| CY5 – HIV1/2 | + | HIV – 1/2 detected |
| FAM - HBV | + | HBV detected |
| HEX / VIC - HCV | + | HCV detected |

- NOTE: Absence of sigmoid curve and Ct for Sample/Positive control/Negative control – If both target and IC are not amplified – There could be problem in the extraction of nucleic acid from sample or there could be problem in the PCR real time reaction reagents. This can be cross checked by re extraction of sample or setting up real time with another valid real time kit.

Interpretation of Quantitative Results

- For the analysis of samples, in each run include four standards and one NTC along with unknown samples. Each sample is analysed along with an internal control and NTC to verify the sample run. After the completion of the run the samples are analysed. Interpret the values for unknown samples - there should be no amplification in negative control.

- Standard Curve Control Parameter

- Biological Reference Interval

- Not detected.

- Reportable Interval

- The reporting of results of HIV, HBV, HCV nucleic acid detection is as per turnaround time (TAT). The turnaround time for reporting is 72 hours from the time of receiving samples to lab. In case of an urgent request from clinician or referral lab, the sample is considered under STAT (Special Turnaround time).

- Procedure when Results are not in Range

| Observation | Possible cause | Recommendation |
| --- | --- | --- |
| Target growth curve in extraction control or NTC | Cross contamination of samples occurred during nucleic acid extraction procedures or assay setup. | Repeat the assay with stricter adherence to good laboratory guidelines. |
| No signal for IC target in sample | Presence of RT-PCR inhibitors from clinical specimens | Repeat the test from fresh sample following good laboratory practices and strict adherence to assay protocol |
|  | Improper assay set up and execution |  |

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- Based on the amplified target, the sample is concluded to have the HIV, HBV, HCV target nucleic acid material.

- Potential Source of Variability

If the nucleic acid recovery in the extraction procedure is poor, the Ct value of target would be high as it picks up at late Cts.

Haemolysed samples and delayed transport samples often give a low yield or no yield of target and often leads to false negative results.

It is important to pipet the indicated quantities and mix well after each reagent addition. Check pipettes regularly.

Strict compliance with the user manual is required for optimal PCR results.

Attention should be paid to expiration dates printed on the kit box and labels of all components. Do not use expired components.

Although rare, mutations within the highly conserved regions of the viral genome covered by the kit’s primers and/or probe may result in under quantitation or failure to detect the presence of the virus in these cases. Validity and performance of the assay design are revised at regular intervals.

The presence of PCR inhibitors may yield false negative or invalid results.

- References

- PathoPlex HIV1, HBV, HCV NeoDx Qualitative PCR Kit Handbook