SOP - 002

PROCEDURE FOR HIV ELISA

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	10

10.	Equipment & Reagents	10

11.	Safety	10

12.	Calibration Procedure	11

13.	Test Procedure	11

14.	Quality Control Procedure	13

15.	Interfering Substances	13

16.	Result Calculation	13

17.	Biological Reference Interval	14

18.	Reportable Interval	14

19.	Procedure when Results are not in Range	14

20.	Critical Alert Values	14

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AIDS | Acquire Immunodeficiency Syndrome |
| 2 | CSD | Customer Service Department |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | ELISA | Enzyme-Linked Immunosorbent Assay |
| 5 | EQAS | External Quality Assurance Scheme |
| 6 | HIV | Human Immunodeficiency Syndrome |
| 7 | HLA | Human Leukocyte Antigen |
| 8 | NACO | National AIDS Control Organization |
| 9 | OD | Optical Density |
| 10 | rpm | Rotations per minute |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |
| 12 | WHO | World Health Organisation |

- Objective

  - This procedure describes the step-by-step tasks involved in the processing of the HIV test by the ELISA method.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

HIV ELISA is an in-vitro diagnostic test for qualitative detection of antibodies against Human Immunodeficiency Virus Type – 1&2 virus in patient serum/ plasma. Human immunodeficiency Virus (HIV) is a member of the Retrovirus family that can cause Acquired Immunodeficiency Syndrome (AIDS), a condition in humans in which the immune system begins to fail. There are 2 strains of HIV known to exist, HIV-1 and HIV-2. HIV-1 is more virulent, relatively easily transmitted, and is the cause of the majority of HIV infections globally. HIV-2 is less transmitted and is largely confined to West Africa and Europe.

- Test Principle

ELISA HIV 1+2 is based on indirect ELISA using a solid phase prepared with the mixture of synthetic peptides i.e., gp41, C terminus of gp120 for HIV1 and gp36 for HIV2. Detection is carried out using anti-human antibodies conjugated with horseradish peroxidase (HRPO).

The performance of the test includes the following reaction steps:

      - The specimens to be tested along with the controls are added to the wells. If antibodies to HIV-1 or HIV-2 are present in the specimen, they will bind to the antigen coated on the wells.

      - After a wash, peroxidase labeled anti-human immunoglobulins (conjugate) is added to the wells which in turn binds to the specific antibodies captured on the solid phase.

      - The bound antigen-antibody complex is identified by addition of a substrate.

      - After the reaction has been stopped, the absorbance values are read. The intensity of the color developed is proportional to the amount of anti- HIV 1+2 antibody bound to the solid phase.

- Specific Performance Characteristics

| Sample Panel | Indian Panel | European Panel | African Panel | Worldwide Panel |
| --- | --- | --- | --- | --- |
| Sensitivity | 100% | 100% | 100% | 100% |
| Specificity | 99.5% | 100% | 100% | 100% |

      - CV%: < 10%

      - Accuracy: > 90%

      - Analytical Measurement Range (Limit of Quantitation): Not applicable

      - Clinical Reportable Range: Not applicable

- Type of Sample

- Serum or Plasma must be used for testing, separated by centrifugation at the earliest from the vacutainer to avoid hemolysis. Should a delay in testing occur, store the serum or plasma at 2-8°C for up to a week, or frozen at -20°C until use. Do not use hemolysed serum. Samples containing aggregates must be centrifuged prior to use. Fresh serum / plasma samples are preferred. Undiluted serum can be stored at 2-8°C for a week. Frozen specimen should be completely thawed and centrifuged. The test should be performed on clear supernatant serum or plasma. Heat inactivated, hemolyzed and icteric hyperlipemic samples may yield erroneous results.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques. Consent form as approval for testing is necessary before testing. Testing will be done as per NACO guidelines.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

Antigen coated microwells

Sample diluent

Negative control (ready to use)

Positive control (ready to use)

Conjugate concentrate, Conjugate Diluent (10 µL of Conjugate concentrate in 1 ml of Conjugate Diluent for 8 tests)

Wash buffer – 25x (2ml of wash buffer in 48ml of distilled water)

TMB Substrate concentrate and diluent (0.8ml of substrate to 0.8ml of deionized water for 8 tests)

Stopping solution

Black plate cover

Strip sealers

Absorbent paper

Disposable gloves and protective glasses

Pipettes capable of delivering 10, 50 and 100 µl volumes

Disposable tips

Graduated cylinder

Waste disposable container

Timer

ELISA reader

Automatic / semi-automated washing system

Distilled / deionized water

Sodium hypochlorite solution for disposal of waste

Centrifuge

- Safety

The use of disposable gloves and proper biohazardous clothing is strongly recommended while running the test.

In case there is a cut or wound in hand, do not perform the test.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in-vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of an accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Centrifuge and Pipette calibrations are done on an annual basis by an NABL accredited calibration agency. Preventive Maintenance of ELISA Reader is done by authorised persons.

- Test Procedure

Precautions

      - Before use, wait for all reagents to stabilize to room temperature.

      - DO NOT use kit components beyond expiration date as indicated on the labels.

      - DO NOT mix reagents between different lots as these are optimized for individual batch performance.

      - All specimens and controls should be considered potentially infectious and discarded appropriately.

      - Use a new tip for each control and specimen to avoid microbial contamination of the reagents.

      - Use clean glassware or plastic ware rinsed with distilled water for reagent preparation.

      - Use only distilled / deionized water to reconstitute the washing solution.

      - DO NOT interchange the caps of the reagents.

      - Run the recommended number of positive and negative controls in each assay to ensure validity of the kit or at least one control of each type.

      - Use Clear serum. Particulate matter should be removed by centrifugation.

      - DO NOT allow the microplate wells to dry between the end of washing and the addition of the reagent.

      - Incubation time should not vary by more than ± 2 minutes.

      - After using the required number of strips, the rest of the strips along with activated silica gel should be kept in sealed condition within a ziplock bag.

      - DO NOT expose the color reagent to light.

      - Level the microwell strips in the holder before reading the results. Wipe the bottom of the wells with a clean tissue to remove any dust or moisture.

      - The test is for in-vitro diagnostic use and should be performed and interpreted by a competent or trained person only.

Procedure

      - Bring all the reagents and test specimens to room temperature and shake well before use.

      - Define the sample/control distribution and identification plan. In each run, assign one well for the Blank (A1), 2 wells for the HIV negative controls (B1, C1) and 3 wells for the anti-HIV positive control (D1, E1, F1) for the reagent verification of First run. One well each for positive and negative control may be used for further runs.

      - Break the number of required wells for a run. Wrap the remaining unused wells tightly in zip-lock pouch with desiccant and return it to 2-8°C immediately.

      - Add 100 µL of the sample diluents to all the wells used for sample testing only and 10 µL of sample as per the arrangement.

      - Add 100 µL of the HIV controls in the respective wells.

      - Mix well, cover the wells with the strip sealers and incubate for 30 minutes at 37°C.

      - Remove the sealer. Discard/aspirate the contents of the well into the waste disposal container. Add a minimum of 350 µL of washing solution to each well. Aspirate again after 30 seconds of soak time. Repeat the washing step 5 times (invert the plate and tap it on absorbent pad to remove the remaining washing solution).

      - Add 100 µL conjugate into all the wells. Cover the wells with the strip sealers and incubate for 30 minutes at room 37°C.

      - Repeat washing step

      - Add 100 µL of the Substrate to all wells. Cover the plate with the black cover provided and allow the reaction to develop in the dark for 30 minutes at room temperature (20-30°C).

      - Add 100 µL of stopping solution to all the wells, Homogenize. After the addition of the stopping solution the blue color of the substrate turns to yellow (for positive samples) or remains colorless (for negative samples).

      - Wipe the plate bottom and read the OD at 450 nm in an ELISA Reader.

- Quality Control Procedure

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – IAMM VIROLOGY EQAS, MHL EQAS

Assay Validity -

      - Blank Value: Absorbance value (OD) of blank should be less than 0.1.

      - Positive Control OD: >0.5

      - Negative Control OD: <0.15

- Interfering Substances

FROZEN SAMPLES:

- Allow the sample to thaw in a vertical position in the rack. Do not shake the sample. This allows particles to settle to the bottom. If a centrifuge is available, the sample can be centrifuged at 3500 r.p.m. for 15 min.

THICK OR VISCOUS SAMPLES:

- Whenever possible, clear specimens should be used. However viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3500 r.p.m. for 15 min. and retested on a fresh device to avoid inconsistent results.

- Result Calculation

Calculate the Mean of the measured absorbance values (O.D.).

For the HIV 1+2 Negative Control (NCx) and Positive Control (PCx)

Cut –Off value (COV) = Mean OD of PCx+ Mean OD of NCx6

Cut –Off value (COV) for further runs = OD of PC+OD of NC6

NON-REACTIVE

      - Non – Reactive for HIV Antibodies 1/2: samples with an OD less than the cut-off.

REACTIVE

      - Reactive for HIV Antibodies 1/2: samples with an OD equal to or more than the cut-off.

INVALID

      - The test is considered invalid if the test validity is not achieved. Repeat testing   with proper corrective actions is to be done.

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 24 Hours

- Procedure when Results are not in Range

The HIV ELISA test is done with HIV ELISA (J’Mitra) kit for all the tests billed under HIV ELISA test.

If the sample is Reactive by HIV ELISA method, then as per NACO guidelines:

      - a second rapid kit method,

      - and a third method by RETROQUIC are used to confirm the results.

- Critical Alert Values

- N/A

- However, all reactive HIV tests should be tested by three methods as per NACO guidelines and released after proper counselling. Pre-test counselling is assumed to be done by the doctor requesting the test. Post-Test Counselling is mandatorily done for all reactive patients, after taking treating doctor’s guidance by the lab doctors through CSD team or by the treating physician himself.

- Laboratory Clinical Interpretation

- The availability of rapid HIV testing in the hospital for emergency pre-operative screening provides a useful tool for clinicians that can inform treatment and disposition decision making for patients, as well as serves to confirm the reactive results by other methods. Although no formal studies have been done yet to evaluate the impact of rapid testing for practice, significant anecdotal experience is being gathered from literature that it has been involved in screening, risk-targeted testing, or testing based on clinical suspicion.

- Potential Source of Variability

The kit works best when used with fresh samples.

Optimum test performance depends on strict adherence to the test procedure as described in this manual. Any deviation from test procedure may lead to erratic results.

Some samples show cross reactivity for HIV antibodies. Following factors are found to cause false positive HIV antibody test results: Naturally occurring antibodies, Passive immunization, Leprosy, Renal Disorders, Tuberculosis, Myco-bacterium avium, Herpes simplex, Hypergammaglobulinemia, Malignant neoplasms, Rheumatoid arthritis, Tetanus vaccination, Autoimmune diseases, Blood Transfusion, Multiple myeloma, Haemophilia, Heat treated specimens, Lipemic serum, Anti-nuclear antibodies, T-cell leukocyte antigen antibodies, Epstein Barr virus, HLA antibodies and other retroviruses.

This is only a screening test. All samples detected reactive must be confirmed by using HIV Western Blot. Therefore, for a definitive diagnosis, the patient’s clinical history, symptomatology as well as serological data, should be considered. The results should be reported only after complying with the above procedure.

- References

HIV MICROLISA – J’Mitra kit insert

NACO Guidelines

- HIV Testing Protocol