SOP - 003

PROCEDURE FOR

HBsAg HEPACARD (RAPID) TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	10

15.	Interfering Substances	11

16.	Result Calculation	11

17.	Biological Reference Interval	11

18.	Reportable Interval	12

19.	Procedure when Results are not in Range	12

20.	Critical Alert Values	12

21.	Laboratory Clinical Interpretation	12

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | EDTA | Ethylenediaminetetraacetic Acid |
| 2 | EQAS | External Quality Assurance Scheme |
| 3 | HBsAg | Hepatitis B Surface Antigen |
| 4 | HBV | Hepatitis B Virus |
| 5 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 6 | PCR | Polymerase Chain Reaction |
| 7 | RF | Rheumatoid Factor |
| 8 | rpm | Revolutions per minute |
| 9 | SE | Serology |
| 10 | SOP | Standard Operating Procedure |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - This procedure describes the step-by-step tasks involved in the process of the HBsAg test using a rapid method.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

HEPACARD is a visual, rapid, sensitive and accurate one step immunoassay for the qualitative detection of Hepatitis B Surface Antigen (HBsAg) in Human Serum or Plasma. The assay is intended to be used as an aid in the recognition and diagnosis of acute infections and chronic infectious carriers of the Hepatitis B Virus (HBV).

- Test Principle

HEPACARD is a one-step immunoassay based on the antigen capture, or “sandwich” principle. The method uses monoclonal antibodies conjugated to colloidal gold and polyclonal antibodies immobilized on a nitrocellulose strip in a thin line. The test sample is introduced to and flows laterally through an absorbent pad where it mixes with the signal reagent. If the sample contains HBsAg, the colloidal gold-antibody conjugate binds to the antigen, forming an antigen-antibody-colloidal gold complex. The complex then migrates through the nitrocellulose strip by capillary action. When the complex meets the line of immobilized antibody (Test line) “T”, the complex is trapped forming an antibody-antigen-antibody colloidal gold complex. This forms a pink band indicating the sample is reactive for HBsAg. To serve as a procedural control, an additional line of anti-mouse antibody (Control line) “C”, has been immobilized at a distance from the test line on the strip. If the test is performed correctly, this will result in the formation of a pink band upon contact with the conjugate.

- Specific Performance Characteristics

Minimal Detection Limit - HEPACARD can detect Hepatitis B Surface Antigen in serum or plasma at a concentration of as low as 0.5 ng/ml. All the eleven HBsAg subtypes can be detected positive with HEPACARD.

| Sample panel | CMC Panel | In-house Evaluation |
| --- | --- | --- |
| Sensitivity | 100% | 99.8% |
| Specificity | 100% | 99% |

- Type of Sample

- Serum or Plasma must be used for testing. Should a delay in testing occur, store the sample at 2-8°C. Samples can be stored for up to a week. Do not use hemolysed, icteric, lipemic serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

HEPACARD Test Device

Buffer Solution

Sample Dropper

Centrifuge to separate serum or plasma

- Safety

The use of disposable gloves and proper biohazardous clothing is strongly recommended while running the test.

In case there is a cut or wound in hand, do not perform the test.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent person only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed off in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material are ingested or come in contact with skin puncture or wounds.

Do not open the foil pouch to remove the product until it attains room temperature and you are ready to perform the test.

Take out the Cards from the pouch just before performing the test to avoid denaturation of antisera due to atmospheric exposure.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Centrifuge calibration is done on an annual basis by an NABL accredited calibration agency.

- Test Procedure

Bring the required number of HEPACARD foil pouches and specimen to room temperature prior to testing.

Take out the HEPACARD device from the foil pouch. In case of 5 test pouch packing, tightly reseal the pouch containing balance devices with clamp and rod so that devices are protected from moisture. Ensure that the pouch is perfectly sealed otherwise the devices will deteriorate in condition leading to erratic results.

Label the test card with the patient’s name or identification number.

Add 2 drops (70 µl) of human serum/plasma specimen into the sample well using the dropper provided (use separate dropper/microtip for each specimen).

Allow reaction to occur during the next 20 minutes.

READ RESULTS AT 20 MINUTES.

- Quality Control Procedure

Each testing device has an inbuilt Control Dot to confirm the validity of the device. The appearance of the control dot after completing the procedural test indicates a valid test.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – IAMM VIROLOGY EQAS, MHL EQAS

- Interfering Substances

If testing within 3 days is not possible, specimen should be stored frozen at - 20°C.

- Specimen containing visible precipitates or cloudy specimens may give inconsistent test results. Such specimens should be clarified prior to testing by high-speed centrifugation i.e. 3500 rpm for 15 minutes before testing. Haemolysed or lipemic specimens may give erroneous results.

- Result Calculation

POSITIVE

      - Appearance of pink-coloured lines, one each in test region “T” and control region “C” indicate that the sample is POSITIVE for HBsAg. A difference of intensity in colour may occur between the Test line & Control line depending on the concentration of the HBsAg in the serum but this does not affect interpretation of the result. If the concentration of HBsAg in the sample is very high, only test line may be observed. This is due to Hook’s effect. Such samples should be diluted at 1:10 or 1:20 in normal saline & the test re-run again. Diluted samples should show both control & test lines. In case the control line still does not appear or is faint, dilute the sample further.

NEGATIVE

      - Appearance of one distinct pink line, in the control region “C” only, indicates that the sample is “NEGATIVE” for HBsAg.

INVALID

- When control line does not appear on the membrane, the test should be treated as invalid. This may be because of the following reasons:

      - Improper storage at temperature other than the recommended temperature.

      - Wrong procedure.

      - Long atmospheric exposure of the test device after opening the pouch. The test should be repeated using a new HEPACARD and test sample.

- Biological Reference Interval

- Negative

- Reportable Interval

- 4 Hours

- Procedure when Results are not in Range

- All samples are reconfirmed with a second device and correlated with clinical history and other laboratory investigations before releasing the result.

- Critical Alert Values

- N/A

- Laboratory Clinical Interpretation

- The antigenic determinant of the HBsAg protein moiety is antigenically heterogenous and it determines specific HBV serotypes and provides a basis for immunodetection. The principal antigenic determinant is “a” which is common to all HBV serotypes. In addition, two pairs of subspecific determinants have been identified, d/y & w/r, which are apparently mutually exclusive. Four antigenic combinations are therefore possible: adw, adr, ayw and ayr.

- Potential Source of Variability

The HEPACARD is for in-vitro diagnostic use only.

The test should be used for the detection of HBsAg in serum or plasma only and not in other body fluids.

This is only a Screening test. All reactive samples should be confirmed by confirmatory tests – HBV PCR and other serological markers of HBV infection. Therefore, for a definitive diagnosis, the patient’s clinical history, symptomatology as well as serological data, should be considered. The results should be reported only after complying with the above procedure.

Additional follow up testing using available clinical methods (along with repeat HEPACARD test) is required, if HEPACARD test is non-reactive with persisting clinical symptoms.

False positive results can be obtained due to the presence of other antigens or elevated levels of RF factor. This occurs in less than 1% of the samples tested.

- References

- HEPACARD – J’Mitra kit insert