SOP - 004

PROCEDURE FOR

HCV RAPID TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	10

15.	Interfering Substances	11

16.	Result Calculation	11

17.	Biological Reference Interval	12

18.	Reportable Interval	12

19.	Procedure when Results are not in Range	12

20.	Critical Alert Values	12

21.	Laboratory Clinical Interpretation	12

22.	Potential Source of Variability	12

23.	References	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ALT | Alanine Aminotransferase |
| 2 | CDC | Centre for Disease Control |
| 3 | EDTA | Ethylenediaminetetraacetic Acid |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | HCV | Hepatitis C virus |
| 6 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 7 | NANBH | Non-A Non-B Hepatitis |
| 8 | PCR | Polymerase Chain Reaction |
| 9 | RF | Rheumatoid Factor |
| 10 | rpm | Revolutions per minute |
| 11 | SE | Serology |
| 12 | SOP | Standard Operating Procedure |
| 13 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - This procedure describes the step-by-step tasks involved in the process of the HCV test using a rapid method.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

The 4th Generation HCV TRI-DOT is a rapid, visual, sensitive and qualitative in vitro diagnostic test for the detection of antibodies to Hepatitis C Virus in human serum or plasma. The 4th Generation HCV TRI-DOT has been developed and designed with increased sensitivity for core and NS3 antibodies using a unique combination of modified HCV antigens. They are for the putative core (structural), protease/helicase NS3 (non-structural), NS4 (nonstructural) and replicase NS5 (non-structural) regions of the virus in the form of two test dots “T1” & “T2” to provide a highly sensitive and specific diagnostic test.

- Test Principle

HCV antigens are immobilized on a porous immunofiltration membrane. Sample and the reagents pass through the membrane and are absorbed into the underlying absorbent pad. As the patient's sample passes through the membrane, HCV antibodies, if present in serum/plasma, bind to the immobilized antigens. In the subsequent washing step, unbound serum/plasma proteins are removed. In the next step, the protein-A conjugate is added which binds to the Fc portion of the HCV antibodies to give distinct pinkish purple dot against a white background at the test region (“T1” &/or “T2”). At the control region (“C”) a “Built-in Quality Control Dot” has been devised to confirm the proper functioning of the device, reagent and correct procedural application.

- Specific Performance Characteristics

| Sample panel | Indian Panel | WHO Evaluation |
| --- | --- | --- |
| Sensitivity | 100% | 100% |
| Specificity | 99.8% | 98.9% |

- Type of Sample

- Serum or Plasma must be used for testing. Should a delay in testing occur, store the sample at 2-8°C. Samples can be stored for up to a week. Do not use hemolysed, icteric, lipemic serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

HCV Test Device

Sample Dropper

Timer

- Safety

The use of disposable gloves and proper biohazardous clothing is strongly recommended while running the test.

In case there is a cut or wound in hand, do not perform the test.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in-vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

Protein-A Conjugate and Buffer Solution contain Sodium Azide as a preservative. If these materials are to be disposed of in a sink or other common plumbing systems, flush with generous amounts of water to prevent accumulation of potentially explosive compounds. In addition, consult the manual guideline "Safety Management No. CDC-22", Decontamination of Laboratory Sink Drains to remove Azide salts" (Centre for Disease Control, Atlanta, Georgia, April 30, 1976.)

- Calibration Procedure

- Centrifuge calibration is done on an annual basis by an NABL accredited calibration agency.

- Test Procedure

Allow the test device and specimen to equilibrate to room temperature (15 - 300c) prior to testing.

Add 3 drops of Buffer Solution to the centre of the device.

Hold the dropper vertically downwards and add 1 drop of patient's sample (50 µl serum or plasma) using the sample dropper provided. (use a separate sample dropper for each specimen to be tested).

Add 5 drops of Buffer Solution.

Add 2 drops of Protein-A Conjugate.

Add 5 drops of Buffer Solution.

READ RESULT IMMEDIATELY.

- IMPORTANT: It is important to allow each solution to soak in the test device before adding the next solution.

- Quality Control Procedure

Each testing device has an inbuilt Control Dot to confirm the validity of the device. The appearance of the control dot after completing the procedural test indicates a valid test.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – IAMM VIROLOGY EQAS, MHL EQAS

- Interfering Substances

- Though HCV TRI-DOT works best when used with fresh samples, however the frozen or viscous samples can also perform well if the following instructions are strictly adhered to:

Frozen Samples

      - Allow the sample to thaw in a vertical position in the rack. Mix the sample thoroughly. If particles are seen, allow them to settle at the bottom or if a centrifuge is available, the sample can be centrifuged at 3500 rpm for 15 minutes.

      - Insert the dropper just below the top surface of the sample and withdraw one drop of the sample.

Thick or Viscous Samples

      - Whenever possible, clear specimens should be used. However, viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3500 rpm for 15 minutes and retested on a fresh device to avoid inconsistent results.

- Result Calculation

NON-REACTIVE

      - Appearance of only one dot at the control region “C” indicates that the sample is NON-REACTIVE for antibodies to HCV.

REACTIVE

      - Appearance of two dots, one at the control region “C” & the other at the test region “T1”, indicates that the sample is REACTIVE for antibodies to HCV.

      - Appearance of two dots, one at the control region “C” & other at the test region “T2” indicates that the sample is REACTIVE for antibodies to HCV.

      - Appearance of all three dots, one each at “C”, “T1” & “T2” regions, indicates that the specimen is REACTIVE for antibodies to HCV.

INVALID

      - If no dot appears after the completion of test, either with a clear background or with a complete pinkish/purplish background the test indicates an ERROR.

      - This may indicate a procedural error or deterioration of the specimen/reagents or particulate matter in the specimen. The specimen should be retested on a fresh device (Refer sample / specimen processing).

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 4 Hours

- Procedure when Results are not in Range

- HCV Tridot is done for tests billed under HCV - Rapid. Invalid results are to be repeated with a new test device. Reactive results are also to be repeated with a new device.

- Critical Alert Values

- N/A

- Laboratory Clinical Interpretation

- Hepatitis C Virus was identified in 1989 as the main aetiological agent of non-A, non-B hepatitis (NANBH) accounting for greater than 90% of post-transfusion hepatitis cases. HCV is a spherical virus of about 30- 60 nm in diameter with single positive stranded RNA and is related to the family flaviviridae. It is considered to be the major cause of acute chronic hepatitis, liver cirrhosis and hepatocellular carcinoma throughout the world. The test for antibodies to HCV was proved to be highly valuable in the diagnosis and study of the infection, especially in the early diagnosis of HCV after transfusion. The diagnosis of Hepatitis C can be easily made by finding elevated serum ALT levels and presence of antibodies to HCV in serum/plasma.

- Potential Source of Variability

The 4th Generation HCV TRI-DOT detects anti-HCV in human serum or plasma and is only a screening test. All reactive samples should be confirmed by supplemental assays like RIBA. Therefore, for a definitive diagnosis, the patient’s clinical history, symptomatology as well as serological data, should be considered. The results should be reported only after complying with the above procedure.

The assay is only validated for serum and plasma from individual bleeds and not for pools of serum or plasma or other body fluids.

A non-reactive result does not exclude the possibility of infection with HCV.

It should be noted that repeated false reactive results may occur due to non-specific binding of the sample to the membrane.

Patients with auto-immune liver diseases, Renal disorders may show falsely reactive results.

The kit works best when used with fresh samples and when all the kit components are at room temperature (20-25°C). Samples which have been frozen and thawed several times contain particulates which can block the membrane, hence resulting in improper flow of reagents and high background colour which may make the interpretation of results difficult.

Optimum test performance depends on strict adherence to the test procedure as described in this SOP. Any deviation from test procedure may lead to erratic results.

- References

- HCV TRIDOT – J’Mitra kit insert