SOP - 005

PROCEDURE FOR DENGUE

NS1, IgG & IgM RAPID TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	4

RECORD OF AMENDMENTS	5

DISTRIBUTION LIST	6

LIST OF ABBREVIATIONS	7

1.	Objective	8

2.	Scope	8

3.	Responsibilities	8

4.	Definition	8

5.	Test Principle	8

6.	Specific Performance Characteristics	9

7.	Type of Sample	9

8.	Patient Preparation	9

9.	Type of Container & Additives	9

10.	Equipment & Reagents	9

11.	Safety	9

12.	Calibration Procedure	10

13.	Test Procedure	10

14.	Quality Control Procedure	11

15.	Interfering Substances	11

16.	Result Calculation	11

17.	Biological Reference Interval	12

18.	Reportable Interval	12

19.	Procedure when Results are not in Range	12

20.	Critical Alert Values	12

21.	Laboratory Clinical Interpretation	12

22.	Potential Source of Variability	12

23.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | CDC | Centre for Disease Control |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | ELISA | Enzyme-Linked Immunosorbent Assay |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | IgG | Immunoglobulin G |
| 6 | IgM | Immunoglobulin M |
| 7 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 8 | NS1 | Non-Structural protein 1 |
| 9 | rpm | Revolutions per minute |
| 10 | SE | Serology |
| 11 | SOP | Standard Operating Procedure |
| 12 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- This procedure describes the step-by-step tasks involved in the process of the Dengue NS1 antigen test as well as Dengue IgG & IgM using a rapid method.

- Scope

- This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- The immunochromatography technique is a rapid method that is used to detect the Dengue Virus NS1 antigen and IgG / IgM antibodies simultaneously, with a high degree of sensitivity and specificity.

- Test Principle

- Dengue NS1 Ag + Ab Day 1 rapid test contains 2 test devices (Dengue NS1 Ag rapid test & Dengue IgG/IgM test). Ns1 Ag device, the rapid test membrane coated with anti-Dengue NS1 on test line. Sample interacts with monoclonal anti-Dengue NS1- gold moves along membrane to test line via capillary action to react with the anti-Dengue NS1 on the test line. If NS1 Ag is present a visible line appears at test line. For IgG/IgM Ab device monoclonal anti-human IgM and monoclonal anti-human IgG are immobilized at two individual test lines respectively on the membrane. Inactivated Dengue virus in the antigen pad and anti-Dengue-env-gold conjugate pad get released by adding buffer and react with anti-Dengue human IgM and IgG respectively on the membrane. If IgG and IgM Ab exist in patient serum, a visible band appears forming a complex with anti-human IgM/IgG, human IgM/IgG, inactivated Dengue virus, anti-Dengue-env-gold which means a positive result. A visible band at control region should always appear.

- Specific Performance Characteristics

| Sample panel | Dengue NS1 | Dengue IgM | Dengue IgG |
| --- | --- | --- | --- |
| Sensitivity | 98% | 97% | 97% |
| Specificity | 100% | 100% | 100% |

- Type of Sample

- Serum/Plasma or Whole Blood must be used for testing. Should a delay in testing occur, store the sample at 2-8°C. Samples can be stored for up to a week. Do not use hemolysed, icteric, lipemic serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel/ EDTA vacutainer.

- Equipment & Reagents

Test sample

Dengue Day1 kit

Buffer bottle

100µl pipette

Disposable dropper.

Centrifuge

- Safety

The use of disposable gloves and proper biohazardous clothing is strongly recommended while running the test.

In case there is a cut or wound in hand, do not perform the test.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in-vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

Protein-A Conjugate and Buffer Solution contain Sodium Azide as a preservative. If these materials are to be disposed of in a sink or other common plumbing systems, flush with generous amounts of water to prevent accumulation of potentially explosive compounds. In addition, consult the manual guideline "Safety Management No. CDC-22", Decontamination of Laboratory Sink Drains to remove Azide salts" (Centre for Disease Control, Atlanta, Georgia, April 30, 1976.)

- Calibration Procedure

- Centrifuge calibration is done on an annual basis by an NABL accredited calibration agency.

- Test Procedure

Dengue NS1 Ag

    - Bring all Kit components and specimen to room temperature prior to testing.

    - Remove the test device from the foil pouch and place it on a flat dry surface.

    - With a disposable dropper, add 2 drops (70µl) of specimen into the sample well.

    - Interpret/record the results at 20 minutes.

Dengue IgG/IgM

    - Bring all kit components and specimen to room temperature prior to testing.

    - Remove the test device from the foil pouch and place it on a flat, dry surface.

    - With 10µl capillary pipette, add 10µl of into the sample well marked “S”.

    - Add 2drops (about 70µl) of assay diluent into buffer well of test device.

    - Interpret/record the results at 20 minutes.

- Quality Control Procedure

Each testing device has an inbuilt Control Dot to confirm the validity of the device. The appearance of the control dot after completing the procedural test indicates a valid test.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – IAMM VIROLOGY EQAS, MHL EQAS

- Interfering Substances

Frozen Samples

      - The Dengue NS1 Test is best when used with fresh samples that have not been frozen and thawed. However, most frozen samples will perform well if the following suggested procedure is followed. 1. Allow the sample to thaw in a vertical position in the rack. Do not shake the sample. This allows particles to settle to the bottom. If a centrifuge is available, the sample can be centrifuged at 3500 r.p.m. for 15 min.

Thick or Viscous Samples

      - Whenever possible, clear specimens should be used. However, viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3500 rpm for 15 minutes and retested on a fresh device to avoid inconsistent results.

- Result Calculation

Dengue NS1 Antigen

    - Negative: The control line (C) is only visible on the test device. No IgG and IgM antibodies were detected.

    - Positive: The control line (C) and IgM line (M) are visible on the test device.

    - Invalid: The control line (C) is not visible. Repeat test with a new device

Dengue IgG/IgM

    - Negative: The control line (C) is only visible on the test device. No IgG and IgM antibodies were detected.

    - IgM Positive Result: The control line (C) and IgM line (M) are visible on the test device.

    - IgM and IgG Positive Result: The control line (C), IgM line (M) and IgG are visible on the test device.

    - IgG Positive Result: The control line (C) and IgG are visible on the test device.

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 4 Hours

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- Dengue NS1 Antigen and Dengue IgM antibody positive is considered as a critical alert.

- Laboratory Clinical Interpretation

- Dengue is considered to be the most important arthropod-borne viral disease due to the human morbidity and mortality it causes. Dengue may present as non-specific febrile illness or classic dengue fever. Dengue Haemorrhagic fever and Dengue Shock Syndrome are known life threatening presentations of infection. Antigen NS1 is found from the first day and up to 9 days after onset of fever in samples of primary or secondary dengue infected patients. The recommended test is NS1 Antigen detection by rapid immunochromatography or ELISA in the first 5 days of fever. After 7-10 days of fever, the recommended test is Dengue fever antibodies IgG & IgM by rapid immunochromatography or ELISA. Results by immunochromatography should by confirmed by ELISA, due to low Sensitivity.

- Potential Source of Variability

A negative result cannot exclude a recent Dengue Virus infection.

A negative result can occur if the quantity of antigen present is below the detection limits. Results should be correlated clinically and with other laboratory investigations.

Serological cross reactivity across the Flavivirus group is common.

- References

- Dengue Day 1 – J’Mitra kit insert